## Instructions
Part 5 Statement of Requirements - Instructions
HSVITS2021-055 Infusion Pumps
Introduction
· In addition to responding to the HSV Procurement Portal questions and Part 6 Tender Response Worksheet (TRW), this Worksheet (Specification) forms part of the Invitation to Supply (ITS) submission.
· It allows suppliers to confirm product compliance to the updated specification for evaluation.
· Please note the specifications advised in this Part 5 Statement of Requirements are the minimum specification requirements only, any equipment that exceeds these requirements are fully compliant.
· HSV request Respondents advise specifications in the unit of measurement outlined in this Part 5 Statement of Requirements for easy comparison between different manufacturers.
· Respondents are required to confirm compliance to the updated specifications, only for the items remaining on contract. Therefore where base units need to be listed, we only need to see the items as per the pricing schedule extract titled "active" and re-confirm compliance only against the cells highlighted in yellow.
Tab B Service, Delivery & Support
· Responses are limited to Fully comply, Partially comply, Does not comply, Not applicable
· Comments are required if item "Partially Comply" or "Does not Comply".
Parts C & D
· Each product in the Pricing Schedule Extract/ Tender Response Worksheet must be listed in each of these tabs and compliance to updated specification confirmed
· Responses are limited to Fully comply, Partially comply, Does not comply, Not applicable
· Comments are required if item "Partially Comply" or "Does not Comply".
· If required, respondents can insert additional columns
Category Tabs (Cat 1 - Cat 6)
· Each product listed within your extract must be listed in applicable "new" category and sub-category and compliance to specification confirmed
· Responses are limited to Fully comply, Partially comply, Does not comply, Not Applicable
· Comments are required if item "Partially Comply" or "Does not Comply".
· If required, respondents can insert additional columns
Tables
· Please confirm standard warranty information for each category
How and When to Use "Fully Comply, Partially Comply, Does Not Comply, Not Applicable"
· When to use "Not Applicable" & "Do Not Comply"
1. If Product ABC has alarms, it must beep for at least 30 seconds
Example 1: My product has no alarm. Response is "Not Applicable"
Example 2: My product has alarms but it only beeps for up to 25 seconds. Response is "Does Not Comply".
· When to use "Fully Comply" & "Do Not Comply"
2. Product XYZ must be between 150 - 200mm in length
Example 1: My product is 165mm in length. Response is "Fully Comply"
Example 2: My product is 145mm in length. Response is "Does Not Comply"
· When to use "Partially Comply"
3. The product must be >210cm in length
Example 1: Product is 195cm in length but with the optional extender, it is 215cm. Response is "Partially Comply".
NOTE:
Any addition to contract must be processed via the standard HSV contract variation process through the help desk. Noting this must be submitted post tender close against the new HSVC2021-055 Infusion Pumps Specification and new category/ sub-category listing.

## A - Introduction
Part A: Introduction
Clause Identifiers |  |  |  | Text
A |  |  |  | Introduction
| 1 |  |  | Purpose
|  | a |  | The purpose of this Part 5 - Statement of Requirements, is to:
|  |  | (i) | detail the scope and range of products sought under this Invitation to Supply (ITS)
|  |  | (ii) | specify the requirements that Respondents and/or their offered products must meet (these requirements also form part of any resulting Agreement between HSV and any successful Respondent)
| 2 |  |  | Scope
|  | a |  | HSV is seeking responses for Infusion Systems and associated consumables for use in Participating Health Services. The envisaged term is an evergreen arrangement (ongoing) with specified periodical reviews
|  | b |  | The scope of this ITS includes:
|  |  | (i) | the supply of Infusion Systems and associated consumables including optional features
|  |  | (ii) | supply of relevant consumables, semi-consumables and parts;
|  |  | (iii) | servicing and maintenance of equipment; and
|  |  | (iv) | delivery of a training package, including relevant training materials.
|  | c |  | The scope of this ITS does not include:
|  |  | (i) | infusion devices that are integrated with other medical devices including, but not limited to: • Phacoemulsifiers; • Endoscopy systems; and • Irrigations systems
|  |  | (ii) | implantable pumps
|  |  | (iii) | enteral feeding pumps;
|  |  | (iv) | enteral feeding sets
|  | d |  | HSV is seeking responses for various supply solutions, including but not limited to:
|  |  | (i) | outright purchase of equipment
|  |  | (ii) | leasing of equipment
|  |  | (iii) | supply of equipment only
|  |  | (iv) | purchase through set agreement
|  |  | (v) | supply of consumables and/or semi-consumables only
|  |  | (vi) | whole-of-life support/managed supply environment, including but not limited to:
• supply of equipment
• maintenance (including scheduled inspection and unscheduled repairs)
• warranty services
|  | e |  | Indicative volumes are yet to be determined.
| 3 |  |  | Categories
|  | a |  | The categories called for under this Invitation to Supply include:
Volumetric pumps
Syringe pumps
Dedicated pumps (e.g. Ambulatory, PCA, TIVA, etc.)
Pump platforms (incorporating multi-pump racks, docking stations, and interoperability)
Dose Error Reduction Systems
IV administration sets (consumables)
|  | b |  | The Respondent may offer products in one, some or all categories.
|  | c |  | HSV reserves the right not to consider any additional products offered.
| 4 |  |  | Product Conditions
Clinical Trials
|  | a |  | Participating Health Services may, at their discretion, research or trial new technology or use non-contracted products to perform clinical trials at any time during the Term of any resulting Agreement
|  | b |  | Participating Health Services may request trials before creating a Statement of Demand which may also provide successful Respondents time to do a site inspection.
Product Duplication
|  | d |  | HSV will not consider any product that is subject to a current HSV Agreement.
|  | e |  | The Respondent must ensure that each product is offered in only one subcategory. It is at the Respondent’s discretion to ensure that each product is submitted in the most appropriate subcategory.
Product Information
|  | f |  | The Respondent must submit a copy of relevant product diagrams, specifications or brochures to assist in accurately identifying products offered.
|  | g |  | Where research papers or relevant scientific information is available, this should be submitted with the Respondent’s tender.
|  | h |  | All product information submitted must:
|  |  | (i) | be in electronic format
|  |  | (ii) | be in English
|  |  | (iii) | be specific to the product offered
|  |  | (iv) | contain the Respondent’s company name
|  |  | (v) | include the product code
|  |  | (vi) | include a detailed specification of the product
|  |  | (vii) | include clear diagrams/pictures of the product.
|  | i |  | To assist in managing this material, all product information submitted must be labelled with the relevant HSV category and subcategory number.
|  | j |  | Electronic copies must include the HSV category and subcategory numbers in the filename or identifying metadata.
|  | k |  | HSV may not consider unlabelled submissions
|  | l |  | Product information will not be evaluated, but will assist in accurately identifying products offered.
|  | m |  | HSV reserves the right to exclude products from evaluation when the product is unidentifiable and provided information is:
|  |  | (i) | Not labelled as per D-F above; or
|  |  | (ii) | Is incomplete as to C.
|  | n |  | Product samples are not to be provided unless specifically requested by HSV, as per Part 3 – 8 Samples.
|  | o |  | The Respondent should not submit information relating to products that are not called for in this ITS.
Other Items
|  | p |  | Products offered in ‘other’ subcategories will only be considered where the product meets the specification and the Respondent is successful in at least one of the specific subcategories of that category.

## B - Service, Delivery & Support
Part B: Service, Delivery and Support
Clause Identifiers |  |  |  | Text
B |  |  |  | Statement of Work | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Delivery
|  | a |  | All deliveries are bound as per the Draft Agreement, Part 7 – 9 Acceptance and Rejection of Deliverables. | --Select--
|  | b |  | Timeframes are as per the table Delivery, unless otherwise agreed with the Participating Health Service. | --Select--
|  | c |  | Successful Respondents must be capable of installing their devices and/or systems and, if requested to do so, must cooperate with Participating Health Services’ biomedical engineering department. | --Select--
|  | d |  | Successful Respondents must provide the Participating Health Service’s authorised users with one (1) copy of the software tool used for configuration and training on using the software tool and configuring the pumps/system prior to delivery. This requirement is per model purchased and includes any special cables, dongles etc. This software and initial training must be provided at no additional charge as part of the supply of infusion pumps/systems, and is in addition to the training requirements under section 5.9 (Training). | --Select--
|  | e |  | Respondents must provide the Participating Health Service with all test reports, if any were conducted on the individual device, prior to the device being shipping. | --Select--
|  | f |  | Successful Respondents must notify HSV and Participating Health Services of ‘end-of-manufacture’ date for all equipment, accessories, and consumables at least twelve (12) months prior to the date manufacturing ends. | --Select--
Backorders and discontinued lines
|  | g |  | In the event that a product is unavailable at the time of placing a Purchase Order, Respondent must (at a minimum) advise the Clinical Products Advisor (where applicable), Supply Departments and HSV. | --Select--
|  | h |  | Respondents must inform the affected Participating Health Service and HSV of:
|  |  | (i) | the anticipated timeframe for resolving the issue/ backorder, and | --Select--
|  |  | (ii) | the availability of an agreed substitute product | --Select--
|  | i |  | In an anticipated stock outage, Respondents are to notify Participating Health Service and HSV of estimated timeframe of the outage, availability of an agreed substitute product and recovery action plan within 48 hours. | --Select--
| 2 |  |  | Training
|  | a |  | Successful Respondents will be required to develop a training package and training materials to facilitate the introduction of their infusion pumps/systems and associated products to clinicians in their operating environment that is relevant, and complies with best practice guidelines for infusions in Victoria, at the time of training. | --Select--
|  |  | (i) | Such training and/or materials must be available to Participating Health Services at the time of purchase, and | --Select--
|  |  | (ii) | may be offered electronically (e.g. online or CD/DVD) in addition to, | --Select--
|  |  | (iii) | face-to-face training, | --Select--
|  |  | (iv) | including DERS training for pump users that will vary between each Participating Health Service. | --Select--
|  | b |  | Staff training requirements for infusion pumps/systems is a significant cost to Participating Health Services. To assist Participating Health Services in minimising these costs, successful Respondents will be required to train a minimum of two (2) supers users (train-the-trainer), or one (1) super user for every ten (10) infusion pumps/system purchased , whichever number is greater. | --Select--
|  | c |  | If requested by a Participating Health Service, successful Respondents must provide a plan detailing how they will provide training to nominated staff. The number of staff involved in training will vary dramatically between Participating Health Services. | --Select--
|  | d |  | Training may be required during both day and night shifts and on weekends. | --Select--
|  | e |  | Successful Respondents must provide training to Participating Health Service users for the effective life of the equipment. This is to ensure that new employees have access to the manufacturers’ level of training. | --Select--
|  | f |  | Training requirements for Category 5 - Dose Error Reduction Systems are expected to be different due to the greater complexity of these systems. As per clause Part B 2a and 2c, successful Respondents must provide a separate training plan for equipment in this category upon request by Participating Health Services. | --Select--
|  | g |  | Infusion pump/systems and associated products should be compatible and easily interface with commonly used training equipment (e.g. manikins). | --Select--
|  | h |  | Infusion pumps that are expected to be used in the home will come with training documentation for each device purchased. This level of documentation is to be aimed at the home user level. | --Select--
| 3 |  |  | Service and Maintenance
|  | a |  | Service and maintenance must be available for a minimum seven (7) years from the date of delivery for each infusion pump/system; preference may be given to Respondents who can offer service and maintenance for longer than seven (7) years. | --Select--
|  | b |  | Servicing and maintenance should be available throughout the state of Victoria. This may be accomplished by a courier service to collect items offered in Categories 1 through 4 and return to “supplier” for servicing, but on-site representation will be required for Dose Error Reduction Systems (Category 5) unless the problem can be resolved via phone support, where requested (as per clause Part B 3e). | --Select--
|  | c |  | Successful respondent must obtain written approval from Participating Health Services before any maintenance is performed on Participating Health Service’s DERS server. | --Select--
|  | d |  | Service Timeframes - see table "Service" | --Select--
|  | e |  | Service support via phone must be available to all Participating Health Services within two (2) business hours of lodging a request at no cost. Successful Respondents’ staff providing phone support must have relevant clinical or technical knowledge of the products they are supporting. | --Select--
|  | f |  | Successful Respondents must provide Participating Health Services with a list of all commercial service providers that are capable of, and authorised to, provide ongoing support. | --Select--
|  | g |  | Respondents must provide a full range of service contract options for each item of equipment offered. This may range from full service agreements (24 hours a day, 7 days a week) to routine testing contracts. Participating Health Services reserve the right to enter into these contracts during the entire, or part of, the effective life of the equipment. | --Select--
|  | h |  | Respondents must provide Participating Health Services with a service report for any works conducted on device, along with test reports that were conducted post works. | --Select--
|  | i |  | Successful Respondent must provide training certificates to Participating Health Services, of staff employed by Respondent to maintain tendered product, when requested by a Participating Health Service. | --Select--
| 4 |  |  | Spare Parts and Repairs
|  | a |  | Spare parts must be available for a minimum of seven (7) years from the date of last manufacture for each tendered product. | --Select--
|  | b |  | Ensuing clause 4a, service and repair spare parts must be available for purchase by Participating Health Services regardless of the service training status of Participating Health Services’ biomedical engineers. | --Select--
|  | c |  | If requested by HSV / Participating Health Services, successful Respondents must provide a comprehensive list of all replacement parts, accessories and assemblies that are likely to be required to repair and maintain the equipment over the supported life. The list must include part numbers and current purchase prices. | --Select--
|  | d |  | If requested by a Participating Health Service, successful Respondents must provide non-warranty repairs for their equipment for its supported life. Where practical, consumables and semi-consumables should be replaced rather than being repaired. | --Select--
|  | e |  | All parts should be available,
|  |  | (i) | within five (5) business days for Categories 1, 2 and 4, and | --Select--
|  |  | (ii) | within 24 business hours for Category 3, as these are mission-critical. | --Select--
|  | f |  | If the successful Respondent(s) cannot meet the timeframes specified in clause e above, they must notify Participating Health Services before the expiry of these timeframes. | --Select--
|  | g |  | In the event that a Category 3 item is damaged (regardless of whether the damage is covered under warranty) and a repair cannot be completed within the timeframes specified in clause e above, successful Respondent should supply an interim comparable item (i.e. a loan unit) to the relevant Participating Health Service. Preference may be given to Respondents with sufficient coverage of loan unit pumps. | --Select--
|  | h |  | Successful Respondents must notify HSV and Participating Health Services of ‘end of support’ for all equipment and associated consumables and semi-consumables at least three (3) years before the end of support date occurs. | --Select--
|  |  | (i) | If an upgrade is required as part of corrective maintenance to rectify a single fault, then the cost of this upgrade must not exceed 25% cost of the tendered product, i.e., if the parts costs including upgrade replacement parts exceed 25% of a new device, then the device will be considered end-of-life. | --Select--
|  | i |  | Ensuing clause h(i) above, the choice of whether to upgrade or purchase new equipment is at the discretion of Participating Health Services. | --Select--
|  | j |  | During the product lifecycle, all firmware must be upgradable as new technology becomes standard. | --Select--
| 5 |  |  | Customer Service and Support
|  | a |  | The successful Respondent must be able to deliver 24-hour, 7-day customer service and support to Participating Health Services. | --Select--
|  | b |  | The successful Respondent must provide Participating Health Services with representatives that are:
|  |  | (i) | inherently familiar with the contracted products | --Select--
|  |  | (ii) | appropriately qualified | --Select--
|  |  | (iii) | technically / clinically knowledgeable about the contracted products | --Select--
|  |  | (iv) | available to respond to Participating Health Services’ queries 24 hours a day. | --Select--
| 6 |  |  | Warranty
|  | a |  | Equipment, parts, cables, batteries and all other semi-consumables must be warranted for normal use from the date introduced to clinical service, i.e., warranty period commences from the first date of clinical use. Participating Health Services are responsible for maintaining accurate records of first clinical service. | --Select--
|  | b |  | All equipment and semi-consumables offered must have a minimum 12-month warranty period. | --Select--
|  |  | (i) | Preference may be given to equipment and semi-consumables with warranty period longer than 12 months. | --Select--
|  | c |  | Where practical, consumables and semi-consumables should be replaced rather than repaired. | --Select--
|  | d |  | The manufacturer’s warranty must not be impacted by the use of generic consumables (including batteries). | --Select--
|  | e |  | If available, Respondents should provide details of any extended warranty that is available beyond the manufacturer’s warranty (including the cost of such warranty per additional year). It will be at the discretion of Participating Health Services whether or not to purchase any extended warranty. | --Select--
Repairs and Replacements under Warranty
|  | f |  | The repair or replacement of any equipment or parts under warranty will be at no cost to Participating Health Services; excluding situations where fault is traced to user error or damage, or non-compliance to manufacturer's recommendation. | --Select--
|  | g |  | The cost of any pickup or delivery associated with a repair or replacement under warranty will be borne by the successful Respondent. | --Select--
|  | h |  | Items repaired under warranty must be returned to Participating Health Services within five (5) business days from when the item is received by the successful Respondent. | --Select--
|  | i |  | If the successful Respondent is not able to meet the timeframes specified in the above clause, then the successful Respondent must provide Participating Health Services with a suitable replacement item of the same make and model until the repaired item is returned. This will be done at no cost to Participating Health Services. | --Select--
| 7 |  |  | Biomedical Equipment Maintenance and Repair
|  | a |  | Respondents agree and understand that Participating Health Services may have an internal biomedical engineering department, or use the services of an external biomedical engineering contractor, to perform the following for the offered systems: | --Select--
|  |  | (i) | acceptance testing of new equipment (required before clinical use) | --Select--
|  |  | (ii) | safety and performance testing | --Select--
|  |  | (iii) | assessing whether a repair is a warranty issue or not and, in the case of a warranty issue, coordinating the repair with the successful Respondent | --Select--
|  |  | (iv) | performing non-warranty repairs when requested | --Select--
|  |  | (v) | performing the required routine preventative maintenance schedule. | --Select--
|  | b |  | Successful Respondents must offer annual service training to Participating Health Services’ biomedical engineering department or other nominated persons, and specify how and when such training will be delivered and cost of (i.e. annual seminar, face-to-face training at Participating Health Service, virtual, costs, etc.). | --Select--
|  |  | (i) | The initial service training for each Participating Health Service must be provided free of charge after the purchase of a new model not currently used at that Participating Health Service. For a complete fleet replacement service training must be provided on-site. | --Select--
|  | c |  | Post-training, successful Respondents who use software tool for servicing and/or configuring their infusion pumps/systems must provide a minimum of one (1) copy of this software to authorised users free of charge. This includes any special cables, dongles etc. | --Select--
|  | d |  | Post-training, successful Respondents who use tools for servicing and/or configuring their infusion pump/system must provide a comprehensive list of common tools that are likely to be required to service and/or configure the equipment over the supported life. The list must include part numbers and current purchase prices. | --Select--
|  | e |  | The training referenced in this section 8, Biomedical Equipment Maintenance and Repair is in addition to the successful Respondents’ training requirements under section 3, Training. | --Select--
|  | f |  | Upon purchase of any infusion pump/system, and upon request by Participating Health Services, successful Respondents must provide both hard and electronic copies of:
|  |  | (i) | complete technical service manuals in English for repairs to the systems and accessories, and | --Select--
|  |  | (ii) | all documentation on the supporting IT systems, hardware, operating systems (OS) and software applications, and | --Select--
|  |  | (iii) | these must be provided at no cost regardless of the service training status of the biomedical engineering department or nominated persons, and | --Select--
|  |  | (iv) | comprehensive operator’s manuals in English provided at no cost to Participating Health Services, as many times as requested, for the life of agreement. | --Select--
|  | g |  | Upon purchase of any infusion pump/system, Participating Health Service must have full and autonomous access to all IT systems, without having to contact successful Respondents for RSA key to gain access. | --Select--
|  | h |  | Preference maybe given to Respondents who offer a web portal for Participating Health Services to register their maintenance requests. | --Select--
| 8 |  |  | Trial and Evaluation
|  | a |  | Where practical, Respondents may be required to arrange a site visit during the evaluation stage for trial and evaluation purposes. | --Select--
|  | b |  | Participating Health Services reserve the right to trial successful Respondents’ products on-site for reference prior to purchase. To limit the costs of any such trial, full integration to hospital systems is preferred but not required. Trials will not exceed four (4) weeks clinical trial per Participating Health Service site. | --Select--
|  | c |  | Suppliers must be capable of demonstrating their products within four (4) weeks of request. | --Select--
|  | d |  | As part of the trial and evaluation, consumables required for the product, should be FOC | --Select--
| 9 |  |  | Key Performance Indicators
|  | a |  | Refer to Draft agreement clause - Key Performance Indicators. | --Select--
| 10 |  |  | Reporting
|  | a |  | Refer to Draft agreement clause - Reporting Requirements. | --Select--
| 11 |  |  | Service Level Agreement
|  | a |  | Participating Health Services may enter into a Service Level Agreement (SLA) with the successful Respondent(s). | --Select--
|  | b |  | The SLA will be in addition to the Agreement between the successful Respondent(s) and HSV, and will not alter any terms of the Agreement. | --Select--
|  | c |  | HSV will not be responsible for monitoring compliance with any SLA. This is a matter of agreement between the parties to the SLA. | --Select--

## C - General Requirements
Part C: General Requirements
Clause Identifiers |  |  |  | Text | Category 1 (Enter model number) |  | Category 1 (Enter model number) |  | Category 2 (Enter model number) |  | Category 2 (Enter model number) |  | Category 3 (Enter model number) |  | Category 3 (Enter model number) |  | Category 4 (Enter model number) |  | Category 4 (Enter model number) |  | Category 5 (Enter model number) |  | Category 5 (Enter model number)
C |  |  |  | General Requirements | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Please specify or comment | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Standards and Compliance
|  | a |  | All items offered must comply with the relevant Australian Standards (or their equivalent, Australian-recognised International Standards). Refer to Appendix 2 – References for a list of the minimum relevant standards. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All items offered must comply with a relevant Electromagnetic Compatibility standard/s. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered products must not cause electromagnetic interference to other medical devices. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Manufacturer of tendered item, must comply with AS ISO 14971. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | All equipment, consumables and semi-consumables offered must be approved by the Australian Therapeutic Goods Administration (TGA). ARTG number for each tendered product must be listed in the Product Catalogue. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | All sterile consumables must have CE / ISO Standard sterilisation certificates. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Successful Respondents must ensure the ongoing inclusion of tendered medical device in the ARTG, for the period of this agreement by meeting the associated ongoing annual TGA charges. Failure by Respondent to maintain the tendered medical device on the ARTG will result in product suspension from HSV Catalogue. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | All respondents (suppliers) must be compliant with AS ISO 13485. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Does the tendered product have the capability to connect to the hospital's IT network infrastructure? (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Does the tendered product have the capability to connect directly to the Internet (i.e., via public Wi-Fi, internal 4G/5G radio)? (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
If yes, to either of the above two questions:
|  |  | (i) | Respondents must include the Manufacturer Disclosure Statement for Medical Device Security (MDS2) form (the latest version of form, at time of tender version ANSI/NEMA HN 1:2019) for each tendered product with their submission. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | State the Document ID number and Document Release Date on the MDS2 form for tendered device | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Respondents must list the Software Bill of Materials (SBoM) for each tendered product in the Product Overview. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Successful Respondents must provide Participating Health Services the MDS2 form for tendered devices upon the Participating Health Service's request. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | Respondent must ensure that the Manufacturer has a Vulnerability Disclosure Program in place for the total tendered product's life cycle; which monitors, informs Participating Health Services, and addresses security vulnerabilities. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (vi) | Respondents must have a Cyber Incident Response Program that will assist a Participating Health Service in the event that the tendered product succumbs to a cyber incursion. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Packaging and Labelling
|  | a |  | All labels must comply with the Therapeutic Goods Order No. 37: General Requirements for Labels for Therapeutic Goods. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The shelf life must be clearly marked on product packaging for all consumables and applicable semi-consumables (such as batteries). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Recall Process
|  | a |  | All recalls must be managed in line with the TGA’s Uniform Recall Procedure for Therapeutic Goods (URPTG) (V2.2, December 2019 or later, whichever most current). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All TGA recalls and/or hazard alerts must be completed using ‘GS1 Recall Health’, as endorsed by NEHTA. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Class I recalls (as defined by the TGA’s Uniform Recall Procedure for Therapeutic Goods) must also meet the requirements under B.7 - Warranty. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Essential software updates that are life threatening must be completed within three (3) months of notification, regardless of warranty status. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Essential software updates that are not life threatening must be completed within six (6) months of notification, regardless of warranty status. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | As soon as intelligence report arises that a vulnerability within tendered product is being actively exploited by threat actors, Successful Respondent must notify Department of Health (Digital Health team), HealthShare Victoria, and Participating Health Services, within 72 hours of report arising. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Product Information and Identification
|  | a |  | A minimum of two (2) electronic copies of the operating manual in English must be provided to Participating Health Services for each equipment system purchased, unless otherwise specified. Additional copies (in hard copy or soft copy) should be provided free of charge upon request by Participating Health Services. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Respondents must provide a 'Quick Reference Guide' on all tendered items upon the request of Participating Health Service. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondents must provide complete Technical Service Manuals that includes full list of all replaceable spare parts, to the requesting Participating Health Service upon their purchase of tendered item. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All associated consumables and semi-consumables for equipment systems must be listed in a separate document and this must be provided to Participating Health Services upon purchase of each device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Within one (1) year of contract commencement, all consumables and semi-consumables must be identifiable using GTINs and must be listed on the National Product Catalogue (NPC). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Cleaning and Infection Control
|  | a |  | To meet the Australian Guidelines for the Prevention and Control of Infection in Healthcare (2010), medical equipment must be capable of being cleaned with a variety of cleaning agents in order to clean and disinfect the equipment, as prescribed in AS/NZS 4187:2014 Reprocessing of reusable medical devices in health service organizations. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All items offered must be capable of being cleaned without degradation with the following cleaning or disinfectant agents:
|  |  | (i) | Neutral Detergent, pH 6 -8 in either an aqueous solution or impregnated wipe, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Alcohol 70-90%, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | At least, one of the following hospital grade disinfectant categories:
|  |  |  | Chlorine and chlorine compounds, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | Hydrogen peroxide, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | Quaternary ammonium compounds, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | Phenolic disinfectants. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Upon request by Participating Health Services, successful Respondents must provide a full list of cleaning and disinfection products (including wipes) approved for use, including maximum permissible concentration level of active ingredient. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Upon request by Participating Health Services, successful Respondents must provide a list of commonly used cleaning and disinfection products that are not suitable and maybe harmful to the surfaces of Tendered equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | To facilitate cleaning, tendered products must have:
|  |  | (i) | smooth surfaces with no sharp edges, deep crevices or other dirt traps and be able to withstand hospital grade cleaning agents | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | no Velcro | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | surface finishes that are compatible with hospital grade commercial cleaning equipment | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Pump Cases
|  | f |  | Pump cases must:
|  |  | (i) | be made from non-porous material that is easy to clean | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | be of a colour that easily shows up contamination, state case colour/s | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | not have any excessive holes, channels or crevices. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | must not have any sharp edges | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Transport Solutions
|  | a |  | Unless otherwise stated, all items must be an appropriate size and weight to satisfy an ergonomic risk assessment (as described in WorkSafe Compliance code – Hazardous manual handling. Edition 2, December 2019). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Systems offered with a pole mounted solution on wheels preference will be given to solutions with a (5) five-legged roll-stand and with a tube management system for the tubes provided. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Any mobile device with wheels/castors should be of a robust design, with optional braking systems offered. Castor size options should be given | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Any roll-stand offered should incorporate mounting solutions for hygiene products (both hand hygiene and equipment wipes). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Durability
|  | a |  | All devices must be able to withstand all tests prescribed in AS/NZS IEC 60601.1:2015 Section 15 Construction of ME Equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All equipment and semi-consumables must be able to endure the use of hospital cleaners and disinfectants as listed in manufacturer's instruction manuals, without prematurely aging or damaging the equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All equipment and semi-consumables must be able to endure the Respondent’s full list of cleaning and disinfection products without prematurely aging or damaging the equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Sustainability
|  | a |  | Preference may be given to items with recyclable packaging. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Preference maybe given where supplier offers to remove transport packaging. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Successful Respondents must have a plan for the removal and environmentally responsible management of tendered equipment from Participating Health Services on the request of the Participating Health Service. Note not all Participating Health Services will require this service. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Any associated costs to this Offer (e.g., a collection charge) should be negotiated under the Site Level Agreement with the Participating Health Service. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Only equipment supplied by the Successful Respondent is required to be subject to this Offer. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Equipment taken-back prior to their end-of-support life may be reconditioned and sold or donated, or disposed of; as appropriate to the specific equipment and circumstances. Preference is that functioning, useful productive and supported equipment, is not disposed of as waste. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Electrical and electronic waste (e-waste) must not be disposed to landfill and must be sent to a recycler with certification under the e-waste standard AS/NZS 5377:2013. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | Successful Respondents must report annually to HealthShare Victoria on the Take Back services provided to Participating Health Services under this Offer using the Reporting Schedule provided. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference may be given to tendered products that still have a 'long time' remaining in their product life cycle | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | state the year manufacturer first sold tendered product
|  |  | (ii) | state estimated length, in years, of tendered product's product life cycle

## D - Product Specifications
Part D: Product Specifications
Clause Identifiers |  |  |  | Text | Category 1 (Enter model number) |  | Category 1 (Enter model number) |  | Category 2 (Enter model number) |  | Category 2 (Enter model number) |  | Category 3 (Enter model number) |  | Category 3 (Enter model number) |  | Category 4 (Enter model number) |  | Category 4 (Enter model number) |  | Category 5 (Enter model number) |  | Category 5 (Enter model number)
D |  |  |  | Product Specifications | Fully comply/ Partially comply/ Does not comply | Please specify | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Power Supply
|  | a |  | Unless otherwise specified (such as ambulatory pumps), all infusion pumps/system must work from standard 240V AC power points (if electrically driven), as supplied in Australia. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference may be given to infusion pumps/system that operate from a standard mains cable and do not require an external power supply plug pack | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Infusion pumps/system that are powered from an external plug pack must include a solution for securing the low-voltage cable to the 'pump'. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | UPS that are required to ensure an orderly shutdown or to deal with temporary power shortages for DERS, must have ‘hot swappable’ battery capability. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | If the DERS requires electrical power for normal operation, a backup power supply should be a part of the machine and permit normal operation for at least 30 minutes after a mains power supply failure. An alarm must be activated at the time of the mains failure and the state of the reserve power supply must be indicated while it is in use. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | It must be possible to immediately restart infusion on tendered product on mains power with a depleted battery | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Infusion pump/system batteries must be capable of charging independent of the state of the on/off switch. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The pump must have an integral battery that provides power for at least 2 hours of normal operation and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | a depleted rechargeable battery must be fully recharged in less than 12 hours. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Pumps should retain all pump settings during a battery change over. (minimum 5 minutes) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | For pumps with rechargeable batteries, preference may be given where charger is built-in the pump unit. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Preference will be given to pumps that provide for removal of the administration line in the absence of power, for example, if the battery is flat. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | If a mains switch is present, it must be protected from unintended activation or deactivation | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | Switching ‘off’ a pump platform/infusion system during normal operation must require a confirmatory step and/or the machine should display a warning of imminent shut-down. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Batteries
|  | l |  | For items that use rechargeable batteries preference may be given to products provided with rechargeable batteries that do not require cycling to maintain capacity. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | The rechargeable battery life must be at least one (1) year, with tolerance to partial charging. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | All items with battery power must have a battery level indicator. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | A depleted battery must be accompanied by an audible alarm conforming to ISO and IEC alarm standard 60601-1-8 and visual indication signalling interruption of infusion. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | p |  | Batteries should have the ability to remain charged for extended periods and should have short recharging times. Upon request, successful Respondents must advise time battery remains charged and charging time. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | ‘Button’ batteries are not desirable due to these being a swallowing hazard. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | r |  | Infusion pumps intended to be run on disposable batteries should be of a standard type such as size “AAA”, “AA”, “C” or 9 volt. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | s |  | For multi-channel pumps the battery life for the category standard flow is:
|  |  | (i) | 2 channels must be minimum 3 hours; | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | 3 Channels must be minimum 2.5 hours | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | 4 Channels must be minimum 2 hours. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | t |  | Preference may be given to battery systems, which can charge/condition the battery separate to the pump. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | u |  | Respondents must advise battery chemistry, e.g. alkaline, SLA, Ni-MH, Li-ion, etc. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | v |  | Upon request, successful Respondents must advise if battery incorporates disabling timers or software | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
UPS
|  | w |  | UPS must have:
|  |  | (i) | hot swappable battery | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | run time capacity of at least 30 minutes without mains power | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | audible alarm of at least a minimum of 65dB | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | connection to the building infrastructure management system at Participating Health Service | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Cables and Giving Sets
|  | a |  | All cabling must be robust in design and fit for purpose. Cabling must be flexible for easy use. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Giving sets must be routed within the device in such a manner that the giving sets do not easily get kinked or damaged during normal use. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Preference may be given to products with a cable management system | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Programming
|  | a |  | It must be possible for authorised users to program pump default start-up parameters. Password protection required. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Certain settings should be programmable, including initial alarm volume, alarm settings, display format etc. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | When a patient demographic is altered, devices should automatically readjust set rates, pressures and volumes and must be confirmed. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Neonatal/Paediatric Compatibility
|  | d |  | Devices must clearly indicate which drug library in in use, if applicable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Consumables and Semi-Consumables
|  | a |  | Note: further specifications are provided for giving sets in Category 6 | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | It is desirable for all infusion pumps/system to be compatible with non-proprietary general consumables and semi-consumables (this includes batteries, where applicable). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The shelf life of all consumables and applicable semi-consumables (such as batteries) must be clearly marked on the product and/or product packaging. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Consumables and semi-consumables must be available as a separate purchase option (i.e. priced separately to starter packs). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Human Factors
|  | a |  | Switches and control must be protected from accidental setting changes. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference may be given to infusion pumps/systems with clearly marked switches, audio prompts and other controls, and to systems that clearly indicate chosen selections. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All pumps must be of robust construction from durable materials. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference may be given to devices with no rough or sharp edges. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Infusion pump/systems must incorporate self-diagnostics and report problems via visual and audio prompts. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Additional quick diagnostic cards may be provided to supplement this. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Infusion pump/systems must log all activities for at least 24 hours (or longer), and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | all data must be available for Participating Health Service to download and interrogate if required. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Infusion pump/systems log must have date/time stamp and include an explanation of each error code and alarm condition logged. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Preference may be given to products that give an explanation for error code | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | If any item contains latex (whether in the delivery tubing or not) this must be clearly labelled on the item and its packaging (where applicable). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Preference may be given to products (including their accompanying packaging) that are latex-free, unless otherwise stated. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Preference may be given to product that are either DEHP-Free, or contains DEHP-Free internal lining | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Preference may be given to products that provide protection against light exposure. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Alarms
|  | i |  | Audible alarms must have adjustable volume, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | minimum alarm volume must remain sufficient to hear in the presence of moderate ambient noise (65dB). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | All alarms should be adjustable at a Health Service level, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | for each alarm condition, health service must be able to select high and low alarm volumes. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Alarms should differ between critical level (i.e. life threatening), medium level and information alert, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | differing criticality level alarms must be easily identified by different tones and colour indicators. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | If alarm volume can be muted, the unit must have a visual indicator that clearly identifies when the volume is muted. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | All pumps must have the following alarms:
|  |  | (i) | High pressure | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | System malfunction | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Low battery | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Battery depleted | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | Downstream occlusion | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | p |  | Preference may be given to systems that have the capabilities of forwarding infusion pump alarms onto a third-party alarm vendor’s system (i.e., Connexall, Spok, Vocera) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | Specific alarms, i.e. for DERS, are further detailed in the various categories | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Displays
|  | r |  | Displays must be clear and easy to read from a variety of angles and distances and in a variety of lighting conditions. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | s |  | Displays should be an appropriate size for the type of system offered and its intended application. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | t |  | Pump must display current flow settings for user to rapidly identify. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | u |  | Preference may be given to equipment with a visual indicator that the pump is running. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | v |  | Respondents may offer pumps with a variety of display formats. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | w |  | Part D, 10 i | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | x |  | Preference may be given to coloured displays that comply with the national standards of medication colour code. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | y |  | Preference may be given to day/night display. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Construction
|  | a |  | Preference may be given to pumps that offer a colour-coded system able to easily identify different clinical uses/applications of the pump. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The infusion pumps/systems must be protected against fluid spills due to breakage or disconnection of IV solution containers, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | patient safety must not be compromised by fluid spills, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | pump's operation must not be adversely affected by fluid spills, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | if fluid penetrates the electronic circuitry, the pump must audibly alarm and cease operating. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The controls must be sealed against penetration of liquids. The respondent must list the level of protection in the Product Catalogue, i.e., IP4 etc. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | MRI Conditional
|  | a |  | All infusion pump/system that are MRI conditional must be offered with appropriate MRI-conditional consumables and semi-consumables. Upon request, successful Respondents must advise appropriate MRI-conditional consumables and semi-consumables. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Preference may be given to infusion pump/system and associated products that can be used in a 3T (Tesla) environment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | state the maximum magnetic field strength (T) tendered product can be used in
|  |  | (ii) | state the maximum gauss line to which tendered product is specified to operate in
|  |  |  | Respondents should provide various mounting solutions for offered MRI-conditional infusion pump/systems. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Remote display of patient parameters maybe desirable. Respondents should provide various optional configurations such as display sizes, enunciator options, printing/recording options etc. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Hyperbaric Chamber Compatibility
|  | a |  | Define product hyperbaric chamber combability rating
| 9 |  |  | Aviation Compatibility
|  | a |  | Define product aviation combability rating
| 10 |  |  | Data Requirements for all Categories
|  | a |  | The trend for patient record keeping is to be in electronic format. To assist this, patient identification of the information from the infusion pump/systems is becoming more common. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Barcode readers may be offered as a purchasable option. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For keyboards offered as a purchasable option, preference may be given to keyboards that use standard types of USB connectors (i.e. standard, mini or micro). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Keyboards used in infection sensitive areas must be able to be cleaned using standard hospital cleaners. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Data integration with HIS (as defined by the Participating Health Service) may be offered as a purchasable option. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Preference may be given to systems that can accept additional data fields (such as weight and height). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Briefly outline connection method and information transfer to a Clinical Hospital Information System (HIS) Specify capability for each of the following modes: Wi-Fi, ethernet, RS232, USB, other.
|  | h |  | Briefly outline connection method and requirements to connect to other patient systems. Include a list of other patient systems (make and model) that the infusion pump/systems can communicate with.
Data Transmission
|  | i |  | Preference will be given to infusion pumps/systems that can communicate wirelessly | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | If data transmission is available, infusion pumps/systems must be able to transfer data via industry standard connections. For example: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | RJ45 (Ethernet) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Wi-Fi | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state wi-fi version
• specify the broadcasting features; broadcast, multicast or unicast
|  |  | (iii) | Zigbee | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | LoRa | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | if other, please specify
|  | k |  | Data transmission via Bluetooth (of any version) is not permitted | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Data transmission must be encrypted for privacy requirements, and state the cryptographic method. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Data transmission to other systems is preferred in standardised languages, such as Health Level Seven (HL7). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free upon purchase. Details of all conditions and subsequent annual costs must also be provided pre-purchase. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Event Reporting
|  | o |  | What event reporting information can you provide
|  | p |  | All infusion pumps/systems with data transmission capabilities must have the capability to pass event data wirelessly, via cable or via a memory device (such as a memory card) to other devices. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | A summary report should be downloadable from the infusion pump/system in a universal format so that it may become a component of patient records. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | r |  | Event information should be able to be stored on the host Participating Health Services’ information systems. This preferably will include assessments of all systems such as keypad interactions, alarm settings & events, battery performance. Etc.. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | s |  | For research and development purposes, data format should be open ECG-SCP, RCG or similar Open Same Standard format. If only a proprietary format is available, a software development kit for Microsoft Windows should be provided as a purchasable option. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | t |  | Patient data that is stored on the internal memory should be encrypted to an open standard (i.e. non-proprietary). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | u |  | A data copy must be able to be stored in such a manner that it is incorruptible for use in any potential legal analysis | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
IT Security
|  | v |  | Any tendered product that can connect onto the hospital's IT network, whether physical or virtual, must:
|  |  | (i) | have unique user login credentials (both generic and group logins must not be used), and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | user level of privileges managed, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | must not use single-factor user login authentication protocols, i.e., must use multi-factor secure user login authentication protocols, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Participating Health Service must be have access to an admin account. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | w |  | Successful Respondents must have a program in place that informs Participating Health Services of cybersecurity vulnerabilities to the tendered device throughout its life cycle. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | x |  | Where a Tendered product is installed with an Operating System (OS), the OS should be actively supported and patchable for the full life of the equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | OS patch updates must be installed in affected tendered product within 30 days from patch being released from OS provider. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | y |  | Tendered product software program (app) must have a patch update installed within 90 days upon an app vulnerability having been identified at no additional cost to Participating Health Services. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | z |  | Where a vulnerability whether in the OS, or the app of tendered device is being actively exploited by threat actors, the Successful Respondents must provide Participating Health Services: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | advice on implementing compensating control measures to mitigate vulnerability being exploited, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | a patch update within 48 hours to remediate the vulnerability. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | aa |  | Health Service Endpoint protection solutions (anti-malware, application control, whitelisting functionality, etc.) must be permitted to be enabled and/or installed on tendered products, with no further on-going additional costs by Respondent to Participating Health Services. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | all endpoint protection solutions for the tendered product must be listed in the Product Catalogue | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | endpoint protection event log should feed into Participating Health Service's cybersecurity systems (i.e., SIEM, event log repository) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | details of the endpoint protection event log must be listed in the Product Catalogue | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | endpoint protection event logs should include:
|  |  |  | • a log of valid and invalid logins and logoffs | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • all security type processes | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | ab |  | All communication ports must be disabled by default, and only the ports required in specific install base to be activated via a secure login by the Participating Health Service's IT Department. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | ac |  | Does tendered product have the capability to utilise multiple security certificates to hospital's IT network infrastructure? (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | ad |  | Tendered products must not have hardcoded or default passwords. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | af |  | Participating Health Services must be permitted to control the use of administrative level privileges on tendered products, if applicable. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | ag |  | Remote Access to any tendered product must:
|  |  | (i) | be through a secure environment, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | be time-limited; access should time out after 30 minutes, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | be encrypted | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • list encryption method in product catalogue, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | have unique user login credentials (both generic and group logins must not be used), and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | must not use single-factor user login authentication protocols, i.e., must use multi-factor secure user login authentication protocols, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (vi) | use the Participating Health Service's existing remote access gateway. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | ah |  | Tendered devices that connect onto Participating Health Service's IT network, should have UL 2900 certification | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | If so, date when certification was acquired for tendered device?
|  |  | (ii) | How regularly is tendered device evaluated to cyber penetration testing and security assessments?
| 11 |  |  | Transport
|  | a |  | Define device transportability & features, e.g. size, construction etc.
|  | b |  | Pumps intended for transport use (by Ambulance Victoria) must:
|  |  | (i) | be compliant with AS/NZS 4535, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | have a pole clamp mounts, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | fully accessorised must weigh less than 6kg, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | for road transport, must withstand a gravitational force of greater than 20G for forward and rearward force and 10G sideways force, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | for helicopter transport, must have CASA certification, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (vi) | for aeroplane transport, must have CASA certification. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (vii) | Preference may be given to pumps that are smaller in mass, both in weigh and size. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Pumps must continue to infuse, without altering dosage delivery, when used in transporting a patient. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference may be given to pumps that can be easily mounted with vertical or horizontal brackets to IV poles, beds, wheelchairs or other patient accessories. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Preference may be given to pumps that are easily removable from “rack mounting systems” but also are designed to avoid accidental release of pumps from their mount. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | Options
|  | a |  | Define which options are available for all categories:
|  |  | (i) | Wireless communication | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Data entry | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Carry handles | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | DERS | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13 |  |  | Accessories
|  | a |  | The below accessories are required across all categories:
|  |  | (i) | Bed Mount | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Pendant mounts | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Bar code reader | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 1 - Volumetric
Category 1: Volumetric Pumps
Clause Identifiers |  |  |  | Volumetric Pumps | (Enter model number) |  | (Enter model number) |  | (Enter model number)
Cat 1 |  |  |  | Volumetric Pumps | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Intended Purpose For Use
|  | a |  | A mains powered device designed to facilitate the accurate and consistent administration of drugs and solutions which can be delivered via intravenous, subcutaneous, arterial, epidural, and intravital routes using a dedicated infusion set. It is used to supply higher pressures than those provided by manually clamped gravity infusion sets or infusion controllers. It has internal batteries that enable operation for a short period when no mains electricity is available. | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Standards and Compliance
|  | a |  | Respondents must list all the Australian and International Standards which tendered product complies with in the Product Catalogue. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Respondents must list all EMC Standards which tendered product complies with in the Product Catalogue. | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Base Unit
|  | a |  | Must be able to infuse fluids from large rigid containers or flexible containers (0.5L to 2.0L). Containers may be of greater or lessor volume. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference maybe given to a pump that has the capability to utilise a syringe as a container and be directly connected to secondary port connector | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The pumps should be capable of infusing a wide variety of medical type fluids from crystalloid fluids to cytotoxic drugs, blood products, etc. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Any fluid that cannot be infused must be clearly indicated in Product Catalogue | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The cassette/giving set must be easy to load correctly into the pump. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The volumetric infusion pump must be able to provide the following flow rates:
|  |  | (i) | rates 0.1 to 99.9 mL/hr in increments of 0.1mL/hr | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | rates 1mL/hr to 999mL/hr in increments of 1mL/hr | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The pump must deliver and maintain accurate flow with any type of infusion fluid and with any appropriate infusion set intended for use with the device. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Flow error must be ±6% of indicated flows. | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The flow must remain accurate and vary <6% from the initial flow error after 72 hours of a continual set use under normal operating conditions. | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The pump must sense an upstream (fluid side) occlusion and should not require a drop/flow sensor to be attached to the intravenous (IV) set drip chamber for upstream occlusion detection. | --Select-- |  | --Select-- |  | --Select--
|  | h |  | The pump must alarm and terminate flow if >50 microlitres of air has entered the line. | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The pump must be capable of reverting to a keep vein open (KVO) rate, less than the programmed flow rate, after the programmed volume has been delivered. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | KVO rate must be programable at a health service level | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | State the range of programable KVO rates
|  | j |  | Preference maybe given to pumps that have the capability of an automated secondary (piggyback) infusion capability that automatically changes from a programmed secondary flow rate to the primary flow rate following the delivery of the secondary volume. | --Select-- |  | --Select-- |  | --Select--
|  | k |  | The interruption to flow, when changing flow setting, must not exceed 10 seconds in flow interruption. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference will be given where there is no interruptions to flow when changing flow setting. | --Select-- |  | --Select-- |  | --Select--
|  | l |  | The pump must have a two step IV set removal technique to initiate gravity flow and must prevent inadvertent free flow, i.e., the set cannot be removed from the pump unclamped. | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Power Supply
|  | a |  | Pump must operate from a standard 240V AC Australian power point. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference may be given to pumps that operate from standard mains cables and do not require external power supply plug packs. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | In the absence of mains power, pump's internal battery must be capable of operating for following times:
|  |  | (i) | for a single channel pump, at least 6 hours infusing at 25 mL/hr | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | for a dual channel pump, at least 4.5 hours infusing at 25 mL/hr | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | for multichannel pumps with all channels running, at least 3 hours infusing at 25 mL/hr | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Alarms
|  | a |  | Volumetric pumps must have the following alarms:
|  |  | (i) | Infusion Complete | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Air in Line | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Occlusion | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Plus all other alarms in the base unit | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Options
|  | a |  | The following are options that may be desired by the end users.
|  |  | (i) | Downstream continuous pressure monitoring | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Accessories
|  | a | (i) | Carry case | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Drop/flow sensor may be a preference | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Product Life
|  | a |  | State the warranty period for pump
|  | b |  | State battery warranty period
|  | c |  | State the expected lifespan of tendered product, in years
|  | d |  | State the expected lifespan of battery under normal conditions, in months

## 2 - Syringe
Category 2: Syringe Pumps
Clause Identifiers |  |  |  | Syringe Pumps | (Enter model number) |  | (Enter model number) |  | (Enter model number)
Cat 2 |  |  |  | Syringe Pumps | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Intended Purpose For Use
|  | a |  | A mains powered device designed to precisely drive the plunger of a syringe down its barrel to infuse a solution when it must be administered with a high degree of volume accuracy and rate consistency. Because of the lower flow settings and flow resolution (e.g., 0.1 ml/hr), it is especially appropriate for neonatal, infant, and critical care applications in which small volumes of concentrated drugs are to be delivered over an extended period. It can also be used to administer epidural analgesia. It has internal batteries that allow the device to operate for a short period of time when no mains electricity is available. | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Standards and Compliance
|  | a |  | Respondents must list in the overview tab of the TRW, all the Australian and International Standards which tendered product complies with. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Respondents must list all EMC Standards which tendered product complies with in the TRW. | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Base Unit
|  | a |  | The base unit must be able to infuse fluids from syringes ranging in size from 10mL to 50mL. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The pumps should be capable of infusing a wide variety of medical type fluids from crystalloid fluids to cytotoxic drugs, blood products etc. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Any fluid that leads to an impediment of flow delivery, or cannot be infused must clearly be stated in the Product Catalogue. | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The syringe must be easy to load correctly into the pump. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference will be given to syringe pumps that can use a multitude of manufacturer’s syringes. | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The syringe pump must be able to provide the following flow rates:
|  |  | (i) | for a 10mL syringe, maximum flow rate 250 mL/hr | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | for a 50mL syringe, maximum flow rate of 999 mL/hr | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | rates less than 10 mL/hr, must increment in 0.1mL/hr | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | rates greater than 10 mL/hr, must increment in 1mL/hr | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The pump must deliver and maintain accurate flow with any type of infusion fluid and with any appropriate infusion set intended for use with the device. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference may be given to pumps with a flow error of ±3% of indicated flows. | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The flow rate must remain accurate and vary less than 3% until the syringe is exhausted under normal operating conditions. | --Select-- |  | --Select-- |  | --Select--
|  | h |  | The pump must sense a downstream occlusion (flow restriction) | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | preference may be given to pumps where occlusion pressures can be set by the health service. | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The pump must be capable of delivering a bolus. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | bolus rate must be selectable by the health service and programmed into the syringe pump. | --Select-- |  | --Select-- |  | --Select--
|  | j |  | The pump must be capable of having a prime/purge feature. | --Select-- |  | --Select-- |  | --Select--
|  | k |  | It is desirable that there is no interruption when changing flow settings but shall not exceed 10 seconds in flow interruption. | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Power Supply
|  | a |  | Pump must operate from a standard 240V AC Australian power point. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference may be given to pumps that operate from standard mains cables and do not require external power supply plug packs. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | In the absence of mains power, pump's internal battery must be capable of operating for following times:
|  |  | (i) | for a single syringe pump, at least 5 hours infusing at 5 mL/hr | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | for multi-syringe pump, at least 2 hours infusing at 5 mL/hr on all syringes | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Alarms
|  | a |  | Syringe pumps must have the following additional alarms:
|  |  | (i) | Syringe dislodged | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Near empty reservoir | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Occlusion | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Plus all other alarms in the base unit | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Options
|  | a |  | The following are options that may be desired by the end users:
|  |  | (i) | Downstream continuous pressure monitoring | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | various colour clamp/door options | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Syringe covers. Successful Respondents must advise whether the syringe cover is lockable. | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Accessories
|  | a |  | Carry case | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Product Life
|  | a |  | State the warranty period for pump
|  | b |  | State battery warranty period
|  | c |  | State the expected lifespan of tendered product, in years
|  | d |  | State the expected lifespan of battery under normal conditions, in months

## 3 - Pump Platform
Category 3: Pump Platforms
Clause Identifiers |  |  |  | Pump Platforms | (Enter model number) |  | (Enter model number) |  | (Enter model number)
Cat 3 |  |  |  | Pump Platforms | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Intended Purpose For Use
|  | a |  | Products tendered within this category are ancillary products to the pumps listed in Category 1 - Large Volume Pumps and Category 2 - Syringe Pumps:
|  |  | (i) | Multi-Pump Rack - To facilitate the linking of numerous single channel infusion and/or syringe pumps onto a single footprint. All pumps function fully independently of each other. | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Docking Stations - To facilitate the connectivity of numerous single channel infusion and/or syringe pumps onto a single footprint, for automated charging of their internal rechargeable batteries and to include wired or wireless connection to other devices and/or communication (e.g. nurse call) systems or information systems facilitating the download of data. | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Pump Controllers - A mains powered device facilitating the connectivity and management of numerous single channel infusion and/or syringe modules to form a single user configurable system/device, providing an alarm (both visible and audible) when pre-set limits are exceeded and also includes wired or wireless connection to other devices and/or communication systems (e.g. nurse call) or information systems facilitating the download of data. | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Patient Parameter Monitoring (modules) - Some pump platforms offer additional patient monitoring options for the purpose of controlling the delivery of opioid medications. | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Standards and Compliance
|  | a |  | Respondents must list all Australian and International Standards which tendered product complies with. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Respondents must list all EMC Standards which tendered product complies with in the TRW. | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Base unit
|  | a |  | The base unit configuration for a Category 3 “Pump Platforms” may vary from manufacturer to manufacturer. The pump platform must be capable of joining a number of pumps from Categories 1-2 together to form a single device and/or system. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Pump Platform must be one of the following:
|  |  | (i) | Multi pump rack; with or without centralised power cord | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Multi-pump rack with power and communication, i.e. Docking Station | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Multi-pump rack with power, communication and common control, i.e. Controller ("brain") | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Patient parameter monitoring (module) components | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Size and Weight
|  | a |  | If the pump platform is designed to be mounted on wheels or castors it must be easily transportable and meet all OH&S manual handling requirements. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If the pump platform is mobile on wheels (Category 3.01 - 3.03), a braking feature must be provided. | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Transportation
|  | a |  | For pump platforms weighing more than 7 kg (including pump heads and options) it is desirable for a trolley/roll-stand option to be available. This trolley shall have wheels/castors that are minimum 5 cm in diameter. Preference may be given to larger wheels/castors. | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Power Supply
|  | a |  | All Category 3 Pump Platforms must have a UPS/battery system to provide a minimum of 120 minutes of support (if of an electronic type with control of the pumps). | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Alarms
|  | a |  | When the pump platform provides central alarm features the alarms should clearly identify which pump module is alarming and for what reason. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | It is desirable for either the pump platform or the pump module to be able to mute the alarm. | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Optional configurations
|  | a |  | It is desirable that pump platforms can include the following optional configurations upon request by Participating Health Services:
|  |  | (i) | touch screen | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | various display sizes | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | battery | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | pole mount | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | interface capabilities with other IT systems | --Select-- |  | --Select-- |  | --Select--
| 9 |  |  | Product Life
|  | a |  | State the warranty period for tendered product
|  | b |  | State battery warranty period
|  | c |  | State the expected lifespan of tendered product, in years
|  | d |  | State the expected lifespan of battery under normal conditions, in months
| 10 |  |  | Patient Parameter Monitoring
|  | a |  | Some infusion pump platforms offer additional options regarding monitoring patient parameters for the purpose of controlling the medication delivery. The following patient parameter monitoring options may be offered:
|  |  | (i) | EtCO2 monitoring | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | O2 monitoring | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | SpO2 monitoring | --Select-- |  | --Select-- |  | --Select--
note: if more patient parameter monitoring options are offered, contact HPV for additional specifications
|  | b |  | Consumables for patient monitoring must be listed under:
|  |  | (i) | HPV Monitoring Products Contract HPVC2016-021 | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | HPV Respiratory Products Contract HPVC2020-015 | --Select-- |  | --Select-- |  | --Select--
| 11 |  |  | EtCO2 Monitoring
|  | a |  | EtCO2 numerical data must be displayed on screen and captured to the event memory. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | User selectable on-screen waveforms must be displayed. | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respiratory range must be 3–60 breaths per minute. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Respiratory rate accuracy must be ± two (2) breaths per minute (bpm) below 60 bpm, ± 3 bpm between 60 and 100 bpm and ±5 above 100 bpm. | --Select-- |  | --Select-- |  | --Select--
|  | e |  | EtCO2 range must be 0–99mmHg in 1mmHg increments, or 0–10% in 0.2% increments. | --Select-- |  | --Select-- |  | --Select--
|  | f |  | EtCO2 accuracy must be ± 2mmHg for range 0–40mmHg; ± 5mmHg for range 41–70mmHg; ± 8mmHg for range 71–100mmHg | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | O2 Monitoring
|  | a |  | O2 monitoring may be offered as a separate option. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | respiratory range must be 2–60 breaths per minute | --Select-- |  | --Select-- |  | --Select--
|  | c |  | respiratory rate accuracy must be ± two (2) breaths per minute | --Select-- |  | --Select-- |  | --Select--
|  | d |  | the warm-up time of the gas monitor/module should be less than five (5) minutes | --Select-- |  | --Select-- |  | --Select--
|  | e |  | self-calibration is preferred. | --Select-- |  | --Select-- |  | --Select--
| 13 |  |  | Pulse Oximetry (SpO2)
|  | a |  | SpO2 data must be captured to the event memory. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | SpO2 accuracy must be within the following ranges:
|  |  | (i) | ± 3% at 70–100% SpO2 | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | ± 4% at 50–70% SpO2. (if capable) | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Pulse rate range must be 30–240 beats per minute. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Pulse rate accuracy should be within ± three (3) beats per minute. | --Select-- |  | --Select-- |  | --Select--

## 4 - Specific Pumps
Category 4: Specific Purpose Pumps
Clause Identifiers |  |  |  | Specific Purpose Pumps | (Enter model number) |  | (Enter model number) |  | (Enter model number)
Cat 4 |  |  |  | Specific Purpose Pumps | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Intended Purpose For Use
|  | a |  | Pumps tendered within this category have a more specific function:
|  |  | (i) | Ambulatory Infusion Pump - A portable, battery-powered, electronic device designed to provide continuous or intermittent delivery of drugs and solutions to patients. It is typically worn in a case attached to the belt or carried in a pocket and includes a reservoir (e.g., a cartridge similar to a syringe), a pumping mechanism, and an infusion-rate controller. The pump infuses periodic microvolume pulses and bolus doses. | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Epidural Infusion Pump - A mains or battery powered infusion device designed to deliver analgesics typically through epidural routes at pre-established accurate flow rates for pain relief. Pumps may use one of a variety of different technologies for propelling the infuscate, including peristaltic, rotary, cassette, syringe-driven, and elastomeric mechanisms. These pumps are capable of delivering large and/or small volumes of liquids through a single channel. Infusion pumps may be intended for mounting on poles or for ambulatory use. | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Patient-Controlled Analgesic Pumps - A mains or battery powered infusion device designed to deliver intermittent bolus doses of opioid medications intravenously (e.g., Fentanyl, Demerol (meperidine), Morphine, or Hydromorphone) upon patient demand (i.e., when activated by the patient handset when they require pain relief), or in a slow, continuous mode for regional anaesthesia. | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Patient-Controlled Epidural Analgesic Pumps - A mains or battery powered infusion device designed to deliver intermittent bolus doses of analgesics upon patient demand (when activated by the patient when he/she requires pain relief), or in a slow, continuous mode for regional anaesthesia. A patient-controlled epidural analgesia (PCEA) infusion pump is used to infuse analgesics into the epidural space for pain management. | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | Total Intravenous Anaesthesia (TIVA) Pumps - A mains powered syringe pump designed for only in administering anaesthesia, usually through intravenous (IV) routes at pre-established accurate flow rates. These pumps usually consist of a syringe, a motor that drives a screw or gear mechanism to advance either the plunger or barrel of the syringe at a predetermined rate, electronic controls used to select the dose or volume to be infused, and audible and/or visual alarms to indicate abnormal conditions. In most pumps, the syringe containing the infuscate must be manually replaced once it has emptied in order to continue the infusion. It controls the infusion pre-programmed according to the patient's (target) weight, height, age and kind of drug being delivered; which is commonly known as the target controlled infusion (TCI). Anaesthesia syringe pumps are used mainly in surgical and special procedures (e.g., catheterization) areas in hospitals and other healthcare facilities. | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Standards and Compliance
|  | a |  | Respondents must list all Australian and International Standards which tendered product complies with. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Respondents must list all EMC Standards which tendered product complies with in the Product Catalogue | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Base Unit
|  | a |  | The base unit must be able to infuse a wide variety of medical type fluids from crystalloid fluids to cytotoxic drugs, blood products etc. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Any fluid that leads to an impediment of flow delivery, or cannot be infused must clearly be stated in the Product Catalogue | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The pump must comply with Category 1 or Category 2 basic specification as outlined in tab "1 - Volumetric" or tab "2 - Syringe" (note devices in this category are NOT limited to syringe pump style of pumps). Respondents must confirm compliance to specification by listing tendered product in applicable tab. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | state if volumetric or syringe type pump
|  |  | (ii) | state if pump intended for ambulatory use
|  | c |  | The base unit must be capable of delivering a bolus. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference maybe given to pumps capable of delivering a bolus dose and/or bolus rate duration. | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Bolus feature must be able to be locked out by the health care worker preferably by the use of a security method (such as password or key). | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The fluid reservoir must have an optional lockable case. | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The pump must have a full and comprehensive log of drug delivery including time and volume. | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Preference maybe given to “on board dose calculators” | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Any data required must be in metrics units of measurement. | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Proprietary required accessories (including drug reservoirs) shall be clearly indicated on the unit and within the response to this ITS. | --Select-- |  | --Select-- |  | --Select--
|  | i |  | If particular algorithms are in use preference may be given to those devices that clearly indicate the algorithm on the pump | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Ambulatory Use
|  | a |  | Tendered product intended for ambulatory use. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Preference will be given to ambulatory use pumps that weigh less than 1 kg and are as small as possible. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Pumps tendered intended for ambulatory use need not comply with "Power Supply" requirements in Category 1 or Category 2 basic specification as outlined in tab "1 - Volumetric" and tab "2 - Syringe" | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Pumps intended for ambulatory use must be handheld devices. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Pumps intended for ambulatory use must be impact resistant to a level of 1 meter. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Respondents must provide evidence of drop test results compliant to an appropriate standard. | --Select-- | <<State Standard>> | --Select-- | <<State Standard>> | --Select-- | <<State Standard>>
|  | e |  | Pumps intended for ambulatory use must have a carry case option available for purchase. | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Alarms
|  | a |  | Dedicated pumps must have the following alarms:
|  |  | (i) | Giving set dislodged | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Dose delivery | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Dose request | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Near empty | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Options
|  | a |  | The following are options that may be desired by the end users:
|  |  | (i) | Patient controlled switch | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Downstream pressure monitor | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Rechargeable battery | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | Battery charger | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Accessories
|  | a | (i) | Carry case | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Disposable carry cases | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Product Life
|  | a |  | State the warranty period for pump
|  | b |  | State battery warranty period
|  | c |  | State the expected lifespan of tendered product, in years
|  | d |  | State the expected lifespan of battery under normal conditional use, in months
4.01 |  |  |  | Ambulatory Infusion Pumps
| 1 |  |  | Base Unit
|  | a |  | Syringe type pumps must be able to infuse fluids from syringes ranging in size from:
|  |  | (i) | pump accommodates a 10mL syringe | --Select-- |  | --Select-- |  | --Select--
|  |  |  | pump administers the full 10mL | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | pump accommodates a 20mL syringe | --Select-- |  | --Select-- |  | --Select--
|  |  |  | pump administers the full 20mL | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | pump accommodates a 30mL syringe | --Select-- |  | --Select-- |  | --Select--
|  |  |  | pump administers the full 30mL | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | pump accommodates a 50mL syringe | --Select-- |  | --Select-- |  | --Select--
|  |  |  | pump administers the full 50mL | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | syringes must be easy to load correctly into pump. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The fluid reservoir should have a lockable case. | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Power Supply
|  | a |  | Batteries must be a standard type, i.e. “AAA”, “AA”, “C” or 9 volt for Ambulatory pumps that use disposable batteries. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Ambulatory pump should run infusion for 24 hours or more, on a single disposable battery, or on a single charge cycle of a reusable battery. | --Select-- |  | --Select-- |  | --Select--
|  | c |  | In Product Catalogue specify the battery run time for:
|  |  | (i) | full charged rechargeable battery | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | disposable battery | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | model number of recommended disposable battery | --Select-- |  | --Select-- |  | --Select--
4.02 |  |  |  | Epidural Infusion Pumps
| 1 |  |  | Base Unit
|  | a |  | Epidural infusion pumps should have a yellow case | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The fluid reservoir must have a lockable case. | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Power Supply
|  | a |  | Epidural pumps intended for ambulatory use must be capable of running infusion for 24 hours or more on battery | --Select-- |  | --Select-- |  | --Select--
4.03 |  |  |  | Patient-Controlled Analgesic Pumps
| 1 |  |  | Base Unit
|  | a |  | Patient controlled pumps must be able to provide the flow rates mentioned in Category 1 or 2 basic specification as outlined in tab "1 - Volumetric" and tab "2 - Syringe"; unless clearly indicated otherwise. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Patient controlled pumps must be supplied with a patient control switch | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The fluid reservoir must have a lockable case. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Pump must have programable accumulated dose limit feature to prevent patient overdose | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Power Supply
|  | a |  | PCA pumps intended for ambulatory use must be capable of running infusion for 24 hours or more on battery | --Select-- |  | --Select-- |  | --Select--
4.04 |  |  |  | Patient-Controlled Epidural Analgesic Pumps
| 1 |  |  | Base Unit
|  | a |  | PCEA pumps should have a yellow case | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Patient controlled pumps must be able to provide the flow rates mentioned in Category 1 or 2 basic specification as outlined in tab "1 - Volumetric" and tab "2 - Syringe"; unless clearly indicated otherwise. | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Patient controlled pumps must be supplied with a patient control switch | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The fluid reservoir must have a lockable case. | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Pump must have programable accumulated dose limit feature to prevent patient overdose | --Select-- |  | --Select-- |  | --Select--
|  | f |  | PCEA pump should be programable with following modes:
|  |  | (i) | continuous infusion | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | patient-controlled infusion | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | programmed intermittent bolus infusion | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Power Supply
|  | a |  | PCEA pumps intended for ambulatory use must be capable of running infusion for 24 hours or more on battery | --Select-- |  | --Select-- |  | --Select--
4.05 |  |  |  | Total Intravenous Anaesthesia (TIVA) Pumps
| 1 |  |  | Base Unit
|  | a |  | TIVA pumps must be supplied with Health Service configurable TCI algorithms | --Select-- |  | --Select-- |  | --Select--

## 5 - DERS Software
Category 5: Dose Error Reduction Systems
Clause Identifiers |  |  |  | DERS | (Enter model number) |  | (Enter model number) |  | (Enter model number)
Cat 5 |  |  |  | Dose Error Reduction Systems | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Intended Purpose For Use
|  | a |  | Dose error reduction system (DERS) software is intended to improve the safety of infusion procedures by warning users of incorrect medication orders, calculation errors, and misses in programming. The software is designed to reduce infusion pump user programming errors, during the administration of medications by infusion pumps in using drug dose libraries and standard limits for medications, electrolytes, and other infusions. The software for future expansion includes interfaces to electronic medical records (eMR), and/or other hospital information systems. | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Standards and Compliance
|  | a |  | Respondents must list all Australian and International Standards which tendered product complies with. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | Respondents must list all EMC Standards which tendered product complies with in the Product Catalogue. | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Base unit
|  | a |  | The base unit configuration for a Dose Error Reduction System (DERS) is the software required to manage DERS settings on 'infusion' pumps. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The software must be able to run on standard computers/servers such as Microsoft based systems. | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The software installation and configuration is to be included. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The DERS must have the following features:
|  |  | (i) | a backup copy disconnected from the 'live' DERS, stored in such a manner that file is incorruptible (for recovery from cyber incident), and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | capability to detect and prevent dose related programming errors and programming of inappropriate or unreasonable infusions, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | be compatible for both Syringe, Patient Controlled, Dedicated Pumps and Volumetric Infusion Pumps, where offered, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | the software must have the ability to require the user to indicate if the pump is being used on a new patient, or confirm, if not a new patient, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | the software must have the ability to require the user to select the clinical location, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (vi) | the system must be simple to operate and require minimal steps or keystrokes to program, compared to conventional rate and VTBI programming | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · however, extra steps must be required in order to disable or enable limits, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (vii) | the pump must be able to be placed in standby mode to allow clinicians to program the pump and the pump must retain programmed settings until infusion commenced, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (viii) | the system must be able to log and analyse all pump activities and events and data must not be able to be modified by the user, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ix) | the system must allow the following pump settings:
|  |  |  | · hard and soft dose limits for each drug, i.e. hard limits that cannot be overridden, and soft limits that can be overridden with acknowledgement of the user; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · limits to be set for drug concentrations | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · the patient’s weight | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · VTBI; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · flow setting; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · bolus delivery, with hands-free option for each individual drug; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · alarm volume levels; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · occlusion and air-in-line threshold values for each location and therapy; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · total dose and time limits for delivering boluses and intermittent infusions; and | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · pre-defined volume for a named drug infusion. | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The DERS should have the following features:
|  |  | (i) | perform integrity check on DERS file; or | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | a "real-time" dashboard indicating:
|  |  |  | · % compliance with drug library vs. basic infusion, or | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · infusion summaries, or | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · pump status, or | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · pump location | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Preference may be given to systems:
|  |  | (i) | where DERS is activated on a pump upon power-up, i.e., pump operates DERS by default, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | where secondary infusions are protected by DERS; or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | that can facilitate multiple entries of a medication, dosed in various ways, with different limits appropriate for each setting; or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | that can provide Body Surface Area based dosing; or | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | that can provide weight based dosing; or | --Select-- |  | --Select-- |  | --Select--
|  |  | (vi) | that provide pump location and status details on "real-time" dashboard. | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Alarms
|  | a |  | Infusion systems must have the following audible alarms tones when:
|  |  | (i) | during the operator programming cycle a soft limit is exceeded, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | during the operator programming cycle a hard limit is exceeded. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A clear WARNING indication must be displayed on the pump when:
|  |  | (i) | any time the DERS is not in use, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | soft limits have been overridden, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | the limit has been overridden, if part of a scrolling display, the status shall be displayed every few seconds, so that can be quickly recognised by users, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | the same drug or solution is programmed on more than one channel, on a multichannel pump. | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Drug Library
|  | a |  | State maximum number of characters pump can display on screen for drug names.
|  |  | (i) | character limits must allow normal naming convention of medications and doses to be displayed on pump | --Select-- |  | --Select-- |  | --Select--
|  | b |  | State the maximum number of drug entries drug library can hold.
|  |  | (i) | state the number of varied configurations for a single drug entry/entity
|  | c |  | State the maximum number of care areas/profiles a drug library can hold.
|  | d |  | State the maximum number of drug entries each care area/profile can hold.
|  | e |  | State the number of drug libraries a single DERS can support.
|  | f |  | State the maximum number of pumps a single DERS can support.
|  | g |  | The system must include a drug library with the following features:
|  |  | (i) | the drug library shall be housed on a single central server for all the pumps that allows management of the drug libraries in the pumps, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | the drug library should be able to be configured to a Health Service’s current ordering and delivery practices, i.e. provide for customisation of drug names, concentrations and dosing units, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | navigation through drug libraries must permit easy, quick, and accurate selection of drug entities. | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Preference will be given to a drug library that:
|  |  | (i) | is easy to review, e.g. a database that can be printed in spreadsheet form and allow batch programming of the drug library for customisation, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | has advisory information about the drug that is being used which the user must acknowledge before continuing, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | allows a Participating Heath Service to specify entire drug names or dosing units (i.e. micrograms or units) without abbreviations, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | allows the use of TALLman lettering to differentiate drugs with similar spelling, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | has sufficient capacity for clinical profiles and drugs to meet the needs of Participating Health Services, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (vi) | prevents primary infusion from being interrupted when secondary infusion being administered. | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The following instructions must be provided to the Participating Health Services:
|  |  | (i) | how drug searching is undertaken, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | the number of clinical ‘locations’ with drug libraries, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | the number of drugs or drug concentrations per library, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | the number of decimal places used in dosage calculations. | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Preference may be given to systems that have the following features:
|  |  | (i) | pumps that have an event log that tracks all programming changes, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | pumps that have an event log that track limit overrides when a programming change is made consequent to a limit warning message being displayed; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · this log be separate from the pumps primary data log, and be accessible via the database server, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | data log analysis software to permit downloading, storing and convenient, e.g. graphical or spreadsheet, analysis of the data aggregated from all pumps, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | a log that retains data for at least 24 hours or the last 1000 incidents. | --Select-- |  | --Select-- |  | --Select--
|  | k |  | Preference may be given to systems that have the following communication and data transfer features:
|  |  | (i) | pumps that are able to receive and send information while in use without disrupting infusion, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | however new information may not become active on the pump unless confirmed by the user on the pump, e.g. at ‘start-up’ when the user is requested to accept the new library. This is to enable users to check which pumps have the latest version, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | the pump is to be able to display the last time the pump accepted a file update, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | the pump data parameters being able to be uploaded from a portable computer, or | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | be able to identify which pumps have received recent updates. | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Implementation Support in developing the drug libraries is essential please indicate the following:
|  |  | (i) | details of their experience, preferably providing examples of hospital projects in Australia, of the following activities:
|  |  |  | · Coordination with the hospital on the drug libraries content, and | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Development of drug libraries, and | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Development of dose limits, and | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Development of location and application specific parameters i.e. maximum infusion rates, patient parameters, VTBI etc., and | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Training of clinicians on the operation of the DERS, and | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Typical implementation costs, and | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Typical time frames for the implementation of the DERS system with libraries. | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Optional configurations
|  | a |  | It is desirable that DERS can include the following optional configurations upon request by Participating Health Services:
|  |  | (i) | the ability to integrate with other Hospital Information Systems such as | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | · Patient Demographic information (UR, age, etc) | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Patient records | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · “e-Scribe” systems | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Drug dispensing systems | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · Supply systems | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Hardware | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | various display sizes with touch screen | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | analysis of pump data for quality purposes. This could be an annual cost or cost per hour. | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Data management
|  | a |  | Data Recording | --Select-- |  | --Select-- |  | --Select--
|  | b |  | DERS with optional data recording must incorporate date and time recording of patient and delivery settings (when available). | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Additional configurable parameters are desirable. | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Event Reporting
|  | a |  | All DERS with data transmission capabilities must have the capability to pass event data wirelessly, via cable or via a memory device (such as a memory card) to other devices. | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A summary report should be downloadable from the DERS in a universal format so that it may become a component of patient records. | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Event information should be able to be stored on the host Participating Health Services’ information systems. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | For research and development purposes, data format should be open ECG-SCP, RCG or similar Open Same Standard format. If only a proprietary format is available, a software development kit for Microsoft Windows should be provided as a purchasable option. | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Patient data that is stored on the internal memory should be encrypted to an open standard (i.e. non-proprietary). | --Select-- |  | --Select-- |  | --Select--
|  | f |  | A data copy must be able to be stored in such a manner that it is incorruptible for use in any potential legal applications. | --Select-- |  | --Select-- |  | --Select--
EHR Interoperability
|  | g |  | Infusion device must have wireless bi-directional interoperability with EHR | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Clinical workflows associated with administering an infusion should be integrated from infusion device to the EHR | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Preference may be given by Participating Health Services to Respondents with demonstrated evidence for data integration at live sites | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | List sites where system has been integrated with EHR
|  |  | (ii) | List any third-party organisations that have recognised EHR integration (i.e., KLAS, ECRI, etc.)
|  | j |  | Preference may be given by Participating Health Services to Respondents that have evidence to show compliance for interoperability usage of 95% or greater | --Select-- |  | --Select-- |  | --Select--
| 9 |  |  | Network
|  | a |  | Tendered pumps with integrated DERS must have wireless capabilities | --Select-- |  | --Select-- |  | --Select--
IT Security
|  | b |  | DERS must have:
|  |  | (i) | unique user login credentials (both generic and group logins must not be used), and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | user level of privileges managed, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | must not use single-factor user login authentication protocols, i.e., must use multi-factor secure user login authentication protocols, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | authentication and event logs. | --Select-- |  | --Select-- |  | --Select--
|  | c |  | DERS must have patch update installed within 90 days upon app vulnerability having been identified at no additional cost to Participating Health Services. | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Where a vulnerability in DERS is being actively exploited by threat actors, the Successful Respondents must provide Participating Health Services: | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | advice on implementing compensating control measures to mitigate vulnerability being exploited, and | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | a patch update within 48 hours to remediate the vulnerability. | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Health Service Endpoint protection solutions (anti-malware, application control, whitelisting functionality, etc.) must be permitted to be enabled and/or installed on tendered products, with no further on-going additional costs by Respondent to Participating Health Services. | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Successful Respondent must permit Participating Health Service have a qualified 3rd-party provider perform a full security assessment and threat-scan assessment of DERS virtual server environment location. | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Does tendered product have the capability to utilise multiple security certificates to hospital's IT network infrastructure? (yes/no) | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | what types and what are they used for?
|  |  | (ii) | does DERS server manage the security certificates? | --Select-- |  | --Select-- |  | --Select--
|  | h |  | DERS should have UL 2900 certification | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | If so, what year was certification first acquired?
|  | i |  | How regularly is tendered device evaluated to cyber penetration testing and security assessments?
|  | j |  | What other cybersecurity certification has the DERS platform been certified to?
|  |  | (i) | What year was certification acquired?
Hardware
|  | k |  | If hardware is provided the mains switch, if present, must be protected from unintended activation or deactivation.
|  | l |  | Switching ‘off’ the hardware during normal operation should require a confirmatory step and/or the machine should display a warning of imminent shut-down.
|  | m |  | All category 5 hardware should have a UPS/battery system to provide a minimum of 10 minutes of support. This is to deal with temporary unintended power interruptions and facilitate an orderly shutdown.
|  | n |  | Preference will be given to UPS/battery systems that have “Hot swappable” batteries.
|  | o |  | A display if provided should be of sufficient quality to easily display the data required.
|  | p |  | Non-industry standard cables/dongles shall be included with the supply of the system.
Data Transmission
|  | q |  | If data transmission is available, DERS must be able to transfer data via one of the following industry standard connections:
|  |  | (i) | RJ45 (Ethernet) | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | Wi-Fi | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | Zigbee | --Select-- |  | --Select-- |  | --Select--
|  |  | (iv) | LoRa | --Select-- |  | --Select-- |  | --Select--
|  |  | (v) | if other, please specify
|  | r |  | Data transmission via Bluetooth (of any version) is not permitted | --Select-- |  | --Select-- |  | --Select--
|  | s |  | Data transmission must be encrypted for privacy requirements. | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | State the cryptographic method used.
|  | t |  | All communication ports must be disabled by default, and only the ports required in specific install base to be activated via a secure login by the Participating Health Service's IT Department. | --Select-- |  | --Select-- |  | --Select--
|  | u |  | Data transmission to other systems is preferred in standardised languages, such as Health Level Seven (HL7). | --Select-- |  | --Select-- |  | --Select--
|  |  | (i) | list languages and versions data is transmitted
|  | v |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free pre-purchase. Details of all conditions and subsequent annual costs must also be provided pre-purchase. | --Select-- |  | --Select-- |  | --Select--
|  | w |  | Describe the method(s) to update DERS settings to pumps
Wireless Environment
|  | x |  | Where pumps have wireless communication capabilities please indicate the following:
|  |  | (i) | availability of bi-directional wireless communication for:
|  |  |  | · transfer of new drug libraries and software updates from the central server to the; | --Select-- |  | --Select-- |  | --Select--
|  |  |  | · transfer of event and alarm logs from the pumps to the server. | --Select-- |  | --Select-- |  | --Select--
|  |  | (ii) | is there a wired communication link available in the event there's a problem with the wireless network; and | --Select-- |  | --Select-- |  | --Select--
|  |  | (iii) | state the wireless communication type utilised (where appropriate) for current IT infrastructure, and
|  |  | (iv) | state the measured average network traffic bandwidth for 600 active wireless LVPs.
| 10 |  |  | Product Life
|  | a |  | State the warranty period for tendered software product/s
|  | b |  | State the warranty period for tendered hardware product/s
|  | c |  | State the expected lifespan of tendered software product/s
|  | d |  | State the expected lifespan of tendered hardware product/s

## 6 - IV Admin Sets
Category 6: IV Administration Sets
Clause Identifiers |  |  |  | Text
Cat 6 |  |  |  | IV Administration Sets | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Intended Purpose
|  | a |  | Intravenous administration sets designed for liquid administration from a bag or bottle through an infusion pump either to a hub catheter or needle at the patient site, or to an extension tube set attached to the patient site. These sets typically consist of plastic tubing, a check valve, a roller clamp, a Y-site connector for secondary tubing attachment, and luer adapters at both ends | --Select--
|  | b |  | Intravenous administration sets are considered as a consumable | --Select--
| 2 |  |  | Base Unit
|  | a |  | Tendered product must be a proprietary (consumable) administration set for exclusive use by a pump that has also been tendered to this contract. Tenderer to state compatible pump/s in the Tender Response Worksheet. | --Select--
|  | b |  | Tendered product must be compatible with only one manufacturer's pumps. | --Select--
|  | c |  | Tendered products that are already subject to another Health Purchasing Victoria contract must not be tendered in this category/contract. | --Select--
|  | d |  | All IV administration sets:
|  |  | (i) | must be a single-use product; | --Select--
|  |  | (ii) | must have luer lock connectors, i.e., needleless ports; | --Select--
|  |  | (iv) | must be available in vented and non-vented or multipurpose configurations; | --Select--
|  |  | (v) | must not kink, restricting flow, or burst, under normal conditions of operation; | --Select--
|  |  | (vi) | must have a low tubing priming volume of less than 15 mL: Tenderer to state priming volume in the Tender Response Worksheet; | --Select--
|  |  | (vii) | the set's dead-space volume must be labelled on the package; and | --Select--
|  |  | (viii) | for micro-infusion (where available), must utilise microbore tubing. | --Select--
|  | e |  | For each administration set offered, Respondents must complete the TRW tab/pricing schedule extract marked as Infusion sets, indicating:
|  |  | (i) | length of tubing, cm | --Select--
|  |  | (ii) | whether if infusion set have access port (yes/no) and distance from patient, cm | --Select--
|  |  | (iii) | whether if Infusion set have side arms | --Select--
|  |  | (iv) | whether if Infusion set have non-return valves | --Select--
|  |  | (v) | whether if Infusion set is colour tinted | --Select--
|  |  | (vi) | whether if Infusion set has filter (yes/no) and filter pore size, μm | --Select--
|  |  | (vii) | whether if Infusion set is vented | --Select--
|  |  | (viii) | whether if Infusion set is PVC free | --Select--
|  |  | (ix) | whether if Infusion set is DEHP free | --Select--
|  |  | (x) | whether if Infusion set is lipid compatible | --Select--
|  |  | (xi) | whether if Infusion set has any drug incompatibilities, list drugs | --Select--
|  |  | (xii) | whether if Infusion set is light sensitive | --Select--
|  |  | (xiii) | whether if Infusion set is MRI conditional (ferrous free) | --Select--
|  | f |  | Respondents are to indicate if the light Sensitive IV tubing is validated for use with Phototherapy lights with a wavelength of 430-490nm and a peak wavelength of 458nm
| 3 |  |  | Volumetric Infusion Pump Sets
|  | a |  | Successful respondent must offer a full range of sterile, single use, IV administration (giving) sets and accessories required for use with tendered volumetric infusion pumps. | --Select--
|  | b |  | Infusion Pump administration sets must have inherent free flow protection built into the IV set. | --Select--
|  | c |  | All secondary (piggyback) infusion pump sets:
|  |  | (i) | must be at least 45cm long and be equipped with a manual roller clamp and needleless injection system; and | --Select--
|  |  | (ii) | must be available in vented and non-vented or multipurpose configurations. | --Select--
|  | d |  | Preference may be given to infusion pump sets that:
|  |  | (i) | allowing pump to prime line | --Select--
|  |  | (ii) | allow gravity-flow priming; | --Select--
|  |  | (iii) | allow manual regulation of gravity flow using a high quality roller clamp; | --Select--
|  |  | (iv) | have access ports above and below the infusion device; | --Select--
|  |  | (v) | have anti-siphon / reflux / check valve built into the IV set; | --Select--
|  |  | (vi) | have low sorbing properties required for chemotherapy | --Select--
|  |  | (vii) | has clear labelling of contents; and | --Select--
|  |  | (viii) | Respondents must provide details of these features in the Tender Response Worksheet | --Select--
| 4 |  |  | Syringe / PCA / TIVA Pump Sets
|  | a |  | The syringe pumps are expected to be able to be used with generic syringes and infusion lines covered by Intravenous Access Devices and Administration Consumables HPV 2013-079 and its successive contracts. | --Select--
|  | b |  | Tendered syringe pump administration sets must be a proprietary consumable, designed to enable an additional pump feature other than administering fluid. | --Select--
|  | c |  | Tendered product is intended for use with either:
|  |  | (i) | syringe pump; | --Select--
|  |  | (ii) | patient controlled analgesia pump; or | --Select--
|  |  | (iii) | total intravenous anaesthesia pump. | --Select--
|  | d |  | PCA and Ketamine giving sets must not have any ports. | --Select--
|  | e |  | PCA and Ketamine giving sets should have anti reflux valves | --Select--
| 5 |  |  | Ambulatory Infusion Pump Sets
|  | a |  | Successful respondent should offer a complete range of sterile, single use, IV administration (giving) sets and accessories required for use with tendered Ambulatory Infusion Pumps. | --Select--
|  | b |  | Ambulatory pump administration sets should be compatible with a range of container options, i.e. syringe, plastic pouch, etc. | --Select--
| 6 |  |  | Epidural / PCEA Pump Sets
|  | a |  | Successful respondent should offer a full range of sterile, single use, neuraxial or regional anaesthesia giving sets and accessories required for use with tendered: | --Select--
|  |  | (i) | epidural pump; or | --Select--
|  |  | (ii) | patient controlled epidural analgesia pump | --Select--
|  | b |  | Epidural giving sets should be compatible with a range of container options, i.e. syringe, plastic pouch, etc. | --Select--
|  | c |  | Epidural giving set tubing should be identified with yellow | --Select--
|  | d |  | Preference maybe given to epidural giving set with neuraxial small-bore connectors (i.e., ISO 80369-6 compliant) on the patient-end connection. | --Select--
|  |  | (i) | HSV intends on running a supplementary tender in future to ensure all epidural giving sets comply with ISO 80369-6. | --Select--

## 7 - Semi Consumable & Accessory
Category 7: Semi-Consumables & Accessories
Clause Identifiers |  |  | Text
Cat 7 |  |  | Semi-Consumables & Accessories | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  | Intended Purpose
|  | a | A semi-consumable is a part that is used multiple times but is not expected to last the entire effective life of the equipment, such as batteries, strip recorder paper, etc. All user-replaceable parts are considered as semi-consumables. | --Select--
|  | b | An accessory is a supplementary part that enhances but does not inhibit the operations of the device, e.g. lockbox, carry bag, barcode reader, etc. | --Select--
| 2 |  | Volumetric Infusion Pumps
|  | a | Respondents must ensure that all semi-consumables and accessories required for use with infusion pumps in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 3 |  | Syringe Pumps
|  | a | Respondents must ensure that all semi-consumables and accessories required for use with syringe pumps in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 4 |  | Ambulatory Infusion Pumps
|  | a | Respondents must ensure that all semi-consumables and accessories required for use with ambulatory infusion pumps in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 5 |  | Patient-Controlled Analgesic Pumps
|  | a | Respondents must ensure that all semi-consumables and accessories required for use with patient controlled analgesic pumps in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 6 |  | Epidural Pumps
|  | a | Respondents must ensure that all semi-consumables and accessories required for use with epidural pumps in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 7 |  | Patient-Controlled Epidural Analgesic Pumps
|  | a | Respondents must ensure that all semi-consumables and accessories required for use with patient controlled epidural analgesic pumps in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 8 |  | Total Intravenous Anaesthesia Pumps
|  | a | Respondents must ensure that all semi-consumables and accessories required for use with total intravenous anaesthesia pumps in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 9 |  | Racks / Docks / Controllers
|  | a | Respondents must ensure that all racks, docks and controllers in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 10 |  | Monitored Patient Parameter
|  | a | Respondents must ensure that all monitored patient parameters in Tender Response Worksheet are listed and new category and sub-category selected. Any new semi-consumable or accessory must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--

## 8 - Options
Category 8: Options
Clause Identifiers |  |  | Text
Cat 8 |  |  | Options | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  | Intended Purpose
|  | a | Options are defined as after sales upgrades, i.e. options are additional features that can be added to or included onto tendered product (pump) that were not included in the original purchase. | --Select--
| 2 |  | OPTIONS
|  | a | Respondents must ensure that all after sales enhancements in Tender Response Worksheet are listed. Ensure current lines include pricing, once a health service has purchased contracted product Any newly added option must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--

## 9 - Spare Parts
Category 9: Spare Parts
Clause Identifiers |  |  | Text
Cat 9 |  |  | Spare Parts | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  | Intended Purpose
|  | a | Spare part is any part for which a tool is required for it to be fitted to a faulty or broken pump during the repair process. | --Select--
|  | b | Respondent must supply any spare part/s and/or maintenance kit/s for tendered product when requested by Participating Health Service | --Select--
|  | c | Maintenance kit is a kit of spare parts required for routine periodic maintenance | --Select--
| 2 |  | Testers / Calibration Kits / Tools / Cables / Dongles / etc.
|  | a | Respondents must include a list of testers, calibration kits, tools, cables, dongles, etc. in Tender Response Worksheet that are required to fully repair and maintain tendered product. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 3 |  | Volumetric Infusion Pumps
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered volumetric infusion pump over the tendered product's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 4 |  | Syringe Pumps
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered syringe pump over the tendered product's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 5 |  | Ambulatory Infusion Pumps
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered ambulatory infusion pump over the tendered product's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 6 |  | Patient-Controlled Analgesic Pumps
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered patient-controlled analgesic pump over the tendered product's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 7 |  | Epidural Pumps
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered epidural pump over the tendered product's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 8 |  | Patient-Controlled Epidural Analgesic Pumps
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered patient-controlled epidural analgesic pump over the tendered product's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 9 |  | Total Intravenous Anaesthesia Pumps
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered TIVA pump over the tendered product's supported lifespan. Any new spare parts being added must be highlighted and listed at the bottom of the pricing schedule extract and relevant new category/ sub-category selected. | --Select--
| 10 |  | Racks / Docks / Controllers
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain tendered racks / docks / controllers over the tendered product's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--
| 11 |  | Monitored Patient Parameter
|  | a | Respondents must list all replacement parts, assemblies and maintenance kits in Tender Response Worksheet that are required to fully repair and maintain the tendered product over it's supported lifespan. Any new spare part must be processed via the standard HPV contract variation process through the help desk. Noting this must be done post tender close against the new Specification and new category/ sub-category. | --Select--

## 10 - Agreements
Category 10: Agreements
Clause Identifiers |  |  |  | Text
Cat 10 |  |  |  | Agreements | Fully comply/ Partially comply/ Does not comply | Please specify or comment
| 1 |  |  | Intended Purpose
|  | a |  | Respondents must submit any and all agreements, which lists any tendered product under this HSV agreement, which Successful Respondent has entered into with any of the Participating Health Services; for the purposes of transparency for all Participating Health Services. | --Select--
|  | b |  | Services that may be offered for products tendered under this HSV agreement:
|  |  | (i) | Corrective Maintenance – A fee paying service for single items where respondent identifies, isolates, and rectifies faults so that failed equipment is restored to operational condition within the manufacturer’s established tolerances. | --Select--
|  |  | (ii) | Calibration and Testing Agreements – A contract between Successful Respondent and Participating Health Service for the regular assessment, calibration and testing of the medical equipment to ensure it is safe, and continues to be safe, for its intended clinical application. | --Select--
|  |  | (iii) | Maintenance Agreements – A contract between Successful Respondent and Participating Health Service for respondent to maintain Participating Health Service’s equipment. Maintenance agreements may range anywhere from full comprehensive to parts-only arrangements. | --Select--
|  |  | (iv) | Software Maintenance Agreements – A contract between Successful Respondent and Participating Health Service for the provisions of all software updates and upgrades, that include code updates, version updates, code fix modifications, enhancements, and added or new functional capabilities loaded into the existing software. | --Select--
|  |  | (v) | Hardware Maintenance Agreements – A contract between Successful Respondent and Participating Health Service that extends the useful life of the tendered medical device by replacing hardware, EOL IT equipment, etc. | --Select--
|  |  | (vi) | Amortisation Agreements – A contract between Successful Respondent and Participating Health Service where capital costs are amortised over a fixed number of consumables for a given length of time. | --Select--
|  |  | (vii) | Extended Warranties – An optional arrangement where Successful Respondent may offer to Participating Health Services a care package to extend the length of time of the manufacturer’s warranty. | --Select--
|  |  | (viii) | BME Training – Any BME service training package offered to Participating Health Services by the Successful Respondent, including the initial free-of-charge servicing training and of any on-going annual training. | --Select--
| 2 |  |  | Base Product
|  | a |  | Respondents must provide a full range of service agreement options for each item of equipment listed in Product Catalogue. This may range from a full comprehensive service agreement (24 hours a day, 7 days a week) to a routine testing only agreement, or anywhere in between. Participating Health Services reserve the right to enter into these contracts during the entire, or part of, the effective life of the equipment. | --Select--
|  | b |  | Respondents must submit any and all agreements entered into with any of the Participating Health Services that pertain any products listed in this HSV agreement. | --Select--
|  | c |  | Respondents must submit Service and Maintenance sales data on the template provided as part of this ITS. Note: Service and Maintenance will not be captured in the pricing schedule, rather in the overview tabs, however subcategories are provided for sales reporting purposes. These sales reports must be mail delivery and cannot be uploaded as part of the standard sales data report. Instructions and template are captured in the Reporting Guidelines. | --Select--
|  | d |  | Respondents must submit all Amortisation Agreements that the Respondent has entered into with any Participating Health Service, that pertain any contracted product listed in HSVC2021-055 Infusion Pumps agreement. Note: Amortisation will not be captured as a category/subcategory in the pricing schedule. All amortised sales must be captured in the sales data template on an individual cost component level. Where there is a price variance as a result of amortisation- comment in comments column "Amortised". Further instructions provided in the sales data template. | --Select--
|  | e |  | Unit pricing for service and maintenance is to be listed in the Product Catalogue. Volume discounts must be applied to agreements for the Participating Health Services. | --Select--
| 3 |  |  | Corrective Maintenance
|  | a |  | Respondents must populate the Product Catalogue for corrective maintenance. | --Select--
|  | b |  | Respondents must submit all corrective maintenance sales data on the template provided as part of this ITS, and | --Select--
|  |  | (i) | must list each part used, and its cost, in the repair process, and | --Select--
|  |  | (ii) | must list total hours taken to conduct the repair. | --Select--
|  | c |  | Any pay-as-you-go, routine preventative maintenance services, must also be included in this subcategory | --Select--
| 4 |  |  | Calibration and Testing Agreements
|  | a |  | Respondents must populate the Product Catalogue for Calibration and Testing Agreements. | --Select--
|  | b |  | Respondents must submit all Calibration and Testing sales data on the template provided as part of this ITS, and | --Select--
|  |  | (i) | must list all parts used, and their costs, and | --Select--
|  |  | (ii) | must list total hours taken to conduct the calibration and/or testing. | --Select--
|  | c |  | As per the Draft Deed of Standing Offer Agreement, copies of all calibration and testing agreements for Participating Health services that include any product that has been tendered within this HSV agreement, must be provide to HSV:
|  |  | (i) | on an annual basis (from date of contract commencement), or | --Select--
|  |  | (ii) | upon HSV request. | --Select--
|  | d |  | Respondents must list the exclusions on the Product Catalogue. | --Select--
| 5 |  |  | Maintenance Agreements
|  | a |  | Respondents must populate the Product Catalogue for all variants of their maintenance agreements. | --Select--
|  | b |  | Respondents must submit all Service and Maintenance sales data on the template provided as part of this ITS, and | --Select--
|  |  | (i) | must list each part used, and its cost, in the repair process, and | --Select--
|  |  | (ii) | must list total hours taken to conduct the repair. | --Select--
|  | c |  | As per the Draft Deed of Standing Offer Agreement, copies of all maintenance agreements for Participating Health services that include products under HSVC2021-055 Infusion Pumps agreement, must be provide to HSV:
|  |  | (i) | on an annual basis (from date of contract commencement), or | --Select--
|  |  | (ii) | upon HSV request. | --Select--
|  |  |  | Noting an agreement template forms part of your TRW response. | --Select--
|  | d |  | Respondents must list the exclusions on the Product Catalogue. | --Select--
| 6 |  |  | Amortisation Agreements
|  | a |  | Respondents must include all amortized sales within their quarterly sales data. Instructions will be provided with the sales reporting template as part of this ITS. | --Select--
|  | b |  | As per the Draft Deed of Standing Offer Agreement, copies of all amortisation agreements for Participating Health Services must be provided to HSV:
|  |  | (i) | on an annual basis, or | --Select--
|  |  | (ii) | when there is an amendment to a contract, or | --Select--
|  |  | (iii) | upon HSV request. | --Select--
|  |  |  | Noting an agreement template forms part of your TRW response. | --Select--
|  | c |  | Respondents must offer a complete range of service and maintenance arrangements for inclusion in their amortisation agreements. It’s at the discretion of Participating Health Services to elect a level of service support the Participating Health Service requires to maintain procured products. | --Select--
| 7 |  |  | Software and Hardware Maintenance Agreements
|  | a |  | Respondents must populate the Product Catalogue for all variants of their Software and Hardware maintenance agreements. | --Select--
|  | b |  | Respondents must submit all Software and Hardware agreement sales data on the template provided as part of this ITS. | --Select--
|  | c |  | As per the Draft Deed of Standing Offer Agreement, copies of all Software and Hardware maintenance agreements for Participating Health services that include products under HSVC2021-055 Infusion Pumps agreement, must be provide to HSV:
|  |  | (i) | on an annual basis (from date of contract commencement), or | --Select--
|  |  | (ii) | upon HSV request. | --Select--
|  |  |  | Noting an agreement template forms part of your TRW response. | --Select--
|  | d |  | Respondents must list the exclusions on the Product Catalogue. | --Select--
| 8 |  |  | Extended Warranties
|  | a |  | Respondents must populate the overview tabs in Tender Response Worksheet for all variants of their extended warrant offerings. | --Select--
|  | b |  | Respondents must submit all Extended Warranty sales data on the template provided as part of this ITS. | --Select--
|  | c |  | Respondents must list the exclusions on the Product Catalogue. | --Select--
| 9 |  |  | BME Training
|  | a |  | Respondents must include all their BME Training sales within their quarterly sales data. Instructions will be provided with the sales reporting template as part of this ITS, and | --Select--
|  |  | (i) | must list each initial free-of-charge training provided, and | --Select--
|  |  | (ii) | must list every annual service training provided, it's cost, and location | --Select--

## Glossary
Glossary
Term | Definition
accessory | product that enhances but not essential for device functionality; expected to last the entire effective life of the equipment
ambulatory use | care provided outside of a hospital
authorised user | Any person that is authorised by a Participating Health Service to perform certain functions (e.g. settings and software configuration) on infusion pump systems (e.g. Biomedical Engineers).
BT | Bolus Tocolysis (pulsating infusions)
business day | Any weekday that is not gazetted as a public holiday in Melbourne, Victoria.
Clinical Service | When a product is placed in a Clinical Area, i.e., Ward, and is ready for use; after all user training completed.
consumable | A part that is used once or only on one patient, such as tubing, paper, soda lime etc.
proprietary consumable | A product that is only available from one supplier and only functions with one brand of device
associated consumable | A product that is available from multiple suppliers and only functions with one brand of device
generic consumable | A product that is available from multiple suppliers and operates on multiple brand of devices
effective life | The minimum lifespan of a piece of equipment as defined by the Department of Health’s Medical Equipment Asset Management Framework (MEAMF).
end of life | A date from which no service support or preventative maintenance parts are available from supplier
end of sales | A date from which device is no longer manufactured
end of service | A date from which no repair spare parts and materials are available from supplier
GS1 | Global Standards One
handheld | Items of equipment that are designed to be transported by, and operated in, one hand.
HIS | Hospital Information Systems
may | Indicates an optional element; it is at the Respondent’s discretion to either meet or not meet this element, and failure to meet this element will not have an impact during evaluation.
must | Indicates a mandatory requirement; failure to meet this requirement will have a significant negative impact during evaluation.
NEHTA | National E-Health Transition Authority
NPC | National Product Catalogue
OEM | Original Equipment Manufacturer
OEM product | A product made by an OEM, which may be a dedicated, associated, generic, or refurbished product
options | A product/service that can be added to enhance device functionality as an after sales upgrade
Participating Health Services | Public Hospitals and other Health or Related Services, as those terms are defined in Section 3 of the Health Services Act 1988 (Vic), that are described in Appendix 4 of Part 8.
PCA | Patient Controlled Analgesia
PCEA | Patient Controlled Epidural Analgesia
PCS | Patient Controlled Sedation
portable | Items of equipment that are designed to be moved by hand.
refurbished product | A reprocessed/remanufactured product, including service exchange
reusable | A device designed or intended by the manufacturer as suitable for reprocessing and reuse.
Security Vulnerability | a weakness in a system’s security requirements, design, implementation or operation that could be accidentally triggered or intentionally exploited and result in a violation of the system’s security policy.
semi-consumable | A part that is used multiple times but is not expected to last the entire effective life of the equipment, such as batteries, reusable patient tubing etc., and is replaced by the device's user without the use of a tool
should | Indicates a highly desirable element; unless justifiable reason exists, not meeting this element may have an impact during evaluation.
single use | A device that is intended to be used on an individual patient, during a single procedure, and then discarded.
single-patient use | A device that can potentially undergo more than one episode of use on one patient only. The device may need to undergo some form of reprocessing between each use (in accordance with manufacturers’ instructions).
SLA | Service Level Agreement
spare part | A duplicate part essential for device functionality, requiring a tool to fit, to replace a lost or damaged part of equipment.
super user | A person employed by a Participating Health Service who has the highest level of access to the product (usually the person with the most expertise with the product).
supported life | The duration after purchase in which the successful Respondent provides ongoing service and/or parts for the equipment (the minimum supported life is specified in clause 5.10.1).
TGA | Therapeutic Goods Administration
TCI | Target Controlled Infusion
TIVA | Total Intravenous Anaesthesia
third-party product | A replica of an OEM product that is still available from the OEM
train-the-trainer | Any person that is authorised by a Participating Health Service to provide internal training to end users on the use of infusion pump systems and associated products.
transportable | Items of equipment that are designed to be moved only with transportation aides (such as roll stands) or that are only movable from bracket to bracket.
updates | To bring up to date by adding new information or making corrections.
upgrades | Technology replacement of older equipment, software, or services.
UPS | Uninterruptable Power Supply
will | Indicates an anticipated future condition or requirement.

## Tables
Tables
Delivery
Order type | Maximum delivery time | Time prior to installation
Standard | 90 days | 7 days
Consumables / semi-consumables | 5 days
Spare Parts | 5 days
Emergency Delivery (consumables only) | > 4 hrs (metro) > 2 Business Days (region)
Service
HSV Categories | 1st Level Support | 2nd Level Support[1] | 3rd Level Support | Part Availability
One, Two, Three, Four | 2 hours phone | Courier to service base | On-site | 5 days
|  | Returned to PHS within 14 days | 7 days whole state
Five | 2 hours phone | On-site | On-site | 24 hours
|  | 4 hours (metro) | 12 hours (metro)
|  | 8 hours (regional & rural) | 24 hours (regional & rural)
[1] Unless resolved by phone support

## Appendix 1 - Product List
Appendix 1 - Product List
Category | Category Name | Sub-Category | Sub-Category Name
1 | Volumetric Pumps | 1.01 | Volumetric Infusion Pumps
2 | Syringe Pumps | 2.01 | Syringe Pumps
3 | Pump Platforms | 3.01 | Multi-pump Racks
|  | 3.02 | Pump Docking Stations
|  | 3.03 | Pump Controllers
|  | 3.04 | Optional Patient Parameters
4 | Specific Purpose Pumps | 4.01 | Ambulatory Infusion Pumps
|  | 4.02 | Epidural Pumps
|  | 4.03 | Patient Controlled Analgesic Pumps
|  | 4.04 | Patient Controlled Epidural Analgesic Pumps
|  | 4.05 | Total Intravenous Anaesthesia (TIVA) Pumps
5 | Dose Error Reduction System | 5.01 | Software - Drug Library
|  | 5.02 | Software - Dose Error Reduction System
|  | 5.03 | Software - Quality Analysis of DERS
|  | 5.04 | Hardware - DERS implementation
6 | IV Administration Sets | 6.01 | IV Admin Sets - Volumetric Infusion Pump Sets
|  | 6.02 | IV Admin Sets - Syringe / PCA / TIVA Pump Sets
|  | 6.03 | IV Admin Sets - Ambulatory Infusion Pump Sets
|  | 6.04 | IV Admin Sets - Epidural / PCEA Pump Sets
7 | Semi-Consumables & Accessories | 7.01 | Semi-Cons - Volumetric Infusion Pumps
|  | 7.02 | Semi-Cons - Syringe Pumps
|  | 7.03 | Semi-Cons - Ambulatory Infusion Pumps
|  | 7.04 | Semi-Cons - Patient Controlled Analgesic Pumps
|  | 7.05 | Semi-Cons - Epidural Pumps
|  | 7.06 | Semi-Cons - Patient Controlled Epidural Analgesic Pumps
|  | 7.07 | Semi-Cons - Total Intravenous Anaesthesia Pumps
|  | 7.08 | Semi-Cons - Racks / Docks / Controllers
|  | 7.09 | Semi-Cons - Monitored Patient Parameter
8 | Options (post purchase upgrades) | 8.01 | Options - Volumetric Infusion Pumps
|  | 8.02 | Options - Syringe Pumps
|  | 8.03 | Options - Ambulatory Infusion Pumps
|  | 8.04 | Options - Patient-Controlled Analgesic Pumps
|  | 8.05 | Options - Epidural Pumps
|  | 8.06 | Options - Patient Controlled Epidural Analgesic Pumps
|  | 8.07 | Options - Total Intravenous Anaesthesia Pumps
|  | 8.08 | Options - Base Stations
|  | 8.09 | Options - Docking Stations
|  | 8.1 | Options - DERS
9 | Spare Parts | 9.01 | Testers / Calibration Kits / Tools / Cables / Dongles / etc.
|  | 9.02 | Spare Parts - Volumetric Infusion Pumps
|  | 9.03 | Spare Parts - Syringe Pumps
|  | 9.04 | Spare Parts - Ambulatory Infusion Pumps
|  | 9.05 | Spare Parts - Patient Controlled Analgesic Pumps
|  | 9.06 | Spare Parts - Epidural Pumps
|  | 9.07 | Spare Parts - Patient Controlled Epidural Analgesic Pumps
|  | 9.08 | Spare Parts - Total Intravenous Anaesthesia Pumps
|  | 9.09 | Spare Parts - Racks / Docks / Controllers
|  | 9.1 | Spare Parts - Monitored Patient Parameter
10 | Agreements | 10.01 | Corrective Maintenance
|  | 10.02 | Calibration and Testing Agreements
|  | 10.03 | Maintenance Agreements
|  | 10.04 | Amortisation Agreements
|  | 10.05 | Extended warranties
|  | 10.06 | BME Training
Out of Scope Products
Implantable pumps
Ambulatory Single-Patient Infusion Devices |  |  | refer to HPVC2019-079
Enteral Feeding Pumps |  |  | refer to HPVC2016-016
Neonatal Enteral Feeding Pumps |  |  | refer to HPVC2016-016
Non-dedicated Sets and Lines |  |  | refer to HPVC2019-079
IV line Accessories |  |  | refer to HPVC2019-079
Syringes |  |  | refer to HPVC2015-005
Insulin Pumps

## Appendix 2 - References
Appendix 2 - References
Standards | Standard Number | Standard Name
References to standards include any amendments, revisions or consolidations to those standards | AS/NZS IEC 60601.1 | Medical electrical equipment Part 1: General requirements for basic safety and essential performance
| AS IEC 60601.1.2 | Medical electrical equipment Part 1.2: General requirements for basic safety and essential performance—Collateral Standard: Electromagnetic disturbances—Requirements and tests
| AS IEC 60601.1.6 | Medical electrical equipment Part 1.6: General requirements for basic safety and essential performance—Collateral standard: Usability
| AS IEC 60601.1.8 | Medical electrical equipment Part 1.8: General requirements for basic safety and essential performance—Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems
| AS IEC 60601.1.9 | Medical electrical equipment Part 1.9: General requirements for basic safety and essential performance—Collateral Standard: Requirements for environmentally conscious design
| AS IEC 60601.1.10 | Medical electrical equipment Part 1.10: General requirements for basic safety and essential performance—Collateral Standard: Requirements for the development of physiologic closed-loop controllers
| AS IEC 60601.1.11 | Medical electrical equipment Part 1.11: General requirements for basic safety and essential performance—Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment
| AS IEC 60601.1.12 | Medical electrical equipment Part 1.12: General requirements for basic safety and essential performance—Collateral Standard: Requirements for medical electrical equipment and medical electrical systems intended for use in the emergency medical services environment
| AS 60601.2.24 | Medical electrical equipment Part 2.24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers
| IEC 80601-2-49 | Medical electrical equipment — Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment
| AS ISO 14971 | Medical devices - Application of risk management to medical devices
| AS ISO 13485 | Medical devices - Quality management systems - Requirements for regulatory processes
| ANSI/NEMA HN 1-2019 | American National Standard — Manufacturer Disclosure Statement for Medical Device Security
| UL 2900 series | Software Cybersecurity for Network-Connectable Products
| AS/NZS 4535 | Ambulance restraint systems
Standards for Associated Services | Standard Number | Standard Name
| AS/NZS 3551 | Management programs for medical equipment
| AS/NZS IEC 60601.1 | Medical electrical equipment Part 1: General requirements for basic safety and essential performance
| AS/NZS 4187 | Reprocessing of reusable medical devices in health service organizations
AS/NZS 2500
Legislation | Name
References to legislation includes any amendments, revisions or consolidations to those references | Electricity Safety Act 1998
Occupational Health and Safety Act 2004
Therapeutic Goods Act
Guidelines and Other References | Issuing Body | Title
References to guidelines include any amendments, revisions or consolidations to those guidelines | DHHS (2021) | Medical Device Security Areas of Practice
| ENISA (2020) | Procurement Guidelines for Cybersecurity in Hospitals
| ISMP | Guidelines for Optimizing Safe Implementation and Use of Smart Infusion Pumps
| NHMRC (2010) | Australian Guidelines for the Prevention and Control of Infection in Healthcare, Commonwealth of Australia
| Therapeutic Goods Administration (2019) | Uniform Recall Procedure for Therapeutic Goods, Commonwealth of Australia
| WorkSafe Victoria (2019) | Compliance Code - Hazardous Manual Handling, 2nd Ed

## Appendix 3 - KPIs
Appendix 3- Key Performance Indicators (KPI's)
Number | KPI | Description | Measure
1. | Product Delivery | Product Delivery: Delivery timeframes quoted by the Supplier are not met (consumables and semi-consumables) Delivery within 5 business days Spare Parts Delivery: Delivery of spare parts within 5 business days from time of order Product Delivery (consumables only) – Emergency: Urgent deliveries must be received by Participating Health Services within the shortest possible timeframe; however, this should not exceed the following timeframes: > 4 hrs (metro) > 2 Business Days (region) | No. of PO’s supplied in full, on time and to the quality specified / Total no of PO’s To be monitored and measured by random audit in collaboration with health services To be monitored and measured by random audit in collaboration with health services
| Product Shelf Life | Consumables with a normal shelf life >12 months, that are delivered with <12 months remaining. | To be monitored and measured by random audit in collaboration with health services. Goods received documents
| Backorders | Supplier exceeds their notified backorder time frames. | To be monitored and measured by random audit in collaboration with health services. Copies of Contractor Quotations, Contract commitments specifying delivery time.
| Catalogue Pricing | The Supplier must ensure 100% compliance of catalogue pricing and invoicing alignment at every transaction. | No. of sales data transactions consistent with HPV pricing / Total No. of sales data transactions
| Product collection and substitution | ≥95% collected (AND substitute provided where available and requested) | To be monitored and measured by random audit in collaboration with health services
| Spare Parts Delivery | Delivery of spare parts within 5 business days from time of order | To be monitored and measured by random audit in collaboration with health services
| Product Returns | Items repaired under warranty must be returned to Participating Health Services within five (5) business days from when the item is received by the successful Respondent. | To be monitored and measured by random audit in collaboration with health services
| Product Returns Collection | Collection of returnable items (according to the suppliers returns policy) within 5 Business Days from advice to Supplier | To be monitored and measured by random audit in collaboration with health services
| Product Trial | Suppliers must be capable of demonstrating their products within four (4) weeks of request. | To be monitored and measured by random audit in collaboration with health services
| Service – Technical Enquiries | Availability of technical resources to provide support to end users when requested | To be monitored and measured by random audit in collaboration with health services
| Victorian Product Catalogue System (VPCS) | Must achieve and maintain VPCS Supplier Scorecard of ≥95% within 3 months of contract implementation | Measurable by VPCS Supplier Scorecard 14 July 2021
| Recallnet Compliance | 100% compliance within three months of contract implementation | Supplier is complying or non-complying on 14 July 2021
| Contract Management Reporting | Making available electronic reporting upon request in the areas of: - DIFOT performance - Back order performance analysis - Spare parts delivery - Service and maintenance performance - Product returns | To be measured upon request at HPV No. of requested reports submitted on time /No. of requested reports due
| Sales & KPI reports | 95% of sales reports (including amortization sales data and service and maintenance sales data) are received on-time as per the contract reporting schedule and are of a standard of accuracy acceptable to HPV 95% of inventory reports (Reporting 1) are received on-time and to a standard of accuracy acceptable to HPV | No. of accurate sales reports submitted on time / No. of sales reports due No. of accurate stock reports submitted on time / No. of reports due
| Patch Updates | OS patches available for medical devices 30 days after the release date from the OS provider

## List (hidden)
--Select-- | --Select--
Fully Comply | Yes
Partially Comply | No
Does Not Comply
Not Applicable