## List (hidden)
--Select-- | --Select--
Fully Comply | Yes
Partially Comply | No
Does Not Comply | Not applicable
Not Applicable

## Definitions
Definitions
Term | Definition
AHA | American Heart Association
Authorised User | Any person that is authorised by a Participating Health Service to perform certain functions (e.g. settings and software configuration) on Ventilator systems (e.g. Biomedical Engineers).
Business Day | Any weekday that is not gazetted as a public holiday in Victoria.
Business Hours | The Hours During The Day In Which Business Is Commonly Conducted (8:00 – 17:00)
CPR | Cardiopulmonary resuscitation
cold start | Starting from mains power being off with configurations and program processing beginning from initial values
Consumable | A part that is intended to be used for a limited time period or limited number of uses on a single patient
proprietary consumable | a product that is only available from one supplier and only functions with one brand of device
associated consumable | a product that is available from one or more suppliers that only functions with one brand of device
generic consumable | a product that is available from multiple suppliers and operates on multiple brand of devices
CTG | Cardiotocography
DIFOT | Delivery in full, On time
ECG | Electrocardiography
EEG | Electroencephalography
Effective Life | The minimum lifespan of a piece of equipment as defined by the Department of Health’s Medical Equipment Asset Management Framework (MEAMF).
eMR | Electronic Medical Records
ESU | Electrosurgical Unit
firmware | Software required by a device, for the device hardware to perform device’s intended purpose.
Fixed | Items of equipment that are not intended to be moved by end users for normal use.
FOC | free of charge
foetus | unborn baby
GS1 | Global Standards One
GTIN | Global Trade Identification Number
Handheld | Items of equipment that are designed to be transported by, and operated in, one hand.
HIS | Hospital Information System
HSV | HealthShare Victoria
IABP | Intra-Aortic Balloon Pump
ICT | Information and Communications Technology
Infant | a patient from 4 weeks of age from term to weighing 10kg
IPXX | Ingress Protection Rating, Dust and Fluid
ITS | Invitation To Supply
Movable / Mobile | Items of equipment that can be transported with transportation aides (such as trolleys, roll-stands etc.).
Must | Indicates a mandatory requirement; failure to meet this requirement will have a significant negative impact during evaluation and may result in non-award. HSV may choose to amend such requirements at any stage during the tendering stage, based on either, the deviation requests received from tenderers, or otherwise.
NEHTA | National E-Health Transition Authority
Neonate | newborn baby less than 4 weeks old, from term
Normal Use | Means that the item has undergone use for which it was manufactured and intended and shows no signs of physical damage other than regular wear and tear.
NPC | National Product Catalogue
OEM | Original Equipment Manufacturer
OMS | Obstetric Management Systems
PACS | Picture Archiving and Communications System
Paediatric | a patient greater than 10kg and less than 18 years old
Participating Health Services (PHS) | Public Hospitals and other Health or Related Services, as those terms are defined in Section 3 of the Health Services Act 1988 (Vic), that are described in Appendix 2 of Parts 1-3 of the ITS.
PC | Personal Computer
PEEP | Positive End Expiratory Pressure
PIP | Peak Inspiratory Pressure
Portable | Items of equipment that are designed to be lifted and carried by hand.
RFI | Request for Information
RFID | Radio Frequency Identification
Reusable | A device designed or intended by the manufacturer as suitable for reprocessing and reuse.
Semi-Consumable | A part that is used multiple times but is not expected to last the entire effective life of the equipment, such as batteries, reusable patient tubing etc.
Should | Indicates a highly desirable element; unless justifiable reason exists, not meeting this element may have a medium impact during evaluation.
SLA | Service Level Agreement
Single Use | A device that is intended to be used on an individual patient, during a single procedure, and then discarded.
Single-Patient Use | A device that can potentially undergo more than one episode of use on one patient only. The device may need to undergo some form of reprocessing between each use (in accordance with manufacturers’ instructions).
Spare Part | Any item required for any preventative or corrective maintenance process in maintaining tendered machine
common | a part that has a Failure Rate (MTBF) of <12 months, or
purchased more than once in a 12 month period by the sum all PHSs on this agreement.
essential | a part required in the equipment for it to function as prescribed in Intended Use, e.g. motherboard, power supply, etc.
generic | a part that is available from multiple Suppliers
infrequent | a part that needs to shipped to Australia from overseas
proprietary | a part that is only supplied by Respondent
Supported Life | The duration after purchase in which the successful Respondent provides ongoing service and/or parts for the equipment (the minimum supported life is specified in clause 5.11.
TCP/IP | Transmission Control Protocol/Internet Protocol
TGA | Therapeutic Goods Administration
Train-The-Trainer | Any person that is authorised by a Participating Health Service to provide internal training to end users on the use of Ventilator systems and associated products.
Transportable | Items of equipment that are designed to be moved only with transportation aides (such as roll stands) or that are only movable from bracket to bracket.
Will | Indicates an anticipated future condition or requirement.
UPS | Uninterruptable Power Supply

## A - Introduction
Part A: Introduction
Clause Identifiers |  |  |  | Text
Part A |  |  |  | Introduction
| 1 |  |  | Purpose
|  | a |  | The purpose of this Part 5 – Statement of Requirements, is to:
|  |  | i | Detail the scope and range of products sought under this Invitation to Supply (ITS)
|  |  | ii | Specify the requirements that Respondents and/or their offered products must meet. These requirements also form part of any resulting Agreement between HSV and any successful Respondent(s).
| 1.1 |  |  | Participating Health Services
|  | a |  | The Participating Health Services for this ITS are:
|  |  | i | All ‘Public Health Services’ (as legislatively defined) referred to in Schedule 1 and Schedule 5 of the Health Services Act 1988; and
|  |  | ii | Any other health related entity receiving public funding that is permitted under the Health Services Act 1988 to access HPVC2024-075.
| 2 |  |  | Scope
|  | a |  | The scope of this Invitation to Supply includes:
|  |  | i | supply of physiological monitors, including optional features
|  |  | ii | supply of basic gas delivery devices (O2, medical air, suction etc.)
|  |  | iii | supply of nitrous oxide and oxygen delivery devices
|  |  | iv | supply of anaesthetic machines, including optional features
|  |  | v | supply of all relevant consumables and accessories
|  |  | vi | servicing and maintenance of equipment
|  |  | vii | delivery of a training package, including relevant training materials.
|  | b |  | The scope of this Invitation to Supply is to include gas delivery systems which are involved in the delivery of care by anaesthetists, that may incorporate a physiological monitor, anaesthetic agents and/or ventilation, used both in and/or outside the operating theatre environment’ e.g. MRI Suite. It also includes other medical gas delivery systems such as nitrous oxide and oxygen delivery systems, O2, medical air, NO2, etc. Refer to Appendix 1 – Product List for a full list of categories and subcategories under this Invitation to Supply.
|  | c |  | The scope of this Invitation to Supply is for ECG machines with built-in printers capable of printing out at least a 3x4 1R ECG report onto one sheet of paper; including ECG machines used for stress testing.
|  | d |  | HSV is seeking responses for various supply solutions, including but not limited to:
|  |  | i | outright purchase of equipment
|  |  | i | leasing of equipment
|  |  | iii | supply of equipment only
|  |  | iv | supply of consumables and/or semi-consumables only
|  |  | v | free-on-loan equipment through consumable volume agreement
|  |  | vi | whole-of-life support/managed supply environment, including but not limited to: supply of equipment and all associated consumables maintenance (including scheduled inspection and unscheduled repairs) warranty services disposal services.
|  | e |  | Respondents may offer products in one, some or all categories listed in Appendix 1 – Product List. However, Respondents who offer a telemetry units must also offer a corresponding data management system/central station.
|  | f |  | Items must only be tendered in one category. It is at the Respondent's discretion to determine the most suitable category
| 3 |  |  | Categories
|  | a |  | Preference will be given to Respondents offering both the best value for money (based on the total cost of ownership) and the best fit for purpose.
|  | b |  | The categories called for under this Invitation to Supply (ITS) include:
|  |  | i | 1.0 - Vital Signs Monitoring
|  |  | ii | 2.0 - Acute Care Monitoring
|  |  | iii | 3.0 - ECG Machines
|  |  | iv | 4.0 - Ambulatory Monitoring
|  |  | v | 5.0 - Intrapartum Monitoring
|  |  | vi | 6.0 - Neurological, Sleep, and Sleep Related Monitoring
|  |  | vii | 7.0 - Basic Gas Delivery Devices
|  |  | viii | 8.0 - Portable Suction Units
|  |  | ix | 9.0 - Nitrous Oxide and Oxygen Delivery Systems
|  |  | x | 10.0 - Anaesthesia Workstations
|  |  | xi | 11.0 - Physiological Monitoring Software Upgrades
|  |  | xii | 12.0 - Physiological Monitoring Associated Consumables
|  |  | xiii | 13.0 - Physiological Monitoring Semi-Consumables
|  |  | xiv | 14.0 - Monitoring Products for Physiological Monitoring
|  |  | xv | 15.0 - Respiratory Products for Physiological Monitoring
|  |  | xvi | 16.0 - Physiological Monitoring Accessories
|  |  | xvii | 17.0 - Physiological Monitoring Equipment Spare Parts
|  |  | xvii | 18.0 - Physiological Monitoring Service & Maintenance (S&M) Agreements
|  | c |  | For a full list of products refer to Appendix 1 – Product List.
|  | d |  | The Respondent may offer products in one, some or all categories.
|  | e |  | HSV reserves the right not to consider any additional products offered.
| 4 |  |  | Product Conditions
| 4.1 |  |  | Product Offering
|  | a |  | All proprietary consumables and semi consumables that can be used on tendered equipment are in scope of this ITS and must be included in your tender response.
|  | b |  | HSV will not consider any product that is subject to a current HSV Agreement, other than the below categories/subcategories on the current Monitoring Products contract (HPVC2023–021).
|  | c |  | HSV will not consider any product that is subject to a current HSV Agreement, other than the below categories/subcategories on the current Respiratory Products contract (HPVC2021–015).
|  | d |  | HSV will not consider any product that is subject to a current HSV Agreement.
|  | e |  | Respondents must ensure that each product is offered in only one subcategory. It is at the Respondent’s discretion to ensure that each product is submitted in the most appropriate subcategory.
|  | f |  | OEM Products, devices and consumables will need to be contracted under this agreement
| 4.2 |  |  | Product Information
|  | a |  | All product information submitted must:
|  |  | i | be in electronic format
|  |  | ii | be in English
|  |  | iii | be specific to the product offered
|  |  | iv | contain the Respondent’s company name
|  |  | v | include the product code
|  |  | vi | include a detailed specification of the product
|  |  | vii | include clear diagrams/pictures of the product.
|  | b |  | To assist in managing this material, all product information submitted must be labelled with the relevant HSV category and subcategory number;
|  |  | i | for electronic copies must include the HSV Category and subcategory numbers in the filename or identifying metadata.
|  | c |  | HSV advises any unlabelled submissions will not be reviewed.
|  | d |  | Product information will not be evaluated, but will assist in accurately identifying products offered.
|  | e |  | HSV reserves the right to exclude products from evaluation when the product is unidentifiable and provided information is:
|  |  | i | not labelled as per clause (b) above; or
|  |  | ii | is incomplete as to clause (a).
|  | f |  | Respondents should note that product samples and research papers are not to be provided, unless specifically requested by HSV.
|  | g |  | The Respondent should not submit information relating to products that are not called for in this Invitation to Supply.
| 4.3 |  |  | Conditional Award
|  | a |  | Products that are awarded conditionally can be trialled and evaluated at any stage of the contract. For products to be moved from conditional award to full award, suppliers will need to provide HSV with supporting documentation, including Victorian public health service references and evidence of immediate health service demand (pending HSV approval). Conditional award requirements are included in the HSV Contract Variation Request Form, available from the HSV website (www.healthsharevic.org.au)
| 4.4 |  |  | Clinical Trials
|  | a |  | Participating Health Services may, at their discretion, research or trial new technology or use non-contracted products to perform clinical trials at any time during the Term of any resulting Agreement.
|  | b |  | Participating Health Services may request trials before creating a Statement of Demand which may also provide successful Respondents time to do a site inspection.

## B - Service, Delivery & Support
Part B: Service, Delivery and Support
Clause Identifiers |  |  |  | Text | Respondent's Level of Compliance |  | No extra columns to be added by Respondents
Part B |  |  |  | Statement of Work | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Delivery
|  | a |  | All tendered equipment must be priced and supplied to Participating Health Services, with all the necessary accessories and consumables, to function as intended by the OEM. | --Select--
|  | b |  | Equipment must be delivered to the location(s) specified by Participating Health Services within the shortest possible timeframe; however, this must not exceed 90 days (3 calendar months) from receipt of order unless agreed by Participating Health Service. | --Select--
|  |  | i | Respondent must immediately inform, both HSV and the Participating Health Service, upon becoming aware that equipment delivery will exceed the 90 days from the purchase order date. | --Select--
|  | c |  | All deliveries are bound as per the Deed of Standing Offer Agreement, Part 5 clause 7.2 Acceptance and Rejection of Deliverables. | --Select--
|  | d |  | Deliverables determined dead at commissioning must:
|  |  | i | be replaced within two (2) weeks from discovery, or | --Select--
|  |  | ii | provide a loan unit that is FOC and acceptable by Participating Health Service within two (2) weeks of discovery. | --Select--
|  | e |  | Consumables and semi-consumables that are purchased as replacements post–original purchase must be delivered to Participating Health Services within the shortest possible timeframe; however, this must not exceed five (5) business days from receipt of order unless otherwise mutually agreed by Participating Health Service and Successful Respondent. | --Select--
|  |  | i | The supplier must notify the purchase requestor in writing (email) if the purchase will not be delivered within five (5) business days at the time of receipt of purchase order. | --Select--
|  | f |  | Equipment, along with AS/NZS 3551 testing documentation (if applicable), must be delivered to Participating Health Service sites seven (7) to fourteen (14) days prior to installation, unless negotiated otherwise with the Participating Health Service. | --Select--
|  | g |  | Successful Respondents must be capable of installing their devices and/or systems and, if requested to do so, must cooperate with Participating Health Services’ biomedical engineering and ICT departments (see Part B Section 7 Biomedical Equipment Maintenance and Repair). | --Select--
|  | h |  | Successful Respondents will be required to configure the devices and/or systems upon delivery, or | --Select--
|  |  | i | if requested by the Participating Health Service, must provide the required configuration tools (software, cables, dongles, etc.), in line with clause Part B 3.2 (c), for the Participating Health Service’s authorised user/s to configure the device/s, or | --Select--
|  |  | ii | if requested by the Participating Health Service, must train Participating Health Service’s authorised user/s to configure the device/s. This must be provided free of charge as part of the supply of Deliverables, and is in addition to the training requirements under Part B Section 3 Training. | --Select--
|  | i |  | Successful Respondents must notify HSV and Participating Health Services of 'end of sales' with a Discontinuation Notice at least six (6) months prior to sales ending, for any tendered equipment, accessory, semi-consumable, or associated consumable. | --Select--
| 1.1 |  |  | Urgent Deliveries
|  | a |  | For the purposes of this section, urgent deliveries refers to urgent requests placed by an individual Participating Health Service, and does not include state-wide emergency situations. | --Select--
|  | b |  | The Respondent must be able to receive and action urgent delivery requests during normal business hours. | --Select--
|  | c |  | Urgent delivery orders that are placed prior to Respondent’s order entry deadline, should be received by Participating Health Services within timelines as prescribed in Tables – Delivery. | --Select--
|  |  | i | Urgent deliveries of Category 12 Associated Consumables that are placed prior to Respondent’s order entry deadline, must be received by Participating Health Services within the timelines as prescribed in Tables – Delivery. | --Select--
|  | d |  | Any additional costs associated with Urgent Deliveries must be clearly outlined in the User Guide tab of the Tender Response Worksheet and shall be discussed and agreed with the ordering health service prior to acceptance of any urgent order. | --Select--
| 1.2 |  |  | Consumables
|  | a |  | All consumables must have a shelf life of at least 12 months upon delivery, unless otherwise agreed by Participating Health Service. | --Select--
|  |  | i | If not agreed to by Participating Health Service and consumable (stock) needs to be disposed of due to short shelf life, then disposed stock must be replaced by successful Respondent at no cost to Participating Health Service. | --Select--
|  | b |  | The Respondent must always have adequate stock-on-hand of the following for the life of this agreement, or until the applicable associated tendered equipment’s end-of-support date:
|  |  | i | for all tendered consumables | --Select--
|  |  | ii | for all tendered semi-consumables | --Select--
|  |  | iii | for all tendered batteries | --Select--
|  |  |  | • If for unforeseen reasons there is a failure to supply any of the above, list alternative option/s in the Tender Response Worksheet. | --Select--
| 2 |  |  | Implementation
|  | a |  | Successful Respondents will be required to provide an implementation plan for Participating Health Services who conduct or participate in an RFQ under this agreement. | --Select--
| 3 |  |  | Training
| 3.1 |  |  | Operator Training
|  | a |  | Upon the request of Participating Health Services, successful Respondents must provide full details of their training package including the training material, the number of trainees to participate, and the duration of each session, to facilitate the introduction of tendered equipment. | --Select--
|  | b |  | Training must:
|  |  | i | be FOC and provided to Participating Health Services anytime during the life of the equipment | --Select--
|  |  | ii | be provided by experienced and qualified personnel, | --Select--
|  |  | iii | be provided as an on-site face-to-face package, unless otherwise requested by Participating Health Service, | --Select--
|  |  |  | • and accommodate all relevant staff, as determined by Participating Health Service, over a number of shifts that may include nights and weekends | --Select--
|  |  | iv | if requested, be provided by videotelephony during business hours, | --Select--
|  |  | v | if requested, be provided as a web-based training package for night shift and weekend workers, | --Select--
|  |  | vi | include the directions on the correct cleaning procedure and appropriate cleaning/disinfectant products to use, and | --Select--
|  |  | vii | if requested by Participating Health Service include face-to-face superuser training, of at least 1 superuser for every 4 units purchased, at no additional cost to Participating Health Service, and | --Select--
|  |  |  | • a minimum of 2 superusers must be trained per Participating Health Service unless otherwise agreed to by the Participating Health Service and successful Respondent. | --Select--
|  | c |  | Educational/Training material must be provided free of charge, and | --Select--
|  |  | i | be in a digital format. | --Select--
|  | d |  | All training aids that are directly related to the Deliverables must be tendered and priced separately. | --Select--
|  |  | i | Where a semi-consumable/consumable is required with training aids, (e.g. chargers, batteries, etc.) this must also be tendered. | --Select--
|  | e |  | Physiological monitors should easily interface with commonly used training equipment (e.g. manikins). | --Select--
|  | f |  | All training programs must include appropriate levels of training to meet Occupational Health & Safety issues as required by Worksafe Victoria | --Select--
|  | g |  | The Respondent is expected to provide training to Participating Health Service users for the effective life of the equipment. This is to ensure that new employees have access to the manufacturers’ training. | --Select--
| 3.2 |  |  | Instruction For Use Manuals
|  | a |  | With each equipment supplied, successful Respondents must provide one complete digital set of operator user manuals and quick user guides. | --Select--
|  | b |  | Respondent’s website should have links to downloadable IFU manuals that Participating Health Services can access. | --Select--
|  | c |  | Instructions for Use manuals must be in English and must include but not be limited to: | --Select--
|  |  | i | Comprehensive operating instructions | --Select--
|  |  | ii | All information necessary to operate the equipment in accordance with specifications and without exceeding safety limits | --Select--
|  |  | iii | All error messages and troubleshooting procedures | --Select--
|  |  | iv | Safety requirements and precautions | --Select--
|  |  | v | Recommended cleaning procedure. | --Select--
|  | d |  | Participating Health services may preference successful Respondents that can offer learning modules via a mobile app, or | --Select--
|  |  | i | QR code on equipment to access quick user guide. | --Select--
| 4 |  |  | Customer Service and Support
|  | a |  | The successful Respondent must be able to deliver customer service and support to Participating Health Services during business hours. | --Select--
|  |  | i | Participating Health Services may preference successful Respondents that can offer product support over the phone outside of business hours. | --Select--
|  | b |  | The successful Respondent must provide Participating Health Services with representatives that are:
|  |  | i | inherently familiar with the contracted products | --Select--
|  |  | ii | appropriately qualified | --Select--
|  |  | iii | technically/clinically knowledgeable about the contracted products | --Select--
|  |  | iv | available to respond to Participating Health Services’ queries during business hours | --Select--
|  | c |  | It is desirable that nominated Representatives have a clinical background or experience. | --Select--
|  | d |  | The level of customer service and support required of Representatives is expected to include (but is not limited to):
|  |  | i | liaising with clinicians to recommend products and solutions | --Select--
|  |  | ii | promptly answering clinicians’ queries (including after hours) | --Select--
|  |  | iii | liaising with various hospital departments (for example: operating theatre, Nurse Unit Managers) | --Select--
|  |  | iv | providing on-site clinical support during cases (if requested) | --Select--
|  |  | v | providing informational materials | --Select--
|  |  | vi | providing education and in-service training upon request. | --Select--
|  | e |  | Representatives must comply with Participating Health Services’ local policies regarding engagement with Participating Health Service staff. | --Select--
| 5 |  |  | Warranty
|  | a |  | Equipment, parts, cables, batteries and all other semi-consumables must be warranted for normal use from the date of first use. | --Select--
|  |  | ii | Participating Health Services are responsible for maintaining accurate records of first use. | --Select--
|  |  | i | Category 2, 6 and 10 items must be warranted for normal use from first date of use. | --Select--
|  | b |  | All equipment offered must have a minimum 24-month warranty period. | --Select--
|  |  | i | All factory prescribed preventive maintenance procedures during the warranty period must be included as a part of warranty and FOC to Participating Health Services, including PM parts, labour, and travel. | --Select--
|  | c |  | OEM batteries must have a minimum 12-month warranty period. | --Select--
|  | d |  | All semi-consumables should offer a minimum 12 month warranty period. If not please list which do not and their warranty in days. | --Select--
|  | e |  | Packaged (still unused) consumables must be warrantied until the expiry date, provided they are handled and stored in as per the manufacturer instructions; mishandling will void warranty. | --Select--
|  | f |  | Consumables and semi-consumables should be replaced rather than repaired. | --Select--
|  | g |  | The manufacturer’s warranty must not be impacted by the use of generic consumables. | --Select--
|  | h |  | If available, Respondents should provide details of any extended warranty that is available beyond the manufacturer’s warranty (including the cost of such warranty per additional year). It will be at the discretion of Participating Health Services whether or not to purchase any extended warranty. | --Select--
|  | i |  | All warranties must clearly indicate exclusions. State all warranty exclusions. | --Select--
| 5.1 |  |  | Repairs and Replacements under Warranty
|  | a |  | The repair or replacement of any item under warranty will be at no cost to the Participating Health Service; excluding accidental and malicious damage. | --Select--
|  | b |  | The cost of any pickup, delivery or on site visit associated with a repair or replacement under warranty must be borne by the successful Respondent | --Select--
|  | c |  | Items repaired under warranty must be returned to Participating Health Services within ten (10) business days from when the item is received by the successful Respondent, and | --Select--
|  |  | i | Respondent must provide a loan unit that is acceptable and FOC to Participating Health Service within two (2) business days of having been notified of faulty equipment during the warranty period. | --Select--
| 6 |  |  | Service and Maintenance
|  | a |  | Maintenance, spare parts, and support services must be available for a minimum of seven (7) years, from the date tendered equipment is delivered to a Participating Health Service. | --Select--
|  | b |  | Service and maintenance must be available throughout the state of Victoria, including Albury-Wodonga. | --Select--
|  | c |  | Respondent must have qualified service engineer/s based in Victoria. | --Select--
|  | d |  | Successful Respondents must adhere to the following Service & Maintenance response times,
|  |  | i | as prescribed in Tables - Service & Maintenance. | --Select--
|  | e |  | Respondent must have a repair facility in an Australian capitol city. | --Select--
|  | f |  | A detailed service report must be provided to Participating Health Services each time their equipment is serviced. | --Select--
|  | g |  | Successful Respondents must provide Participating Health Services with a list of all commercial service providers that are capable of, and authorised to, provide ongoing support. | --Select--
|  | h |  | Respondents must provide a full range of service contract options for each equipment or application offered. | --Select--
|  |  | i | This may range from comprehensive maintenance agreements (24 hours a day, 7 days a week) to routine testing contracts. | --Select--
|  |  | ii | Participating Health Services reserve the right to enter into these contracts during the entire, or part of the, effective life of the equipment. | --Select--
| 6.1 |  |  | Spare Parts and Repairs
|  | a |  | Spare parts must be available for a minimum of five (5) years from the date of last manufacture of each tendered product. Written notification of end of life notification must be provided to relevant department head. | --Select--
|  | b |  | Ensuing clause 6.1(a), service and repair spare parts should be available for purchase by Participating Health Services regardless of the service training status of Participating Health Service's staff. | --Select--
|  | c |  | If requested by a Participating Health Service, successful Respondents must provide a comprehensive list of replacement parts, accessories and assemblies that may be required to repair and maintain the equipment over the supported life. The list must include part numbers and current purchase prices. | --Select--
|  | d |  | If requested by a Participating Health Service, successful Respondents must provide non-warranty repairs for their equipment for its supported life. Where practical, consumables and semi-consumables should be replaced, rather than being repaired. | --Select--
|  | e |  | All 'common' and 'essential' spare parts for tendered equipment must be delivered to Participating Health Service within timeframes prescribed in Tables - Service & Maintenance from Respondent receiving purchase order. | --Select--
|  |  | i | All 'common' and 'essential' spare parts must be stocked in Australia. | --Select--
|  | f |  | If the successful Respondent(s) cannot meet the timeframes specified in clause 5(e), they must notify Participating Health Services before the expiry of these timeframes. | --Select--
|  | g |  | In the event that equipment is damaged (regardless of whether the damage is covered under warranty) and a repair cannot be completed within the timeframes (Parts Availability) specified in clause 6.1e, it is desirable for the successful Respondent to supply an interim comparable item, (i.e. a loan unit) to the relevant Participating Health Service. | --Select--
|  | h |  | Successful Respondents must notify HPV and Participating Health Services of 'end of support' for all equipment and associated consumables & semi-consumables at least three (3) years before the end of support date occurs. For equipment that requires an upgrade, as part of a repair during its life cycle, the cost of this upgrade must not exceed 25% of a new replacement device. If the upgrade cost exceeds this amount, then the device model will be considered as end-of-support. | --Select--
|  |  | i | For ICT Systems, successful Respondents must notify HPV and Participating Health Services of end of support at least 18 months before the end-of-support date occurs. Wherever possible, Respondents should give notice as early as possible. | --Select--
|  | i |  | Ensuing clause 5(h), the choice of whether to upgrade or purchase new items is at the discretion of Participating Health Services. | --Select--
|  | j |  | During the product lifecycle, all firmware must be upgradable as new technology becomes standard. | --Select--
| 6.2 |  |  | Biomedical Equipment Maintenance and Repair
|  | a |  | Respondents agree and understand that Participating Health Services may have an internal biomedical engineering department, or use the services of an external biomedical engineering contractor, to perform the following for tendered equipment: | --Select--
|  |  | i | safety and performance testing | --Select--
|  |  | ii | acceptance testing for new equipment (required before clinical use) | --Select--
|  |  | iii | assessing whether a repair is a warranty issue or not and, in the case of a warranty issue, coordinating the repair with the successful Respondent | --Select--
|  |  | iv | performing non-warranty repairs when requested. | --Select--
|  |  | v | performing the required routine preventative maintenance schedule | --Select--
|  |  | vi | installing user configuration settings | --Select--
|  | b |  | Successful Respondents must offer ongoing service training to Participating Health Services’ biomedical engineering department. | --Select--
|  |  | i | Specify how and when the above training will be delivered (e.g. annual seminar, face-to-face training at Participating Health Service, virtual, etc.) and the costs for such training | --Select--
|  |  | ii | The first Biomed service training for each Participating Health Service must be provided free of charge after the Participating Health Service purchase of a model that is currently not used by them, and | --Select--
|  |  |  | • where there is a complete fleet replacement, the initial biomed service training must be provided on-site at the health service for no additional cost. | --Select--
|  | c |  | Post-training, successful Respondents who use a software tool for servicing and/or configuring their equipment must provide a minimum of one (1) copy of this software to authorised users free of charge. This includes any special cables, dongles etc. | --Select--
|  | d |  | Post-training, successful Respondents who use tools for servicing and/or configuring their equipment must list in the Tender Response Worksheet all tools that are required to service and/or configure the equipment over its supported life. | --Select--
|  |  | i | The list must include part numbers and current purchase prices. | --Select--
|  | d |  | The training referenced in clause 7(b) is in addition to the successful Respondents’ training requirements under Part B Section 3 Training. | --Select--
|  | e |  | Upon purchase of any tendered equipment, and upon request by Participating Health Services, the successful Respondent must provide a full technical service manuals in English for repairs of the system and accessories. | --Select--
|  |  | i | The initial (first) set of manuals must be provided at no cost and regardless of the service training status of the biomedical engineering department or nominated persons. | --Select--
| 7 |  |  | Trial and Evaluation
|  | a |  | Where practical, Respondents may be required to arrange a site visit during the evaluation stage for trial and evaluation purposes. | --Select--
|  | b |  | Participating Health Services reserve the right to trial successful Respondents’ products on-site for reference prior to purchase. To limit the costs of any such trials, full integration to hospital systems is preferred but not required. | --Select--
|  | c |  | Clinical trials will be for a minimum of one (1) week and maximum of up to four (4) weeks per Participating Health Service site, as mutually agreed with Supplier. | --Select--
|  | d |  | Suppliers should be capable of demonstrating their products within four (4) weeks of request. | --Select--
|  | e |  | As part of the trial and evaluation, consumables required for the product, must be FOC | --Select--
| 8 |  |  | Award
| 8.1 |  |  | Conditional Acceptance
|  | a |  | Products may be designated as ‘Conditionally Accepted’ for the following reasons:
|  |  | i | Where products offered are not ‘known and accepted’ but represent value for money; or | --Select--
|  |  | ii | Where products are inactive and have not been in use for at least 12 months. | --Select--
|  |  | iii | Where minimum data information is not provided e.g. UNSPSC code. | --Select--
|  | b |  | Clause 3.10 of the Draft Deed of Standing Offer Agreement sets out terms relating to Conditionally Accepted Deliverables. | --Select--
|  | c |  | Products designated as ‘Conditionally Accepted’ may be subject to desktop and functional evaluation by health services. | --Select--
|  | d |  | “Acceptance for use” is confirmed by issuing a product acceptance certificate or any other form signed by an appropriate clinical representative of a Participating Health Service | --Select--
|  | e |  | Pricing will be as per response position and cannot be amended post Agreement acceptance. | --Select--
| 8.2 |  |  | Key Performance Indicators
|  | a |  | Requirements as Agreement - Key Performance Indicators. | --Select--
|  | b |  | The Key Performance Indicators are to be agreed upon between each Participating Health Service and the successful Respondent. | --Select--
|  | c |  | KPI No. 6 is for HSV only, to manage Sales and Stock-on-Hand Reports as part of contract management. If a supplier were to amass significant number of points under this agreement, the result maybe they are excluded from the next agreement, yet to be determined as an evaluation step in the next agreement. | --Select--
| 8.3 |  |  | Reporting
|  | a |  | Requirements as Agreement - Reporting Requirements. | --Select--
| 8.4 |  |  | Service Level Agreement
|  | a |  | Participating Health Services may enter into a Service Level Agreement (SLA) with the successful Respondent(s), on the terms set out in the Draft Agreement clause 4.3 Service Level Agreement. The SLA may cover the following arrangements : | --Select--
|  |  | i | the provision of products on consignment | --Select--
|  |  | ii | requirements for stock management and rotation | --Select--
|  |  | iii | arrangements for ordering, invoicing and delivery | --Select--
|  |  | iv | social procurement commitments or framework set out by a Participating Health Service that may be linked to the Social Procurement Framework set out in the Agreement; and / or | --Select--
|  |  | v | clinical support, including attendance requirements for Representatives in relation to education and training | --Select--
|  |  | vi | communication arrangements for product recalls and safety alerts (refer Part C Section 3 Recall Process). | --Select--
|  | b |  | The SLA will be in addition to the Agreement between the successful Respondent(s) and HSV and will not alter any terms of the Agreement. Any SLA entered into between the Supplier and a Participating Health Service must be established in accordance with the framework of the Agreement and must not contravene or undermine the terms of the Agreement. | --Select--
|  | c |  | HSV will not be responsible for monitoring compliance with any SLA. This is a matter of agreement between the parties to the SLA. | --Select--
|  | d |  | Successful Respondent(s) must provide a copy of all Service Level Agreements/ Order Contracts to HSV within one (1) week of being finalised. | --Select--

## C - General Requirements
Part C: General Requirements
Clause Identifiers |  |  |  | Text |  | Compliance of Tendered Item |  | Compliance of Tendered Item |  | Compliance of Tendered Item |  | Compliance of Tendered Item |  | Compliance of Tendered Item |  | Compliance of Tendered Item
Part C |  |  |  | General Requirements | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Standards and Compliance
|  | a |  | All equipment, consumables and semi-consumables offered must be approved by the Australian Therapeutic Goods Administration (TGA), unless exempt. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Respondents must provide the ARTG number for each tendered item in the Product Description of the Tender Response Worksheet, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The successful Respondent must provide evidence of ARTG certification to Participating Health Services upon request. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Successful Respondents must ensure the ongoing inclusion of all tendered medical products (equipment, consumable, semi-consumable, and accessory) in the ARTG, for the period of this agreement by meeting the associated ongoing annual TGA charges. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Failure by Respondent to maintain the tendered medical product on the ARTG will result in product suspension from HSV Catalogue. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Manufacturer of tendered product (equipment, consumable, semi-consumable, or accessory) must conform with AS ISO 14971 requirements. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | All sterile consumables must have CE / ISO Standard sterilisation certificates. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Respondents must have a certified QMS, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | preference is for AS ISO 13485 certification, otherwise |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | ISO 9001 certification, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | QMS must also include a section on: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • required regulatory documentation (i.e., TGA registration, TGAs Uniform Recall Procedure, etc.), and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • servicing provisions that include full traceability when servicing tendered product. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | All equipment offered must comply with the relevant Australian Standards (or equivalent, Australian recognised International Standard if no current Australian Standard exists), Orders, Legislation and Regulations – collective ‘Compliance Requirement’. Refer to Appendix 2 - Compliance Requirements for a list of the minimum relevant Compliance Requirements. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Tendered products must not cause electromagnetic interference to other medical devices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | All equipment offered must comply with a relevant Electromagnetic Compatibility standard/s. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 1.1 |  |  | Cybersecurity
|  | a |  | Does the tendered product have the capability to connect to the hospital's IT network infrastructure? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Does the tendered product have the capability to connect directly to the Internet (i.e., via public Wi-Fi, internal 4G/5G radio), or connect with additional hardware offered by Respondent? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | If yes to either of the above two questions, then
|  |  | i | Respondents must include the Manufacturer Disclosure Statement for Medical Device Security (MDS2) form (the latest version of form, at time of tender version ANSI/NEMA HN 1:2019) for each tendered product with their submission. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Successful Respondents must provide Participating Health Services the Software Bill of Materials (SBoM) upon their request. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Successful Respondents must provide Participating Health Services the MDS2 form for tendered devices upon their request. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Respondent must ensure that the Manufacturer has a Vulnerability Management Program in place for the total tendered product's life cycle; which monitors, informs Participating Health Services, and addresses security vulnerabilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Respondents must have a Cyber Incident Response Program that will assist a Participating Health Service in the event that the tendered product succumbs to a cyber incursion. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Does the tendered product store or capture patient's personal information? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | If yes to the above question, then Respondent must treat patient's personal data in a secure manner and comply with:
|  |  | i | Commonwealth Privacy Act 1988 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | Victorian Health Records Act 2001 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Victorian Privacy and Data Protection Act 2014 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Victorian Health Records Act 2001 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Does the tendered product have the capability to be connected to, or installed on a computer/server? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Packaging and Labelling
|  | a |  | All tendered products must be labelled in accordance with Essential Principle 13 of Schedule 1 of the Therapeutic Goods (Medical Devices) Regulations 2002. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Any carton/box greater than 15kg must have a "Heavy" label externally adhered on |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All cartons/boxes containing consumables must have:
|  |  | i | the consumable batch number clearly visible externally on each UOM (carton/box), |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | the Product Expiry Date clearly visible on each UOM (carton/box), |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Manufacturing dates clearly visible on each UOM (carton/box), unless otherwise mutually agreed to with relevant Participating Health Service, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | any special storage requirements or instructions clearly visible on each UOM (carton/box) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | instructions on any warnings, restrictions, or precautions that are to be taken in relation to the use of tendered product contained within each UOM (carton/box). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Each individually packaged tendered product must have a label that:
|  |  | i | is visually legible to read |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | in English |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | includes Batch, Lot, or serial number |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | includes expiry date, if applicable |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | classifies the level of reusability (e.g., single use only, single patient, reusable, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | for Category 12 Associated Consumables, has the word 'STERILE', if product sterile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | for Category 12 Associated Consumables, indicating the presence or absence of latex in either the product or its packaging |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | MRI Conditional (for implantable products) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | tracking/traceable identification labels, if applicable |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The packaging for each individually packaged consumable must:
|  |  | i | be easy to open, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | robust to protect the quality of tendered product, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | if the tendered product is supplied as sterile, then
|  |  |  | • openable in a way, that the sterile integrity of tendered product remains unaffected, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • complies with ISO 11607-1 Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • packaged in a manner that protects the contents from contamination during transportation, storage and handling. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Tendered products must be delivered in accordance with the manufacturer’s instructions. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Recall Process
|  | a |  | All recalls must be managed in line with the TGA’s Uniform Recall Procedure for Therapeutic Goods (URPTG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All and any TGA recalls and/or hazard alerts which the Respondent issues, must also be logged on GS1 Australia’s ‘Recall Health’ digital platform. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Successful Respondents should be capable of receiving and accepting Recall Notice Acknowledgement Forms electronically from Participating Health Services via GS1 Recall Health. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • It will be at the discretion of the Participating Health Service whether or not to submit Acknowledge Form via GS1 Recall Health. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All TGA classified Class I recall actions (as defined in TGA’s URPTG) must also comply with the requirements of ‘Part B section 5.1 - Repairs and replacements under warranty‘ located within this document. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All other TGA recall/non-recall actions (other than Class I Recalls), whether software or hardware, must be completed by supplier/contractor within six (6) months of notification otherwise product will be suspended from HPV catalogue. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Essential software upgrades that are not life threatening must be completed within three (3) months of notification, regardless of warranty status. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | As soon as intelligence report arises that a vulnerability within tendered product is being actively exploited by threat actors, Successful Respondent must notify within 72 hours, the Department of Health (Digital Health team), HealthShare Victoria, and Participating Health Services. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Product Information and Identification
|  | a |  | For each different model equipment, Respondent must submit as part of their tender response, a complete set of operator user manuals in portable document format (i.e., .pdf). | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each different model equipment, Respondent must submit as part of their tender response, a complete product brochure in portable document format (i.e., .pdf). | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For each device purchased by Participating Health Service, the successful Respondent must provide at least one (1) complete digital set of operator user manuals. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Additional electronic copies must be provided free of charge upon the request of Participating Health Services. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Successful Respondent must provide, upon the request of Participating Health Service, a 'Quick Reference Guide' for any tendered device. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Successful Respondent must provide a full set of Technical Service Manuals that includes a full list of all replaceable spare parts, upon the request of the Participating Health Service having purchased tendered device. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | All associated consumables and semi-consumables must be listed in a separate document and this must be provided to Participating Health Services upon purchase of each device. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | All consumables and semi-consumables must be identifiable using GTINs, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | must be listed on the National Product Catalogue (NPC). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | the model of device compatible with tendered consumable |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | For all consumables, the manufacturing plant address must be listed in the Tender Response Worksheet |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | For non-OEM consumables, Respondents should submit a Certificate of Validation to which brand and model the consumable is compatible for use with. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | For each tendered equipment, Respondents must provide the following information in the Product Catalogue:
|  | k |  | Brand name / Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Model number |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | ARTG ID number for the device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Dimensions in centimetres |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Weight in kilograms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | warranty period for the tendered device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Participating Health Services may preference equipment with extended warranties |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Participating Health Services may prefer products that still have a 'long time' remaining on the product's life cycle |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the year manufacturer first sold tendered product |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the anticipated year tendered product is expected to reach it’s 'end-of-sales' life |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the average lifespan (of use) of tendered product, in number of years |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Cleaning and Infection Control
|  | a |  | Respondents must list in the Product Catalogue, all the individual components on the tendered equipment that are reusable and require cleaning between patients. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Reprocessing of reusable components must comply with the requirements of AS/NZ 4187 Reprocessing of reusable medical devices in health service, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | only using TGA approved agents on the relevant components of the reprocessing cycle. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Upon the request of a Participating Health Service, successful Respondent must provide reprocessing instructions, including the list of TGA approved agents, prior to purchase of equipment as part of the product evaluation, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | provide the cleaning details of all reusable products when requested by a Participating Health Service, for the life of the product. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | To meet the Australian Guidelines for the Prevention and Control of Infection in Healthcare (2010), medical equipment must be capable of being cleaned with a variety of cleaning agents in order to clean and disinfect the equipment, as prescribed in AS/NZS 4187:2014 Reprocessing of reusable medical devices in health service organizations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Respondents must provide cleaning instructions based on infection control best practice for cleaning and disinfecting all reusable items. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The external covering of all tendered equipment (excluding touchscreen displays), must be capable of being cleaned with one of the following types of disinfectant agents without causing any deterioration to the covering:
|  |  | i | Chlorine and Chlorine compounds, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Hydrogen peroxide, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Quaternary ammonium compounds. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Respondents must ensure there are suitable disinfectant products (including wipes), that don’t cause degradation to the equipment case, available for Participating Health Services to purchase from an Australian supplier, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | a validation certificate should be included Respondent’s submission, stating verified/certified disinfectant for use on tendered equipment that does not cause degradation. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Suitable disinfectant products should be listed on HSV's HSVC2016-124 Hand Hygiene, Disinfectants and Chemical Products agreement |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Respondents must list the disinfectant products suitable for use on tendered equipment in the Product Catalogue, as well as |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | list all suitable cleaning products (e.g., which detergent, % alcohol wipe, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Upon the request of a Participating Health Service, successful Respondents must provide a full list of suitable cleaning and disinfection products (including wipes), which also includes the maximum permissible concentration level of each active ingredient. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | To facilitate cleaning, tendered equipment must have:
|  |  | i | smooth surfaces with no sharp edges, deep crevices or other dirt traps and be able to withstand hospital grade cleaning agents |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | no hook & loop fastener (e.g. Velcro®), and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | surface finishes that are suitable (do not deteriorate) with the continual use of hospital grade cleaning and/or disinfectant agents. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.1 |  |  | Equipment Casing
|  | a |  | The equipment external case must:
|  |  | i | be made from non-porous material that is easy to clean, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | be of a colour that easily shows up contamination, state case colour/s, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | not have any excessive holes, channels or crevices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Materials of Concern
| 6.1 |  |  | Latex
|  | a |  | For each product tendered, Respondents must provide in the relevant columns of the Tender Response Worksheet, the presence of natural rubber latex in either: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The product (yes/no), or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Its accompanying packaging (yes/no). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If product or packaging contains latex, then this must be clearly indicated on the product or its accompanying packaging. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Participating Health Services will preference items that are latex free, unless otherwise indicated. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6.2 |  |  | PVC Plasticisers
|  | a |  | For any tendered product made from polyvinyl chloride (PVC), Respondents must list in the relevant columns of the Tender Response Worksheet, if a plasticisers (i.e., DEHP) has been added to the PVC. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | If the tendered PVC product is flexible, then Respondent must list the plasticiser added to the PVC in the Tender Response Worksheet. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Under normal intended use, plasticiser must not leach/migrate from the tendered PVC product |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Participating Health Service may preference PVC products that have the plasticiser listed on EU’s REACH regulation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6.3 |  |  | Mercury
|  | a |  | Tendered products must not contain Mercury. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Reference Sites
|  | a |  | For each tendered product Respondents must provide upto a maximum of three (3) references with complete and accurate referee contact details including, their position title, phone number/s, and a health service email address. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Referees should either be from a clinical background, technical background or an educator. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Preference will be given where references: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | be one of the Australasian health services listed in Tender Response Worksheet |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Categories 1 & 2 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
- References from a College of Intensive Care Medicine accredited ICU offering 24 months training in intensive care medicine. List of hospitals are provided in the Tender Response Worksheet from the drop down list.
https://www.cicm.org.au/Hospitals/Accredited-Sites-Accordion/Accredited-Units
|  |  |  | • Categories 3, 4 & 6 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
- References from a Royal Australasian College of Physicians accredited training site offering 36 months training in Cardiology. List of hospitals are provided in the Tender Response Worksheet from the drop down list.
https://www.racp.edu.au/trainees/accredited-settings
|  |  |  | • Category 5 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
- References from a Royal Australasian College of Physicians accredited training site offering 18 months perinatal training. List of hospitals are provided in the Tender Response Worksheet from the drop down list.
https://www.racp.edu.au/trainees/accredited-settings
|  |  |  | • Categories 7, 8, 9 & 10 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
- References from an Australian & New Zealand College of Anaesthetists accredited training site offering 36 months anaesthesia training. List of hospitals are provided in the Tender Response Worksheet from the drop down list.
https://www.anzca.edu.au/education-training/training-site-accreditation/anaesthesia-training-sites-and-rotations
|  |  | ii | either an Educator, a Clinician or a Technical person |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | be currently using, or have trialled and evaluated, the product offered in this submission. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• Where usage is
|  |  |  | - equipment: the quantity of the model equipment in use at referee's hospital |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - consumables: the quantity of consumable purchased by referee's hospital for FY23 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Where a product category contains a variety of specific subcategories, Respondents are to ensure that the reference sites provided are representative of the full range of products tendered. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | HSV reserves the right to verify this information with the nominated referees and seek user feedback as to the acceptability of the products. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Respondents must not nominate a referee without their express permission. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Environmental, Social & Governance (ESG)
| 8.1 |  |  | Modern Slavery
|  | a |  | Successful Respondents must comply with the Modern Slavery Act 2018 (Cwth). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8.2 |  | i | Social Procurement Framework
|  | a |  | The Victorian Government released the Social Procurement Framework in April 2018, which HealthShare Victoria and all Participating Health Services apply to their procurement & purchasing decisions; for further details refer to:
|  |  | i | https://www.buyingfor.vic.gov.au/social-procurement-framework-supplier-guidance
|  | b |  | Social and sustainable objectives prioritised under this ITS are:
|  |  | i | Women's equality and safety
|  |  | ii | Environmentally sustainable business practices
|  |  | iii | Environmentally sustainable outcomes
|  | d |  | Respondents should provide details through the Procurement Portal on their approach to these objectives. This information from successful Respondents will be provided to Participating Health Services. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Within twelve (12) month of contract commencement, successful Respondents must have a documented:
|  |  | iii | Family Violence policy |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Gender Equality strategy |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Under the social procurement framework, successful Respondents should be providing opportunities for the following cohorts:
|  |  | i | Victorian First Nations people, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Victorians with a disability, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | disadvantaged Victorians, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | offer apprenticeships or traineeships. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8.3 |  |  | Environmental management systems
|  | a |  | Respondents should have a documented Environmental Management System or Plan in place appropriate to the environmental impacts and risks being managed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Preference may be given to Respondents that have an Environmental Management System independently certified to meet ISO14001 or equivalent standard |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8.4 |  |  | Greenhouse Gas Emissions
|  | a |  | Victoria has legislated a target of Net Zero Carbon by 2050 under the Climate Change Act 2017 (Vic), with interim targets for state-wide emissions to be 28-35% below 2005 levels by 2025, and 45-50% below 2005 levels by 2030. See: https://www.climatechange.vic.gov.au/victorias-climate-change-strategy HealthShare Victoria is looking for awarded suppliers to commit to voluntarily report on their annual Scope 1 and Scope 2 Greenhouse Gas emissions, emissions efficiency metrics, and progress on their own emissions reduction plans.
|  |  | i | Respondents should provide details of any Greenhouse Gas emissions reduction plans and/or targets set by their organisation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents should indicate if they are able to report their Greenhouse Gas emissions and progress on emissions reduction targets to HealthShare Victoria and/or Participating Health Services during the lifetime of the Agreement |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8.5 |  |  | Waste reduction and end-of-life equipment
|  | a |  | Respondents should offer waste reduction initiatives or recycling collection services for applicable products for access by Participating Health Services. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | In the Tender Response Worksheet, list the waste reduction / recycling initiatives on offer to Participating Health Services. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Successful Respondents must have a plan for the removal and environmentally responsible management of tendered equipment from Participating Health Services on the request of the Participating Health Service. Note not all Participating Health Services will require this service. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Any associated costs to this Offer (e.g. a collection charge) should be negotiated under the site Service Level Agreement with the Participating Health Service. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Only equipment supplied by the Successful Respondent is required to be subject to this Offer unless otherwise negotiated between Health Service and Successful Respondent. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Equipment taken-back prior to their end-of-support life may be reconditioned and sold or donated, or disposed of; as appropriate to the specific equipment and circumstances. Preference is that functioning, useful productive and supported equipment, is not disposed of as waste. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Electrical and electronic waste (e-waste) must not be disposed to landfill, and must be sent to a recycler with certification under the e-waste standard AS/NZS 5377:2013. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Successful Respondents must report annually to HealthShare Victoria on the Take Back services provided to Participating Health Services under this Offer using the Reporting Schedule provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8.6 |  |  | Environmental attributes of products and packaging
|  | a |  | For each product tendered, the Respondent must provide the information requested below for the product and its accompanying packaging in the relevant columns of the Tender Response Worksheet or upon request by a Participating Health Service: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | For the primary packaging:
|  |  |  | • The material(s) used |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Weight in grams |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • The percentage of recycled content, if known |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • If the packaging is recyclable |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • If items of packaging are marked with information to assist end users with their waste management (e.g. an Australasian Recycling Label (ARL)) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | For plastic items:
|  |  |  | • The Plastics Identification Code |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • If the Plastics Identification Code is marked on the item. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each product tendered, the Respondent may nominate if the product is an environmentally-preferable product within its category in the relevant column of the Tender Response Worksheet. Respondents must provide additional information as an attachment to substantiate any claim. Respondents may either: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Provide certification by an independent third-party that the product conforms to a relevant environmental specification (an “Eco-label”), which is compliant to ISO Standard 14024 (preferred) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Provide an Environmental Product Declaration for self-declared environmental claims. In this case the Respondent is referred to AS/NZS ISO Standard 14021, ISO 14025, and the Australian Competition and Consumer Commission guidance on Green Marketing and the Consumer Law for general advice. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## D - Product Specifications
Part D: Product Specifications
Clause Identifiers |  |  |  | Text | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Part D |  |  |  | Product Specifications | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Patient Profile
|  | a |  | Respondents must clearly identify the patient profile and limitations of their device (e.g. adult/paediatric/neonatal, minimum/maximum weight etc.). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Power Supply
|  | a |  | All mains powered tendered products must work from standard 240V AC power points, as supplied in Australia. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services may preference devices that operate from a standard mains cable and do not require an external power supply plug pack. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Devices that are powered from an external plug pack should include a solution for securing the low-voltage cable to the 'device'. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Power cord must have an appropriate restrainer/locking feature to prevent unintended disconnection from the device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | UPS that are required to ensure an orderly shutdown or uninterrupted supply, should have ‘hot swappable’ battery capability. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.1 |  |  | Batteries
|  | a |  | All items that use rechargeable batteries, must be provided with rechargeable batteries that do not require cycling to maintain battery capacity. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The expected life of rechargeable battery must be at least two (2) years, with tolerance to partial charging. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State the expected life of the rechargeable battery | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All devices with that can operate on battery power, must have a low-battery indicator. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Rechargeable batteries should have the ability to remain charged for extended periods and should have short recharging times. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | ‘Button’ batteries are not desirable due to being a swallowing hazard. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | A tool must be used to gain access to 'button' batteries | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Disposable batteries (i.e. consumables) must be of a standard alkaline type. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | All rechargeable batteries should include smarts, built-in management systems monitoring remaining usable life (smart batteries), and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | provide a battery imminent end-of-life warning message, with at least 50 remaining cycle counts. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | By Q4 2026, all tendered equipment with rechargeable batteries must include a battery imminent end-of-life alarm message, with at least 50 remaining cycle counts. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Cables and Electrodes
|  | a |  | Cables for ECG monitoring, NIBP, ETCO2 etc. must be of a sufficient length to allow adequate patient access. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A visual and/or audible indicator must be present to indicate lead/cable faults. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All ECG leads and cables must follow the colour coding standards set by the Association for the Advancement of Medical Instrumentation (AAMI)/American Heart Association (AHA). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Excluding portable monitors, all monitors (regardless of whether they are used for transport or are fixed/mounted) must have an optional purchasable cable and/or consumable management system. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Programming
|  | a |  | It must be possible for authorised users to program default start-up parameters. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Password protection is preferred. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Certain settings should be programmable, including initial alarm volume, alarm settings, display format etc. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Consumables and Semi-Consumables
|  | a |  | It is desirable for all tendered products to be compatible with non-proprietary general consumables and semi-consumables (this includes batteries, where applicable). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Consumables and semi-consumables must be available as a separate purchase option (i.e. priced separately to starter packs). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All consumables and semi-consumables (whether OEM or third-party) must be fit for purpose. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Respondents must state which equipment (by manufacturer) their tendered consumable/semi-consumable can be used with. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | All consumables must be intended to be used for a limited time period, or limited number of uses, on a single patient | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Human Factors
|  | a |  | Tendered equipment should include clearly marked switches, audio prompts and other controls, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | with the control switches orderly located and visually easy to follow sequentially. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Any tendered product that contain electronics, must incorporate self-diagnostics and report problems via visual and audio prompts. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Additional quick diagnostic cards may be provided to supplement this. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6.1 |  |  | Alarms
|  | a |  | Tendered device must include a comprehensive graded alarm system for all measured physiological parameters. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device must have the ability for the user to predefined default alarm settings for Adult, Paediatric and neonatal patient populations. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Alarm settings must return to default settings when tendered device is turned off then back on. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Audible alarms must have adjustable volume, but volume must remain sufficient to hear in the presence of moderate ambient noise (65dB). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | It must not be possible to turn off (deactivate) or turn down the volume of an alarm to and in audible level. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | If alarm volume can be muted, the unit must have a visual indicator that clearly identifies when the volume is muted. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | If alarm volume can be muted, the muted alarm should be overridden by the equipment's next critical alarm. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6.2 |  |  | Displays
|  | a |  | Displays should be clear and easy to read from a variety of angles and in a variety of lighting conditions. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | It is desirable for displays to be:
|  |  | i | configurable (including observation cycles) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | in colour | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | auto-dimming. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Displays should be an appropriate size for the type of system offered and its intended application. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Respondents may offer a variety of display formats. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The number of waveforms may be a key purchasing decision for Participating Health Services. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | Occupational Health and Safety
|  | a |  | Tendered equipment, that is considered mobile, must be of an appropriate size and weight for the average nursing staff to handle (lift, lower, carry) the device without requiring high force. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | (5) five-legged roll-stand with a cable/tube management system for the leads/hoses and/or trolley/bed mounts should be offered with equipment greater than 7 kilograms. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Any mobile device with wheels/castors should be of a robust design, with optional braking systems offered. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Any roll-stand offered should incorporate mounting solutions for hygiene products (both hand hygiene, i.e. hand gel, and equipment wipes). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | All equipment/cases must have no sharp edges and no pinch points. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.1 |  |  | Carry Bags
|  | a |  | Each portable tendered equipment should have an optional purpose-designed protective carry bag available for purchase. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Disposable carry bags will be considered, especially for Telemetry Units. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Carry bags offered as a purchase option for physiological monitors, must allow the monitor to be fully operated while inside the bag. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Carry bags should:
|  |  | i | have easily accessible storage compartments for relevant accessories and consumables (such as leads, monitoring electrodes, pads and paper) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | be completely dust-proof | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The carry bag handle should be:
|  |  | i | positioned to ensure that the weight of the case is evenly distributed when carried | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | of a diameter that allows a small adult to be able to grasp the handle so that the thumb can touch any finger of the hand to make a pincher grip. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | If the carry bag has a shoulder strap, this should also be positioned so that the weight (when fully loaded) is evenly distributed when carried and should be appropriately padded to ensure user comfort. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.2 |  |  | Mounting Hardware
|  | a |  | If mounting brackets are offered as an optional accessory, these must have appropriate restraints. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All brackets must be of a design that offers varying height, tilt and pan (exceptions may be given to MRI-Conditional equipment). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Cable management solutions must be offered. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Upon purchase and upon request by a Participating Health Service, a variety of mounting brackets must be available that fit the successful Respondent’s device(s) (such as pendants, anaesthetic machines etc.). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Durability
|  | a |  | All tendered equipment is expected to be durable, and must be able to withstand all tests as prescribed in AS IEC 60601 and sub-standards. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All tendered equipment, whether or not electrical, must withstand all the mechanical tests that are prescribed in AS/NZS IEC 60601.1:2015 Section 15.3 Mechanical Strength (i.e., drop test, etc.). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | state the 'Drop Test' height | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | state the IP rating | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All tendered equipment should have a lifecycle of greater than 7 years. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | All equipment and semi-consumables must be able to endure the use of standard hospital cleaners, and disinfectants, without prematurely aging or damaging the equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9 |  |  | MRI Conditional
|  | a |  | Is the tendered device MRI conditional? (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If yes to the above question, then
|  |  | i | the tendered equipment must be offered with appropriate MRI-Conditional consumables and semi-consumables. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | offered with various mounting solutions | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Preference may be given to MRI-Conditional products that can be used in a 3T (Tesla) environment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the maximum magnetic field strength (T) tendered product can be used in | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the maximum gauss line to which tendered product is specified to operate to | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference may be given to MRI-Conditional monitors that have the ability to ‘gate’ with other devices. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Remote display of patient parameters is highly desirable. Respondents should provide various optional configurations such as display sizes, enunciator options, printing/recording options, etc. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 10 |  |  | Data Requirements
|  | a |  | All devices should, high acuity devices must, be able to communicate and share data with hospital systems. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Barcode readers may be offered as a purchasable option. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For keyboards offered as a purchasable option, preference may be given to keyboards that use standard types of USB connectors (i.e. standard, mini or micro). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Keyboards used in infection sensitive areas must be able to be disinfected using standard hospital disinfectant. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Data integration with HIS (as defined by the Participating Health Service) may be offered as a purchasable option. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Preference may be given to systems that can accept additional data fields (such as weight and height). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Where available, all monitors must record data in a manner that is easily transferrable to the successful Respondent’s data management system/central station, and preferably to other proprietary data management systems, via cable, Bluetooth and Wi-Fi. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Respondent's Cybersecurity team must be able to support Participating Health Service with their Risk Management processes around cybersecurity of networked tendered products | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11 |  |  | Cybersecurity
|  | a |  | Does tendered product include the ability to connect onto the hospital's IT network infrastructure? (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Does tendered product include the ability to connect onto mobile phone (4G/5G) network? (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• If yes to either of the above two questions, then compete all questions in section 11.0
|  | c |  | Any tendered product that can connect onto the hospital's IT network, whether physical or virtual, must:
|  |  | i | have unique user login credentials (both generic and group logins must not be used), and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | have user level of privileges managed, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | must use multi-factor secure user login authentication protocols, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Participating Health Service must be have access to an admin account. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Successful Respondents must have a program in place that informs Participating Health Services of cybersecurity vulnerabilities to the tendered device throughout its life cycle. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Where a Tendered product is installed with an Operating System (OS), the OS should be actively supported and patchable for the full life of the equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | OS patch updates must be installed in affected tendered product within 30 days from patch being released from OS provider. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered product software program (app) must have a patch update installed within 90 days upon an app vulnerability having been identified at no additional cost to Participating Health Services. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Where a vulnerability in a tendered device, whether in the OS or the app, is being actively exploited by threat actors, the Successful Respondents must:
|  |  | i | immediately upon being made aware that threat actors are actively exploiting a vulnerability inform all the following: • all affected Participating Health Services, • Department of Health (Digital Health team), and • HealthShare Victoria; and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | provide Participating Health Services advice on implementing compensating control measures to mitigate the vulnerability being exploited, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | provide a patch update within 48 hours to affected Participating Health Services to remediate exploited vulnerability. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Participating Health Service's Endpoint protection solutions (anti-malware, application control, whitelisting functionality, etc.) must be permitted to be enabled and/or installed on tendered products, with no further on-going additional costs by Respondents to Participating Health Services. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | all endpoint protection solutions for the tendered product must be listed in the Product Catalogue | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | endpoint protection event log should feed into Participating Health Service's cybersecurity systems (i.e., SIEM, event log repository) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | details of the endpoint protection event log must be listed in the Product Catalogue | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | endpoint protection event logs should include: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • a log of valid and invalid logins and logoffs | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • all security type processes | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | All communication ports on tendered product must be disabled by default, and only the ports required in specific install base to be activated via a secure login by the Participating Health Service's IT Department. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Does the tendered product have the capability to utilise multiple security certificates to Participating Health Service's IT network infrastructure? (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | Tendered products must not have hardcoded or default passwords. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Participating Health Services must be permitted to control the use of administrative level privileges on tendered products, if applicable. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Remote Access to any tendered product must:
|  |  | i | be through a secure environment, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | be time-limited; access should time out after 30 minutes, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | be encrypted | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• list encryption method
|  |  | iv | have unique user login credentials (both generic and group logins must not be used), and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | must not use single-factor user login authentication protocols, i.e., must use multi-factor secure user login authentication protocols, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | use the Participating Health Service's existing remote access gateway. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | Respondents must ensure, for any tendered product that connects onto Participating Health Service's ICT network, that the manufacturer routinely performs cyber penetration testing and security assessments on the tendered product. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | How regularly does the manufacturer perform cyber vulnerability assessments/scans on tendered product?
|  |  | ii | How regularly does the manufacturer perform cyber penetration testing on tendered product?
| 11.1 |  |  | Software Updates
|  | a |  | In the event of a software update (or patch), Respondents must:
|  |  | i | provide such updates at no cost, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | perform supplementary in-service training at no cost in the event that the software update changes the operation of the equipment. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Successful Respondents must be obtain written approval from Participating Health Service before installing any software update or upgrade. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## E - Physiological Parameters
Part E: Specifications for Physiological Parameters
Clause Identifiers |  |  |  | Text | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Part E |  |  |  | Specifications for Optional Physiological Parameters | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1.0 |  |  | General Specifications
|  | a |  | Any patient parameters included upon request by a Participating Health Service should be able to be activated and deactivated by authorised users via the unit’s software system. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Password protection must be provided for the purposes of activating and deactivating unit configurations. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.0 |  |  | Neonatal/Paediatric Compatibility
|  | a |  | Electronic tendered products must clearly indicate whether they are in a neonatal or paediatric mode. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | When in either neonatal or paediatric mode, tendered product must automatically adjust set rates, pressures and volumes, or other parameters, to reflect the patient’s weight. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.0 |  |  | ECG Monitoring
|  | a |  | Must be capable of withstanding electrosurgical and other electrical interference without creating undue ECG artifact. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | It must be possible to select and display any available lead. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All leads must have a minimum diagnostic bandwidth of 0.05–100Hz, with the ability to display to screen, print and data storage. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Monitoring bandwidth of 0.5–40Hz is desirable. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Bandwidth must be displayed on the monitor. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | There must be an ECG output for defibrillation synchronisation, IABP slaving etc. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | All monitoring devices must be able to detect and differentiate pacing spikes. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | ECG leads must be AHA colour coded | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Must be capable of performing ST segment analysis on all monitored ECG leads | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Must be capable of performing QT interval analysis on all monitored ECG leads | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | Must provide arrhythmia detection on the following life-threatening conditions:
|  |  | i | Asystole | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Ventricular Fibrillation | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Ventricular Tachycardia | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Supra Ventricular Tachycardia | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Bigeminy | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Tachycardia | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Bradycardia | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.1 |  |  | 3- and 5-Lead ECG
|  | a |  | ESU-protected ECG leads must be available. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Derived 12-lead ECG may be offered as a purchasable option. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.2 |  |  | 12-Lead ECG
|  | a |  | If 12-lead ECG is acquired, a rhythm strip must be provided. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If 12-lead ECG is acquired, it must be diagnostic quality as standard. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | 12-lead ECG interpretation software must be priced separately. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Monitors and data management systems/central stations should be capable of 12-lead ECG interpretation with an arrhythmia analysis feature. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4.0 |  |  | Respiration (Impedance, i.e. from ECG leads)
|  | a |  | Respiratory range must be 0–80 breaths per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Respiratory rate accuracy must be ± two (2) breaths per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Apnoea monitoring is required, and an alarm must sound within 20 seconds of no breath detected. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | In Neonatal mode, must be user adjustable from at least 10 to 20 seconds | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.0 |  |  | Pulse Oximetry (SpO2)
|  | a |  | Pulse oximeter cable should be compatible with only one (1) brand of pulse oximeter probe; the pulse oximeter probe connector on the cable should only be apt for a single brand of probe. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List the brand/s of pulse oximeter probes compatible with tendered equipment
|  |  | ii | List the probe connector type/s on each cable, including number of pins
|  | b |  | SpO2 data must be captured to the event memory, other than for handheld pulse oximeters. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | SpO2 range must be at least 70-100% | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State the SpO2 measurement range that tendered device can measure
|  | d |  | SpO2 accuracy must be within ± 3% at 70–100% SpO2 | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State the SpO2 accuracy for both non-motion and motion artifact
|  |  | ii | State the oximeter-derived Respiratory Rate measurement accuracy for both non-motion and motion artifact
|  |  | iii | State the accuracy of any other oximeter-derived measurements, such as haemoglobin measurements, for both non-motion and motion artifact.
|  | e |  | Pulse rate range must be 30–240 beats per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State the Pulse Rate measurement range that tendered device can measure
|  | f |  | Pulse rate accuracy should be within ± three (3) beats per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State the Pulse rate accuracy for both non-motion and motion artifact
|  | g |  | Tendered device must display the Perfusion Index and/or have a signal strength indicator, other than for handheld pulse oximeters. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Tendered device must offer the user ability to enable an audible variable pitch pulse tone. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Tendered device must display real-time values of:
|  |  | i | SpO2 | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Pulse Rate | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Tendered device must include the following alarms, other than for handheld pulse oximeters.
|  |  | i | high/low SpO2 | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | high/low Pulse Rate | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Sensor failure | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Sensor off | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Low battery | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6.0 |  |  | Non-Invasive Blood Pressure (NIBP) Measurement
|  | a |  | For Acute care monitors, systolic, diastolic and mean arterial pressure (all 3) must be displayed on screen, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | captured to the event memory. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must be able to display either the Heart Rate or Pulse Rate , and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | captured to the event memory for Acute care monitors. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondents must offer both reusable and disposable cuffs for each patient type (i.e. adult, paediatric, neonate and bariatric patients) that their monitor is designed for. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Minimum blood pressure ranges must be as follows:
|  |  | i | systolic:
|  |  |  | adult: 40–270mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | paediatric: 40–180mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | neonate: 40–130mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | diastolic:
|  |  |  | adult: 20–245mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | paediatric: 20–150mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | neonate: 20–100mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | mean:
|  |  |  | adult: 30–255mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | paediatric: 30–160mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | neonate: 30–120mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Measurement time must be less than 40 seconds for a standard adult cuff. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Measurement tolerance must have a mean of ≤ ± 5mmHg with a standard deviation of ≤ 8mmHg over the full range. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Must be able to initiate an immediate single instant measurement. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Must be able to initiate timed interval measurements and continuous repeated measurements, other than for portable NIBP Monitors. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7 |  |  | EtCO2 Monitoring
|  | a |  | EtCO2 numerical data must be displayed on screen, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | captured to the event memory. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For Category 2 tendered devices, when user selected, the EtCO2 waveform must be displayed on-screen. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respiratory range must be 3–80 breaths per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Respiratory rate accuracy must be ± two (2) breaths per minute (bpm) below 60 bpm, ± 3 bpm between 60 and 100 bpm and ±5 above 100 bpm. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | When in use the EtCO2 parameter must be selectable as the method to measure respiration. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | EtCO2 range must be 0–99mmHg in 1mmHg increments, or 0–10%. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | EtCO2 accuracy must be ± 2mmHg for range 0–40mmHg; ± 5mmHg for range 41–70mmHg; ± 8mmHg for range 71–100mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Apnoea monitoring is required, and an alarm must sound within 20 seconds of no breath detected. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | State whether mainstream / side stream / both
|  | j |  | State the side stream delay time
|  | k |  | State the rise time to 90% of final value
|  | l |  | State the EtCO2 warmup time
| 8 |  |  | O2 Monitoring
|  | a |  | O2 monitoring may be offered as a separate option. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | When O2 monitoring is not integrated with anaesthetic gas monitoring, it must comply with the following requirements:
|  |  | i | respiratory range must be 0–80 breaths per minute | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | respiratory rate accuracy must be ± two (2) breaths per minute | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | the warm-up time of the gas monitor/module should be less than five (5) minutes | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | self-calibration is preferred. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9 |  |  | Level of Consciousness
|  | a |  | A minimum of (1) one–channel EEG is required for Anaesthetic Machines, but more channels may be preferred. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Display options must be available (e.g. EEG waveform, spectral). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 10 |  |  | Cardiac Output
|  | a |  | Accuracy must be 0.15 L/min or ± 5% of reading. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Cardiac Output range must be a minimum 1.0 L/min to 15.0 L/min. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Minimum resolution must be 0.1 L/min. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Thermal injectate injection volume must be either 3mL, 5mL or 10mL | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Temperature range of thermal injectate must be between 0°C and 24°C | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Must have the capacity to average 5 or more curves per measurement. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Other measured physiologic parameters should be available. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11 |  |  | Anaesthetic Gas / Integrated Patient Gas Monitoring
|  | a |  | Respiratory range must be 3–80 breaths per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Respiratory rate accuracy must be ± two (2) breaths per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The minimum number of gas waveforms displayed is one, although more are preferred. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The warm-up time of the gas monitor/module should be less than 10 minutes. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Self-calibration is preferred. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered device must accurately measure each of the following gas concentrations over the range that is encountered clinically:
|  |  | i | 0-80% for Nitrous Oxide with an accuracy of 10% of reading or 5 volume %, whichever greater | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | 0-10% for Carbon Dioxide with an accuracy of 10% of reading or 0.4 volume %, whichever greater | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | 0-100% for Oxygen with an accuracy of 5% of reading or 2 volume %, whichever greater | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Tendered device must accurately measure some of the following gas concentrations over the range that is clinically encountered (must have the accuracy for few of the following gases):
|  |  | i | 0-10% for Sevoflurane with an accuracy of 0.25 volume % | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | 0-20% for Desflurane with an accuracy of 0.25 volume % | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | 0-6% for Isoflurane with an accuracy of 0.25 volume % | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | 0-6% for Halothane with an accuracy of 0.25 volume % | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | 0-6% for Enflurane with an accuracy of 0.25 volume % | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Agent monitoring must activate automatically when the device is turned on, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | should automatically identify the administered volatile agent. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Tendered device must eliminate or automatically compensate for presence of water vapour, aspirated fluid, or pressure in the breathing circuit. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Tendered device should only require calibration at six-month intervals. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | Apnoea alarm (associated with CO2 monitoring) must be active by default and only be possible to disable one case at a time. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | FiO2 lower alarm limit for Adults must be no less than 18%. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | FiO2 lower alarm limit for Neonates must be 16%. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | Spirometry
|  | a |  | Respiratory range must be 3–80 breaths per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Respiratory rate accuracy must be ± two (2) breaths per minute. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Accuracy of all measurements, such as rates, volumes, and pressures must be stated in Product Catalogue | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | In spirometry mode, the following must be displayed on monitor:
|  |  | i | tidal volume | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | minute volume | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | PEEP | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | PIP | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | airway respiration rate. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13 |  |  | Invasive Pressure
|  | a |  | A digital readout must be captured to the event memory, when monitored. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | There must be an option of at least one (1) IBP waveform. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must be able to assign labels or appropriately name each pressure | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Minimum range must be -30–300mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | There should be the option to Zero each IBP with a single key stroke. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | There should be an IBP output signal for IABPs. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 14 |  |  | Patient Temperature
|  | a |  | Monitors must display patient temperate in degrees Celsius. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For continuous temperature monitoring, the following is required:
|  |  | i | digital readout must be captured to the event memory, when monitored. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | temperature range must be 0-45°C | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | temperature accuracy must be ± 0.1°C | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For non-continuous temperature monitoring, the following is required:
|  |  | i | oral/axial temperature measurement within 10 seconds | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | rectal temperature measurement within 15 seconds | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | temperature range must be 30–42°C. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | temperature accuracy must be ± 0.2°C | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15 |  |  | Electroencephalograph (EEG)
|  | a |  | Monitors with the EEG option must be capable of displaying a minimum of two (2) channels of EEG and be capable of having a minimum of five (5) electrodes. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | EEG bandwidth must be at least 0.3 Hz – 30 Hz. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Notch filter must be able to be set to 50 Hz. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Sample rate resolution must be at least 0.2 Hz. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 16 |  |  | Recorder
|  | a |  | Strip recorders must provide a diagnostic-quality recording of the event, or when manually activated. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Strip recorders must use thermal paper. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Preference may be given to Respondents who devices can use non-proprietary strip recorder paper. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 17 |  |  | Intravascular Oxygen Saturation (SO2, SvO2, ScvO2)
|  | a |  | A digital readout must be captured to the event memory. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Saturation range must be 60–100%. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Saturation accuracy must be ± 4%. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | There must be a screen display and an audible alarm with adjustable volume, and users must be able to select high and low limits. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 18 |  |  | Transcutaneous Gases (tcpO2, tcpCO2)
|  | a |  | Must have the ability to display the Heart Rate (HR) derived from a pulsatile waveform | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A digital readout must be captured to the event memory, when monitored. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For oxygenation (tcpO2), the following is required: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | tcpO2 range must be 0–400mmHg in 1mmHg increments | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | tcpO2 accuracy must be ± 5mmHg or 10% of reading (whichever is greater) over full range | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | For ventilation (tcpCO2), the following is required: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | tcpCO2 range must be 5–100mmHg in 1mmHg increments | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | tcpCO2 accuracy must be ± 3mmHg or 10% of reading (whichever is greater) over full range | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Must have the following alarms:
|  |  | i | high/low limits for both TcPO2 and TcPCO2 | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | low sensor polarization | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | disconnected sensor | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | unbalanced or inoperative thermistors | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | deviation from preset temperature | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | system malfunction | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Must have T90 response time of within 60 seconds for CO2, and | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | within 50 seconds for O2 | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The monitoring range of the temperature sensor must be 37°C to 45°C | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Sensor must be user-calibratable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The minimum sensor lifetime must be > 14days. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 19 |  |  | Neuromuscular Transmission
|  | a |  | Must use standard 'ECG' electrodes for neuromuscular transmission measurements | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 20 |  |  | Closed-Loop Gas Systems
|  | a |  | To reduce the amount of anaesthetic gases consumed, certain machines may be offered with closed-loop systems (opposed to venting scavenging systems). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Closed-loop gas systems must be able to introduce fresh gas. It is preferable that this can be achieved with a single action (i.e. with no secondary confirmation required). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 21 |  |  | IV Pump Control
|  | a |  | An override feature of the IV pump mechanism must be available. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 22 |  |  | Advanced Ventilation Modes
|  | a |  | If the ventilator can provide Spontaneous Intermittent Mandatory Ventilation (SIMV), it must work from 6–50BPM. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If the ventilator can provide Continuous Positive Airway Pressure (CPAP) as a ventilation mode, the pressure delivered must be 5–30cm/H2O. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Other ventilation modes must be clearly displayed on screen. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 23 |  |  | Independent Air Supply
|  | a |  | Anaesthetic machines that provide their own medical air (whether through compressor or turbine) must clean the air to a 0.3 micron level. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The independent air supply must be of a sufficient capacity to ventilate a patient without relying on bottled medical air. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 24 |  |  | Aneroid Sphygmomanometers
|  | a |  | Respondents must offer both reusable and disposable cuffs for each patient type (i.e. adult, children and bariatric patients) that their sphygmomanometer is designed for. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Blood pressure range must be as follows:
|  |  | i | inflation pressure range: 20–290mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | resolution: 2mmHg. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Accuracy must be ± 4mmHg or ± 2% of reading (whichever is greater) over full range. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | - | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 25 |  |  | Foetal Ultrasound (FHR)
|  | a |  | A waveform must be recorded onto paper, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | captured to the event memory, when monitored. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Highly sensitive ultrasound transducer, with a transmit frequency of 1 to 2.5 MHz | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | FHR range must be at least 50–210 beats per minute, | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | with a resolution of 1 bpm | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Must have the following alarms:
|  |  | i | high/low limits for foetal ECG | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 26 |  |  | Uterine Activity Monitoring
|  | a |  | A waveform of uterine activity must be recorded onto paper, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | captured to the event memory, when monitored. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For Toco transducer monitoring, the following is required:
|  |  | i | measurement range must be 0–100 relative units. State relative unit measurement range and sensitivity of each unit (in grams). | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | waveform must have a reset baseline feature to 20 relative units | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | maternal pulse rate measurement (if capable) must:
|  |  |  | • pulse rate range must be 30–240 beats per minute | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • pulse rate accuracy should be within ± three (3) beats per minute | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For intrauterine pressure monitoring, the following is required:
|  |  | i | pressure range must be 0–100mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | must have a reset waveform baseline feature to 0mmHg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 27 |  |  | Direct Foetal and Maternal ECG Monitoring (FECG, MECG)
|  | a |  | Heart rate range must be 30–240 beats per minute | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Heart rate accuracy must be ± 1bpm or 1% of reading (whichever is greater) over full range | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For FECG, a waveform must be recorded onto paper or captured to the event memory, when monitored. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | For MECG, a digital readout must be captured to the event memory, when monitored. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 1-VitalSigns
Category 1: Vital Signs Monitoring
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 1.0 |  |  |  | Vital Signs Monitoring | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | A device intended for either manual or automated measuring and/or display of one or more of the vital signs (e.g. oxygen saturation in the blood (SpO2), heart/pulse rate, non-invasive blood pressure, respiratory rate, temperature). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | May be a Vital Signs Monitor
|  |  | i | A device intended all to electronically measure and display of two or more patient physiological parameters for intermittent spot-checking and/or quasi-continuous real-time monitoring. These monitors are typically portable or mobile units that facilitate movement from one location to the other and may be connected to mains power and/or powered by internal batteries. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | May include Pulse Oximetry
|  |  | i | A device intended for the transcutaneous measurement of oxygenated and deoxygenated haemoglobin by measuring light absorbance changes resulting from arterial blood flow pulsations and displaying the results as percent saturation of haemoglobin in arterial blood (SpO2); and may calculate/display other parameters, e.g. pulse rate. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | May be a manual Aneroid Sphygmomanometer
|  |  | i | A device designed to measure blood pressure consisting of an inflatable cuff that fits around a limb, and inflation bulb for controlling the air pressure within the cuff and an aneroid manometer. The manometer may be mounted to a wall, placed on a table, or hand-held (portable); blood pressure measurement is taken in conjunction with a stethoscope. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | May include electronic Sphygmomanometry
|  |  | i | A device intended to noninvasively measure blood pressure using a self-contained software program that regulates automatic arm-cuff inflation and measurement cycles. The device typically displays current heart rate and mean arterial pressure in addition to systolic and diastolic blood pressures. It may have memory to store blood pressure values and may sound alarm if blood pressure exceeds preset limits. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | May include Capnography
|  |  | i | A device intended to continuously measure and display the carbon dioxide content in a patient’s exhaled gas, i.e. to measure and-tidal carbon dioxide (EtCO2) in exploratory gases. The sampling method can be “mainstream” and/or “side stream”. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | May consist of wearable sensors
|  |  | i | a medical-grade lightweight wireless and wearable adhesive biosensor 'patch' that measures a number of vital signs continuously |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Base Unit
|  | d |  | Requirements for each physiological measure must be met, as outline in 'Part E: Specifications for Physiological Parameters'. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | a |  | Tendered device must be capable of monitoring Adult population |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device should be capable of monitoring:
|  |  | i | paediatric population |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | neonatal population |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The tendered product must be able to non-invasively measure at least one of the following physiological parameters:
|  |  | i | Haemoglobin-oxygen saturations |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Heart/Pulse rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Blood pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Respiratory rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Temperature |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Others |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• if so, then list
|  | d |  | Respondents are to note that manual sphygmomanometers are also in scope. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.1 |  |  | Size and Weight
|  | a |  | If the tendered product is designed to be used as a portable device (used without a roll-stand), then its total weight (including batteries, cables and carry bag) must not exceed seven (7) kilograms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.2 |  |  | Transport Solutions
|  | a |  | Tendered products greater than seven (7) kilograms must be offered with an optional roll-stand solution. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.3 |  |  | Power Supply
|  | a |  | Electronic tendered devices must be capable of continuously operating for a minimum of 60 minutes on battery only power, with all available physiological parameters being monitored. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered devices with batteries must have a low battery indicator. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered devices that operate from standard 240V AC power points, as supplied in Australia, should: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | have an internally rechargeable battery, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | be charged by the power supply within tendered device, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | rechargeable battery have a lifespan of at least two (2) years, tolerating plentiful discharges. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Data Management
|  | a |  | Data capture and transmission to other systems (such as patient physiological central monitoring systems) is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Additional configurable parameters, such as weight, pain etc. are desirable for manual entry |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Monitors should record and store multiple events |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free upon purchase. Details of all conditions and subsequent annual costs must also be provided pre-purchase. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
See Category 3.06: Data Management Systems and Central Stations for further data management requirements.
| 3.1 |  |  | Data Transmission
|  | a |  | Data transfer from tendered equipment must be by one of the following industry standard connection types:
|  |  | i | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state Wi-Fi technology used
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | iii | IR connection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Bluetooth |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • the operator should be able to disabled/enabled the Bluetooth connection in the defibrillator settings |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | USB drive |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Other, please specify
|  | b |  | Data transmission must ensure that the privacy of the patient's personal information is not breached. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All data output ports must be activated and ready for use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference may be given to vital signs monitors capable of bidirectional communication. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.2 |  |  | Data Recording
|  | a |  | Tendered equipment with optional data recording must incorporate date and time recording of patient observations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered equipment should be able to printout any captured sequences of events during use on a patient (including the times of events, waveforms etc.). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Preference may be for an integrated recorder |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Captured data on Category 1 tendered devices should be easily exported for printing, to file into medical records. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.01 - Pulse Oximeters, Handheld
| 4 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 5 Pulse Oximetry. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered pulse oximeter must be portable handheld sized, suitable for rapid spot checks of SpO2%. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Total weight, including batteries, probe and case, must not exceed one (1) kilogram. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must be battery operated, either using standard alkaline batteries or a rechargeable battery. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Should be capable of monitoring Paediatric and Neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered device must simultaneously display real-time values of both SpO2% and Pulse Rate. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Tendered device must not use button batteries |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each pulse oximeter offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | oximetry technology brand utilised within device, e.g. Masimo®, NellcorTM, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | patient population profile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | the SpO2% measurement range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • accuracy within the various measurement ranges for both non-motion and motion artifact |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Pulse Rate measurement range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • accuracy for both non-motion and motion artifact |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Pleth waveform displayed (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | List all other monitored patient physiological parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • the number of expected readings per battery life. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Battery life if in continual use (hours) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Automatic turn off after (sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.02 - Pulse Oximeters, Desktop
| 6 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 5 Pulse Oximetry. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device must operate from mains power |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered device must be capable of monitoring all patient populations, i.e. Adult, Paediatric and Neonatal. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The display monitor on the pulse oximeter must also show the SpO2 waveform when a probe is connected to a patient. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6.1 |  |  | Optional Parameters
|  | a |  | Respondents are encouraged to offer EtCO2 monitoring as a purchasable option, if available. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Note: additional monitored parameters are described in greater detail under Part E: Specifications for Optional Physiological Parameters.
| 7 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each pulse oximeter offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | oximetry technology brand utilised within device, e.g. Masimo®, NellcorTM, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | patient population profile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • waveforms shown |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | the SpO2 measurement range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • accuracy within the various measurement ranges for both non-motion and motion artifact |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Pulse Rate measurement range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • accuracy for both non-motion and motion artifact |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • battery capacity (no. hrs in continuous monitoring) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.03 - Aneroid Sphygmomanometers, handheld
| 8 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 24 Aneroid Sphygmomanometers. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The sphygmomanometer must consist of an inflatable cuff (interchangeable with various sized cuffs) a measuring unit (aneroid gauge) and manually operated inflation bulb with valve. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The aneroid gauge may be offered as either a stand-alone gauge, a gauge on the cuff, or with the bulb as one assembly. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered sphygmomanometer must be provided with a small android gauge. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Must be suitable for use with all patient populations: Adult, Paediatric and Neonatal. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Respondent must provide information on the sphygmomanometer measurement range in the product catalogue |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each aneroid sphygmomanometer offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of gauge, diameter (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g), including all items (cuff, bulb, tubes, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | carry case included (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.04 - Aneroid Sphygmomanometers, large dial
| 10 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 24 Aneroid Sphygmomanometers. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The sphygmomanometer must consist of an inflatable cuff (interchangeable with various sized cuffs) a measuring unit (aneroid gauge) and manually operated inflation bulb with valve. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | Tendered sphygmomanometer must be provided with a large android gauge. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must be suitable for use with all patient populations: Adult, Paediatric and Neonatal. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Respondent must provide information on the sphygmomanometer measurement range in the product catalogue |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered sphygmomanometers should be offered as either a wall, roll stand or desktop option. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each aneroid sphygmomanometer offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of gauge, diameter (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g), including all items (cuff, bulb, tubes, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | configuration (wall mount, roll stand, desktop, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.05 - NIBP monitors, portable
| 12 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 6 Non-Invasive Blood Pressure (NIBP) Measurement. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered NIBP monitor must be suitable for use and location on a desktop, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | intended to be portable carried by hand between locations, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | total weight, including batteries, should not exceed two (2) kilograms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Should be capable of monitoring Paediatric and Neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The unit must auto inflate cuff to occlusion pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each NIBP monitor offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | patient population profile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Maximum cuff inflation pressure (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | measurement time (sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Systolic pressure measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Diastolic pressure measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • battery capacity (no. of measurements) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.06 - NIBP monitors
| 14 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 6 Non-Invasive Blood Pressure (NIBP) Measurement. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device must operate from mains power |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered device must be capable of monitoring all patient populations, i.e. Adult, Paediatric and Neonatal. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered device should be offered with an optional roll stand accessory. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The unit must auto inflate cuff to occlusion pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each NIBP monitor offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | patient population profile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Maximum cuff inflation pressure (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | measurement time (sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Systolic pressure measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Diastolic pressure measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.07 - End-tidal CO2 Monitors
| 16 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 7 EtCO2 Monitoring. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device must operate from mains power. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered device must be capable of monitoring all patient populations, i.e. Adult, Paediatric and Neonatal. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 17 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each End-tidal CO2 monitor offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | patient population profile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Sensor type, i.e. Mainstream/Side stream/Micro-stream |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • sampling flow for side-stream systems (ml/min) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Response time (msec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | CO2 measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.08 - Vital Signs Monitors
| 18 |  |  | General Requirements
|  | a |  | Tendered device must comply with all relevant requirements stated in Part E Optional Physiological Parameters. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device must operate from mains power. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered device must be capable of monitoring all patient populations, i.e. Adult, Paediatric and Neonatal. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered device must measure blood pressure noninvasively and also one of the following physiological parameters:
|  |  | i | pulse oximetry (SpO2) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | pulse rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | etCO2 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Respiratory Rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | non-continuous patient temperature monitoring. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • patient temperature monitoring must be non-continuous |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tendered device should be offered with an optional roll stand accessory. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered devices weighing more than seven (7) kilograms must be provided with a roll-stand |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 18.1 |  |  | Optional Configurations
|  | a |  | Desirable for vital signs monitors to include the following:
|  |  | i | manual entry of other patient physiological parameter values |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | manual entry of patient observations and assessments |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | interface capabilities with other ICT systems |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | configurable/customisable Early Warning Scoring (EWS) protocols |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | display patient's trends in either tabular or graphical form |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | offer other additional patient enhancement software packages |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 19 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each End-tidal CO2 monitor offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | patient population profile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | NIBP
|  |  |  | • maximum cuff inflation pressure (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • measurement time (sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • systolic pressure measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • diastolic pressure measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Pulse Oximetry
|  |  |  | • oximetry technology brand utilised within device, e.g. Masimo®, NellcorTM, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • the SpO2% measurement range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - accuracy within the various measurement ranges for both non-motion and motion artifact |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Pulse Rate measurement range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - accuracy for both non-motion and motion artifact |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Pleth waveform displayed (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Thermometer
|  |  |  | • probe technology, e.g. tympanic, oral, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • brand, e.g. Exergen, Genius, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • response time (sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • temperature measurement range (°C) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | End-tidal Carbon Dioxide monitoring
|  |  |  | • sensor type, i.e. Mainstream/Side stream/Micro-stream |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - sampling flow for side-stream systems (ml/min) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • response time (msec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • CO2 measurement range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Early Warning Scoring protocols included |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | List physiological monitored parameters that can be manually entered |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | List ICT systems device can interface with |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xix | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xx | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 1.09 - Wearable Devices for Continuous Vital Signs Monitoring
| 20 |  |  | General Requirements
|  | a |  | Requirements for each physiological measurement must be met, as outline in 'Part E: Specifications for Physiological Parameters'. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 21 |  |  | Product Information
|  | a |  | Must include a product brochure in ITS response for each tendered device | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each wearable device, Respondent must submit as part of their tender response, a complete set of operator user manuals in portable document format (i.e., .pdf). | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 2-AcuteCare
Category 2: Acute Care Monitoring
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 2.0 |  |  |  | Acute Care Monitoring | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Vascular Dopplers, Blood-flow detectors
|  |  | i | generally hand-held battery-powered devices, that use Doppler effect techniques to determine transcutaneously the extent of peripheral vascular blood flow. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Bladder Scanners
|  |  | i | Ultrasonic scanning systems intended for the transabdominal examination of the bladder. These systems perform an analysis of the generated ultrasound images automatically calculating measurements such as bladder volume and post-void residual urine volume. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Mid-Acuity Monitors
|  |  | i | 12 inch or less display monitor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | A bedside patient monitor designed to acquire, measure, display, monitor, record/store, and in some cases analyse information related to one or more physiologic parameters reflecting a patient’s clinical condition; by externally connected cables/leads to appropriate sensors (e.g. electrodes, catheters, probes, transducer sets) attached to or inserted in the patient. Allows users to change, alter or add monitored parameters (i.e. adding additional modules) for an individual patient. Typically connects to a central station monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | High-Acuity Monitors
|  |  | i | 15 inch or more display monitor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | A bedside patient monitor designed to acquire, measure, display, monitor, record/store, and in some cases analyse information related to one or more physiologic parameters reflecting a patient’s clinical condition; by externally connected cables/leads to appropriate sensors (e.g. electrodes, catheters, probes, transducer sets) attached to or inserted in the patient. Allows users to change, alter or add monitored parameters (i.e. adding additional modules) for an individual patient. Typically connects to a central station monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Patient Monitoring System Modules
|  |  | i | A small unit designed to be used with and enhance functionality of the bedside monitor dedicated to the measurement of a single or several physiologic parameters. The module automatically plugs into the parent device when placed into a standardised slot or a connected module rack; the parent device operates as a mainframe displaying the parameter information measured by the provided module. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Includes devices/modules that have incorporated a small display monitor which may also be used as a stand-alone monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | MRI Monitors
|  |  | i | A bedside patient monitor designed for continuous assessment of several vital physiologic parameters (e.g. ECG, blood pressure, heart rate, temperature, cardiac output, apnoea, and respiratory/anaesthetic gas concentrations) of infant to adult intensive care patients during a magnetic resonance imaging (MRI) procedure. It includes MRI safe sensors/electrodes, and typically with an additional remote monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Telemetry Units
|  |  | i | A device designed to continuously measure and wirelessly transmit typically electrocardiographic (ECG) signals, and may include other vital signs (e.g. non-invasive blood pressure, temperature, oxygen saturation, etc.), from a patient to a receiving location (e.g. central station, bedside monitor) for viewing. Typically used for ambulatory patients who require continuous monitoring (e.g. patients in a step-down care area). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Central Station Monitors
|  |  | i | A device that continuously receives and displays physiological data from one or more bedside patient monitors and/or telemetry units and displays this information at a central patient monitoring station where staff can monitor many patients simultaneously. It produces visible and audible alarms when adverse conditions are registered. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Patient Monitoring Information Management Systems
|  |  | i | A dedicated assembly of devices (hardware and/or software) designed to electronically receive, collect, store, manage, assist in the analysis of, display, output, and distribute data, within healthcare facilities, to support the administrative and clinical activities associated with the information acquired on the central station; whether this is web-based RDS clinical support, connectivity to eMR/ADT systems, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Patient Monitoring Asset Condition Management Software
|  |  | i | An application software program intended for use to electronically receive, collect, store data, to manage the maintenance of the physiological monitoring systems (bedside monitors and Central stations). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | Transcutaneous Partial Pressure Monitors
|  |  | i | A bedside monitor designed to continuously measure the partial pressure of carbon dioxide (tcpCO2) and/or oxygen (tcpO2) at the skin surface. The device includes a special sensor that is typically heated to increase the blood flow and gas diffusion at the point of application on the skin surface; results are digitally displayed and/or printed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Cardiac Output Monitors
|  |  | i | A device designed for continuous or intermittent measurement and display of cardiac output and other haemodynamic variables from one or more physiologic non-invasive, minimally invasive, or invasive sensors. Cardiac output monitors usually measure and display the cardiac output values, cardiac output index, stroke volume, venous oxygen saturation, systemic vascular resistance, blood pressure, temperature, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Regional Oximetry Monitors
|  |  | i | A device that noninvasively measures regional oxygen saturation in the deep intracranial microvasculature of the brain, which primarily represents cerebral venous saturation (SvO2). The technique used is like that used in pulse oximetry except arterial pulse is not measured. The probes are similar to pulse oximetry but instead are placed on the scalp/head. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | Respiratory Gas Monitors
|  |  | i | A device that continuously and simultaneously samples, measures, and displays concentrations on respiratory and administered anaesthetic gases during general peri-operative and intraoperative surgical procedures. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | Depth-of- Anaesthesia Monitors
|  |  | i | A device designed for continuous intraoperative measurement and analysis of the electroencephalogram waveform and its other features enabling detection of, e.g. levels of sedation, loss/level of consciousness. It is used during anaesthesia in administration and in trauma. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Data Management
| 2.1 |  |  | Data Transmission
|  | a |  | Tendered device should be able to transfer data via one of the following connection types: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state Wi-Fi technology used
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | iii | IR connection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Bluetooth |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • the operator should be able to disabled/enabled the Bluetooth connection in the defibrillator settings |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | USB drive |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Other, please specify
|  | b |  | Data transmission must ensure that the privacy of the patient's personal information is not breached. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All data output ports must be activated and ready for use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference may be given to acute care monitors capable of bidirectional communication. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.2 |  |  | Data Recording and Management
|  | a |  | Data capture and transmission to other systems (such as patient physiological central monitoring systems) is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered devices that capture/record data must date and time stamp each captured/recorded event. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered devices should be capable to record/store/capture multiple events. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Additional configurable parameters, such as weight, pain etc. are desirable for manual entry |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tendered devices should be able to printout any captured sequences of events during use on a patient (including the times of events, waveforms etc.). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Preference may be for an integrated recorder |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Captured data on tendered devices should be easily exported for printing, to file into medical records. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | This must be achieved on both mid- and high-acuity monitors without the use of ethernet, Wi-Fi, Bluetooth, or central monitoring connectivity. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free upon purchase. Details of all conditions and subsequent annual costs must also be provided pre-purchase. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.3 |  |  | Point-of-Care Integration
Point-of-care integration is an important feature for some Participating Health Services. This feature provides all patient-related information to the bedside for timely decision making for clinical staff (e.g. path results, PACS, HIS).
|  | a |  | Point-of-care integration systems must be able to: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | follow standard TCP/IP protocols |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | interface with Participating Health Services’ intranet. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Preference may be given to systems that have proven records of integration with path systems, PACS systems etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
See Category 3.06: Data Management Systems and Central Stations for further data management requirements.
Subcategory 2.01 - Vascular Dopplers, Blood flow Detectors
| 3 |  |  | General Requirements
|  | a |  | Tendered device must be capable of detecting blood flow. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Should be capable of monitoring Paediatric and Neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must be battery operated, either using standard alkaline batteries or a rechargeable battery. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Must be portable and handheld sized, suitable for rapid spot checks. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Must include an integrated speaker, with |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | user selectable volume control of Doppler flow. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Must offer multiple frequency blood-flow detectors, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | adjustable frequency blood-flow detector. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each vascular Doppler blood flow detector offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | probe frequency/ies (MHz) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Continuous-wave or pulsed-wave Doppler system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Bidirectional detection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List information that is displayed on screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | headphone jack |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • list other type of output connectors |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | integrated printer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Interfaces to a computer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xix | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.02 - Bladder Scanners
| 5 |  |  | General Requirements
|  | a |  | Bladder scanner must include a setting for select gender. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Bladder scanner must use 3D ultrasonography technology |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Scanning angles must be greater than 120 degrees |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Bladder volume calculation accuracy must be at least ±15mL. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Scanner must have at least 12 scan planes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Scan time must be less than 5 seconds |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Scan penetration depth must be more than 15 centimetres |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Bladder scanner must be able to save each performed scan. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each bladder scanner offered, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • console |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • probe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • console + probe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • probe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | patient population profile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) if relevant |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • resolution, pixels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List information that is displayed on screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | State memory capacity, maximum number of scans |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | State how data is download from machine |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | integrated printer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | configuration (roll stand, desktop, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | probe "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | IPX rating, probe (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | IPX rating, console (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.03 - Mid-Acuity Monitors
| 7 |  |  | General Requirements
|  | a |  | The minimum base unit configuration for acute care monitors are:
|  |  | i | colour display |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | ECG monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | pulse oximetry (SpO2) monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Non-Invasive Blood Pressure (NIBP) measurement. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must be capable of monitoring both Adult and Paediatric populations, while Neonatal population is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must have the ability to display the Heart Rate (HR) derived from a pulsatile waveform other than the ECG. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Must have the ability to trend (both tabular and graphical) all parameters for no less than 24 hours |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Participating Health Services may preference mid-acuity monitors that provide bed to bed overview monitoring. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.1 |  |  | Power Supply
|  | a |  | Batteries must operate continuously for a minimum period of 60 minutes with all optional parameters in use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For acute care monitors provided with a battery, Respondent must be able offer Participating Health Services an external battery charger to recharge battery. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.2 |  |  | Size and Weight
|  | a |  | Portable bedside monitors should not exceed a total weight (including any auxiliary equipment) of seven (7) kilograms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Mobile bedside monitors that exceed the total weight of seven (7) kilograms must be offered with an optional roll-stand accessory. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Bedside monitors must be a suitable size and weight to be compatible with wall mount option provided by Respondent. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.3 |  |  | Display
|  | a |  | Display screen should not be bigger than 12 inches. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | When ECG is connected, ECG waveform must be displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The layout on the display must be user configurable; waveform order, numerical field order |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.4 |  |  | Self-Diagnostics
|  | a |  | The ability for the monitor to provide alerts when accessed or unserviceable is optional. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.5 |  |  | Optional Patient Parameters
|  | a |  | It is desirable that acute care monitors can include the following patient parameters upon request by Participating Health Services:
|  |  | i | 12-lead ECG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | respiration rates (acquired via ECG leads) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | EtCO2 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | level of consciousness |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | cardiac output |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | anaesthetic gas |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | spirometry |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | invasive pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | continuous patient temperature |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | electroencephalograph (EEG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | neuromuscular transmission |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | rhythm strip recording (paper) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | intravascular oxygen saturation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.6 |  |  | Optional Configurations
|  | a |  | It is desirable that acute care monitors can include the following configurations upon request by Participating Health Services:
|  |  | i | the ability to display various numbers of waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | touch screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | various display sizes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | slave/remote displays |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | battery |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | patient profile selection (such ICU, neonate, anaesthesia etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | software packages (such as metabolic, spectral analysis etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | interface capabilities with other ICT systems |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | connectivity to other devices (e.g. ventilators, monitoring solutions) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | point-of-care integration solutions |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | local (to monitor) physiological parameter trending |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | data capture and transmission to other systems (such as patient physiological central monitoring systems). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | ST Mapping |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Horizon Trends |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.7 |  |  | Data Management
|  | a |  | Section 2.0 Data Management above has been completed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Refer to subcategory 3.06 Central Station Monitors for further data management requirements |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Bedside monitors must record and store multiple events along with patient details and ECG information. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • number of selectable ECG leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Describe how captured data is able to be printed from device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • list all physiological monitored parameters that can be numerically printed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • list all physiological monitored waveforms that can be printed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • detected arrhythmias |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Number of additional modules able to be connected |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • additional physiologic parameters that can be monitored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.04 - High-Acuity Monitors
| 9 |  |  | General Requirements
|  | a |  | The minimum base unit configuration for acute care monitors are:
|  |  | i | colour display |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | ECG monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | pulse oximetry (SpO2) monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Non-Invasive Blood Pressure (NIBP) measurement. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must be capable of monitoring both Adult and Paediatric populations, while Neonatal population is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must have the ability to display the Heart Rate (HR) derived from a pulsatile waveform other than the ECG. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Must have the ability to trend (both tabular and graphical) all parameters for no less than 24 hours |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Networked bedside monitors must support a bed to bed monitoring overview capability |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.1 |  |  | Power Supply
|  | a |  | Batteries must operate continuously for a minimum period of 60 minutes with all optional parameters in use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For acute care monitors provided with a battery, Respondent must be able offer Participating Health Services an external battery charger to recharge battery. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.2 |  |  | Size and Weight
|  | a |  | Portable bedside monitors should not exceed a total weight (including any auxiliary equipment) of seven (7) kilograms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Mobile bedside monitors that exceed the total weight of seven (7) kilograms must be offered with an optional roll-stand accessory. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Bedside monitors must be a suitable size and weight to be compatible with wall mount option provided by Respondent. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.3 |  |  | Display
|  | a |  | Display screen should not be small than 15 inches. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | When ECG is connected, ECG waveform must be displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The layout on the display must be user configurable; waveform order, numerical field order |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.4 |  |  | Self-Diagnostics
|  | a |  | The ability for the monitor to provide alerts when accessed or unserviceable is optional. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.5 |  |  | Optional Patient Parameters
|  | a |  | It is desirable that acute care monitors can include the following patient parameters upon request by Participating Health Services:
|  |  | i | 12-lead ECG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | respiration rates (acquired via ECG leads) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | EtCO2 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | level of consciousness |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | cardiac output |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | anaesthetic gas |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | spirometry |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | invasive pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | continuous patient temperature |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | electroencephalograph (EEG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | neuromuscular transmission |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | rhythm strip recording (paper) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | intravascular oxygen saturation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.6 |  |  | Optional Configurations
|  | a |  | It is desirable that acute care monitors can include the following configurations upon request by Participating Health Services:
|  |  | i | the ability to display various numbers of waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | touch screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | various display sizes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | slave/remote displays |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | battery |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | patient profile selection (such ICU, neonate, anaesthesia etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | software packages (such as metabolic, spectral analysis etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | interface capabilities with other ICT systems |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | connectivity to other devices (e.g. ventilators, monitoring solutions) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | point-of-care integration solutions |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | local (to monitor) physiological parameter trending |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | data capture and transmission to other systems (such as patient physiological central monitoring systems). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | ST Mapping |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Horizon Trends |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.7 |  |  | Data Management
|  | a |  | Section 2.0 Data Management above has been completed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Refer to subcategory 3.06 Central Station Monitors for further data management requirements |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Bedside monitors must record and store multiple events along with patient details and ECG information. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 10 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • number of selectable ECG leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Describe how captured data is able to be printed from device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | list all physiological monitored parameters that can be numerically printed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | list all physiological monitored waveforms that can be printed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • detected arrhythmias |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Number of additional modules able to be connected |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • additional physiologic parameters that can be monitored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) (where available) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.05 - Patient Monitoring System Modules
| 11 |  |  | General Requirements
|  | a |  | Tendered module must function with the
|  |  | i | tendered mid-acuity monitor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | tendered high-acuity monitor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All accessories required for the module to function with a bedside monitor must be offered as part of this ITS. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered module in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • number of selectable ECG leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) (where available) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of monitored parameter numerical fields |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Module connects to monitor via a rack mount or plugs into monitor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) (where applicable) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.06 - MRI Monitors
| 13 |  |  | General Requirements
|  | a |  | Tendered device must be:
|  |  | i | MR safe, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | MR Conditional. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | MRI monitor must meet all requirements as per section:
|  |  | i | 3.01 Mid-Acuity Monitors, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | 3.02 High-Acuity Monitors |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondent must offer any and all required MR safe accessory and/or consumable required for device to function as intended in a MRI room |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All patient leads/cables/electrodes must be almost MR safe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | MRI monitors must be offered with a remote slave display for the Control Room |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 14 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered module in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | state the maximum magnetic field strength (T) tendered product can be used in |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | state the maximum gauss line to which tendered product is specified to operate to |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • number of selectable ECG leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Display size, H x W (mm) (where available) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of monitored parameter numerical fields |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Module connects to monitor via a rack mount or plugs into monitor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | connects with a Central Station when in use in MRI Room |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) (where applicable) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.07 - Telemetry Units
| 15 |  |  | General Requirements
|  | a |  | The minimum base unit configuration for telemetry units are:
|  |  | i | 3-lead ECG monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | battery-operated transmitter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | receiver unit with a display that uses mains supply power. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15.1 |  |  | Size and Weight
|  | a |  | Telemetry/transmitter units should weigh less than 500g. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Telemetry/transmitter unit dimensions should not exceed 200mm x 100mm x 50mm. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15.2 |  |  | Power Supply
|  | a |  | Rechargeable batteries must operate continuously for a minimum period of 12 hours with all optional parameters in use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | External battery charger units must be offered to allow spare rechargeable batteries to be charged. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Disposable batteries must last a minimum of 24 hours with all parameters in use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Telemetry/transmitters units must have a battery life indicator on the unit. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | It is desirable for the receiver unit (i.e. central station) to indicate state of charge of each transmitter unit. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15.3 |  |  | Display (where available)
|  | a |  | Telemetry/transmitter units with a display must have: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | a battery life indicator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a lead quality indicator. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15.4 |  |  | Durability
|  | a |  | Preference may be given to telemetry/transmitter units (without covers) that are water-resistant. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15.5 |  |  | Optional Patient Parameters
|  | a |  | It is desirable that telemetry units can include the following patient parameters upon request by Participating Health Services:
|  |  | i | 5- or more ECG lead capability |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | SpO2 monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | NIBP. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Note that patient parameters are described in greater detail under Part E: Specifications for Optional Patient Parameters.
| 15.6 |  |  | Optional Configurations
|  | a |  | It is desirable that telemetry/transmitter units can include the following configurations upon request by Participating Health Services:
|  |  | i | the ability to display various numbers of waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | various display sizes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | patient profile selection (such ICU, neonate, anaesthesia etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | interface capabilities with other ICT systems |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | touch screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15.7 |  |  | Human Factors
|  | a |  | The telemetry/transmitter unit should have a ‘call nurse’ function, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | an option for the patient to mark an event for recording. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | It is desirable that there is a signal strength indicator on both the transmitter and receiver (i.e. central station), and that central stations incorporate an audible alarm if the unit becomes out of range of the receiver. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 15.8 |  |  | Data Management
|  | a |  | Data transmission must be secure to ensure patient privacy. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Telemetry/transmitter units must be able to transfer data to the central station in real-time. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The transmitter must transmit on a frequency within the spectrum that is allowed by Australian Communications and Media Authority (ACMA). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | If a standard ICT spectrum is used, it must work with other systems without any transmission issues for both telemetry units and other devices using the system. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | When using any spectrum, it must not interfere with any other current, installed systems. A pre-install survey must be completed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | All transmission information must be provided to the Participating Health Services’ nominated department, including frequencies used. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 16 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered module in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • number of selectable ECG leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) (where available) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of monitored parameter numerical fields |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | State the type of spectrum connection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the spectrum frequency range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | "Drop Test" height, as per IEC60601.1 Sec. 15.3 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) (where applicable) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.08-2.10 - Central Station Monitors / Patient Monitoring Information Management Systems / Patient Monitoring Asset Condition Management Software
| 17 |  |  | Standards and Compliance
|  | a |  | Tendered Information/Software Management Systems must only collect, store, and access the minimum amount of patient's personal data, necessary to perform its function. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List the patient's personal data to be collected, stored and accessed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Stored patient's personal data must be encrypted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Where Tendered Information/Software Management Systems are cloud based, then the location of the cloud-based servers must be within the boarders of Australia. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services may prefer that the cloud based servers are located within the State of Victoria. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must maintain and ensure its permitted third-parties, maintain the physical and virtual security of any facilities owned, managed, licensed, or controlled by Respondent that stores patient data by implementing industry best security practices at the locations where patient data is stored to ensure the confidentiality, integrity, and availability of the patient data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondent, where applicable, must maintain backup policies and procedures of Data containing patient's personal data, image repositories, and provisioned environments. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The backup storage infrastructure must be located in physically protected, limited-access facilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Prior to disposal of any Device returned or discontinued by or on behalf of the Participating Health Service, Respondents must securely wipe or destroy all patient data consistent with industry standards, such as NIST 800-88, to ensure that no patient data is retrievable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Upon request of the Participating Health Service, Respondent must provide confirmation of patient's personal data wipe or destruction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 18 |  |  | General Requirements
|  | a |  | With the exception of where a bedside patient monitor is used as a receiver of telemetry information, the base unit configuration for data management systems and central stations is:
|  |  | i | a desktop computer with associated hardware, such as a keyboard, mouse etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a flat panel display (minimum size 19 inches/48 centimetres) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | an annunciator capable of being heard over moderate ambient noise (65dB) and which is adjustable in volume |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | the ability to admit, discharge and transfer patients. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All data management systems/central stations must have the capability for remote connection by successful Respondent's service provider to access. It's the Participating Health Services' discretion whether connection is enabled. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 18.1 |  |  | Event Reporting
|  | a |  | All monitors and data management systems/central stations must have the capability to pass event data wirelessly, via cable or via a memory device (such as a memory card) to other devices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A summary report should be downloadable from monitors and data management systems/central stations in a universal format so that it may become a component of hospital records. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Monitors and data management systems/central stations should have built-in software to manipulate event information to a de facto format used by common PC applications. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Event information must be able to be stored on Participating Health Services’ information systems. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | For research and development purposes, data format should be open ECG-SCP, RCG or similar Open Same Standard format. If only a proprietary format is available, a software development kit for Microsoft Windows should be provided as a purchasable option. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Patient data that is stored on the internal memory should be encrypted to an open standard (i.e. non-proprietary). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | A data copy must be able to be stored in such a manner that it is incorruptible for use in any potential legal analysis. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 18.2 |  |  | Optional Configurations
|  | a |  | It is desirable that data management systems and central stations can include the following configurations upon request by Participating Health Services:
|  |  | i | touch screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | various display sizes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | display of 12-lead ECG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | full disclosure of ECG waveforms that are 24 hours or greater |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | full disclosure of other waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | the ability to reset basic bedside alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | web services (for remote viewing by medical staff) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | integration with other systems |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | strip recorder |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | A4 printer. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 19 |  |  | ICT
|  | a |  | Access to data repository must only be by system authorised users. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 19.1 |  |  | Data Recording
|  | a |  | Web-enabled formats are preferred. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Monitors and data management systems/central stations must automatically record ECG before and after defined events. All other parameters should be recorded as well. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | If the entire ECG, multi-function monitoring, summary report and other defined events is recorded electronically, then it must be uploadable to a computer. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 19.2 |  |  | Data Transmission
|  | a |  | Data transmission must be encrypted to comply with privacy requirements, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | state the cryptographic method employed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Data transmission to other systems must be using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Data transmission via Bluetooth (of any version) is not permitted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Mobile/cellular/wireless data transmission over public networks must be secure. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | System should allow the export of data, i.e. VEMD, from the patient reports. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | All software must conform to the hospital system for data export and interchange. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 19.3 |  |  | Data Hardware and Software
|  | a |  | Software upgrades should occur without having to replace the hardware on each occurrence. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Health Services must be notified of validated ‘Fixes’/’patches’ within one (1) week of software vendor release. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Authorised users must be able to upgrade software, troubleshoot and configure the system settings. System must be capable of running an antivirus software or equivalent. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Printouts must be of an appropriate diagnostic quality. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The ability to print to non-thermal paper is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | All fixed monitors are expected to have standard TCP/IP connections. Cabling installation and management may become a requirement of the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | UPS should be supplied with hardware |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Permissions and Participating Health Services’ protocols must be adhered to. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Vendor supplied networking equipment and servers should be located in the Participating Health Services’ ICT areas, such as communications rooms and server farms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | For Vendor supplied networks, all connecting cables must be provided and installed by the successful Respondent(s) (including fibre optics, where applicable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | For Hospital supplied networks, all patch cables must be provided by the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Patch cables must be green, unless specifically requested to be another colour by a Participating Health Service. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Upon purchase of any data management system or central station, successful Respondents must provide to Participating Health Services:
|  |  | i | technical diagrams of the installation, including a network diagram |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a technical address spreadsheet with all IP, subnet masks and MAC addresses (where appropriate locations should be indicated). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 20 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered module in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Display size |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of displays per system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of patients shown on a single display |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - number of waveforms shown for display with maximum number of patients |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - number of monitored parameter numerical fields |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | List all physiological monitored parameters that can be numerically displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all physiological monitored waveforms that can be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.11 - Transcutaneous Partial Pressure Monitors
| 21 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 18 Transcutaneous Gases (tcpO2, tcpCO2). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | tendered device must be capable of monitoring either or both:
|  |  | i | tcpO2 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | tcpCO2 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | and should include SpO2 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered device should be capable of monitoring the Neonatal population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered device should be capable of monitoring Adult and Paediatric populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 22 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.12 - Cardiac Output Monitors
| 23 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 10 Cardiac Output |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device should measure other haemodynamic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Should be capable of monitoring Paediatric and Neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 24 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List compatible catheter brands |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.13 - Regional Oximetry Monitors
| 25 |  |  | General Requirements
|  | a |  | Tendered device must use Near Infra Red Spectroscopy to non-invasively continuously monitor cerebral oxygenation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Should be capable of monitoring Paediatric and Neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 26 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.14 - Respiratory Gas Monitors
| 27 |  |  | General Requirements
|  | a |  | Tendered device must comply with all requirements stated in Part E section 11 Anaesthetic Gas / Integrated Patient Gas Monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered device must continuously sample and measure inspired and expired concentrations of respiratory and/or anaesthetic gases during and immediately following anaesthetic administration. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered device must display inspired and expired gas concentrations of CO2 and halogenated agent, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | inspired (or mean) concentrations of O2 and N2O, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | respiration rate. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered device must display a CO2 waveform when running. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 27.1 |  |  | Optional Configurations
|  | a |  | It is desirable that Respiratory Gas monitors can include Spirometry |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Spirometry option must comply with all requirements as stated in Part E section 12 Spirometry. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 28 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored gas parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored gas parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | State the gas sampling flow rate, mL/min |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | State the rise time for each gas, msec |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | State the water trap volume, mL (where available) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 2.15 - Depth-of-Anaesthesia Monitors
| 29 |  |  | General Requirements
|  | a |  | Tendered device must be an EEG system for the purpose of determining patient sedation during anaesthesia. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Should be capable of monitoring Paediatric and Neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered device must:
|  |  | i | have artifact rejection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • can artifact rejection be bypassed (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | include an impedance checking ability that can be bypassed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tendered device should compute and display the following parameters:
|  |  | i | Bispectral Index |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Suppression Ratio |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | EMG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Signal Quality Indicator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Burst Count |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Please state if a bilateral in addition to a unilateral capability is available.
|  |  | vii | Please state Epoch duration
|  |  | viii | Please list any other parameters
| 30 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | List all monitored physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological monitored parameters that are displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | state the amplifier input range (mV) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | state sweep speed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 3-ECGs
Category 3: ECG Machines
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 3.0 |  |  |  | ECG Machines | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | ECG Machines
|  |  | i | A device that is used to detect, register, and record the electrical signals produced through the activity of the heart and to reproduce these signals, (voltage versus time) as an electrocardiogram (ECG), typically in the form of a print-out, but also on a display or in digital media for further use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | ECG Machines, Handheld
|  |  | i | A small portable ECG machine that that produces diagnostic quality electrocardiograms and is able to be carried by hand; |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | includes PC-based ECG recorders that utilise a module to capture all 12 ecg leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | ECG Machines, Cardiac Stress Test
|  |  | i | An ECG machine that records electrocardiograms while the patient's heart is being exercised. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Cardiac Stress Testing Systems
|  |  | i | System must include an ECG machine and a device that can exercise patient's heart, e.g. treadmill, exercise bike, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | ECG Management Systems
|  |  | i | Database for electrocardiogram acquired from ECG machines, Holter acquisition devices, stress systems, bedside monitors, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 3.01-3.03 - ECG Machines / Handheld ECG Machines / PC Based ECG Machines
| 2 |  |  | Base Units
|  | a |  | Must be capable of monitoring Adult population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Should be capable of monitoring Paediatric and Neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Heart Rate range must be 30–300 beats per minute. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Heart Rate accuracy must be ± 10% or ± five (5) beats per minute, whichever greater. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | ECG leads must be AHA colour coded |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Leads off advisory for disconnected leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Electrocardiogram must offer the following ECG filtering:
|  |  | i | selectable mains noise, muscle artifact, and baseline wander filtering |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | wide filter selection between 0.05 to 150Hz. State configurable filter settings for: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• low cut-off frequencies
• high cut-off frequencies
|  | h |  | ECG waveform quality indicator should be provided |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.1 |  |  | Power Supply
|  | a |  | ECG machine must be capable of acquiring and printing at least 25 resting ECGs on a single charge |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.2 |  |  | Printer
|  | a |  | High resolution print-out of ECG, must be at least 500dpi (horizontal) at 25mm/sec by 200dpi (vertical). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State printer resolution, (dpi)
|  | b |  | ECG machine must be able to print at 25mm/s. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State any other printing speeds
| 3 |  |  | Data Management
| 3.1 |  |  | Data Recording and Management
|  | a |  | ECG machine must be able to store ECGs on memory |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Each individual ECG must be time/date stamped |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | ECG machine should be able to store ECG in following formats:
|  |  | i | PDF |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | XML |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference may be given to ECG machines that can store ECG in DICOM format |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Patient data that is stored on the internal memory should be encrypted to an open standard (i.e. non-proprietary). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.2 |  |  | Data Transmission
|  | a |  | ECG Machines must be able to transfer data via one of the following connection types: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state Wi-Fi technology used
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  | b |  | Data transmission must be encrypted for privacy requirements, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | state the cryptographic method.
|  | c |  | ECG machine should be capable of bidirectional communication. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Time should be automatically synchronised to a Network Time Server |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Preference may be given to ECG machines:
|  |  | i | capable of retrieving patient demographic information |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | can download orders worklists |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | have proven records of integration with eMR systems. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Data transmission to other systems is preferred, using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free upon purchase. Details of all conditions and subsequent annual costs must also be provided pre-purchase. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | General Requirements - ECG Machines
|  | a |  | The minimum base unit configuration for ECG machine must include the following:
|  |  | i | at least or greater than 10-ECG lead capability |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | simultaneously acquire all 12 leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | display patient's heart rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | print out all 12 leads on a single A4 sheet |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | algorithm for rhythm analysis (automatic measurements) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | interpretation capabilities |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | simultaneously display of up to 12 leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | full alphanumeric keyboard |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Mobile cart with lockable castors and storage space |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | In manual mode, print all 12 leads continuously at once. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Full screen preview of complete 12 lead report should be shown prior to printing |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | General Requirements - ECG Machines, Handheld
|  | a |  | The minimum base unit configuration for ECG machine should include the following:
|  |  | i | simultaneously acquire all 12 leads |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | display patient's heart rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | print out all 12 leads on a strip recorder |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | algorithm for rhythm analysis (automatic measurements) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | interpretation capabilities |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | carry case |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Can include pc based ECG machines |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | how many leads attach to patient |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | state start-up time (when ready to record ECG) from a cold start (seconds) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List the format of all reports |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • resolution, number of pixels horizontally by vertically |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List all features |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • touch screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • barcode reader |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • paediatric interpretive analysis packages |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • ST segment analysis package |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Stress testing package |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Pharma package |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • other |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 3.04-3.05 - Cardiac Stress Test ECG Machines / Cardiac Stress Testing Systems
| 7 |  |  | General Requirements - ECG Machines, Cardiac Stress Test
|  | a |  | Tendered ECG machine must
|  |  | i | comply with all Category 4.01 ECG Machine requirements |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | automatically record and analyse ST-wave activity throughout the test period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | automatically detect and analyse the arrhythmia throughout the test period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • be able to label/document detected a arrhythmias |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | graph the interval between S and T in the ST wave |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • also on the medians |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | automatically detect I, J and J+x all points on an ST segment Analysis |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | provide Pacemaker Analysis |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | General Requirements - Cardiac Stress Testing Systems
|  | a |  | ECG machine included in Tendered system must comply to Category 4.03 Cardiac Stress Test ECG Machines |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered system must included an exercise device, e.g. treadmill or bicycle ergometer connected with an ECG machine |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | ECG Machine must control the exercise device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Stand-alone exercise devices must not be tendered |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | What parameters are measured and displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | what recoding speeds are available (mm/sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Report Formats
|  |  |  | • pre-programmed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • user customisable |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Patient interface
|  |  |  | • maximum number of lead wires |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Is there wireless connectivity to the patient (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - if wireless, type of battery used & how long between charges/changes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • How many patient interfaces can be simultaneously connected |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List the parameters available for analysis |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Exercise Protocols:
|  |  |  | • list the number of pre-programmed Exercise Protocols offered |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • what standard protocols are available, e.g. Bruce, Naughton, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • what is the maximum number of User programmable Exercise Protocols possible |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • list the user programmable parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Peripherals:
|  |  |  | • list name and type# of exercise device supported |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Treadmills: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - speed range (km/hr) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - elevation range (%) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - maximum patient weight capacity (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - size (L x W) cm |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• Ergometers:
|  |  |  | - Maximum Patient weight capacity (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - Load Range (Watts) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - Handlebar height Range (cm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - Handlebar Rotation Range (degrees) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - Must have Adjustable Crank Length, please describe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - Must have Adjustable Seat Height, please describe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - Footprint (L x W) cm |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | Full Disclosure:
|  |  |  | • what is the length of time available for Full Disclosure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • what parameters are stored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • what is the number of waveforms that can be stored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | In Trending, what parameters are captured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | IPX rating (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 3.06 - ECG Management Systems
| 10 |  |  | Standards and Compliance
|  | a |  | Tendered System must only collect, store, and access the minimum amount of patient's personal data, necessary to perform its function. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List the patient's personal data to be collected, stored and accessed
|  |  | ii | Stored patient's personal data must be encrypted
|  | b |  | If Tendered Systems are cloud based, then the location of the cloud-based servers must be within the boarders of Australia. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services may prefer that the cloud based servers are located within the State of Victoria. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must maintain and ensure its permitted third-parties, maintain the physical and virtual security of any facilities owned, managed, licensed, or controlled by Respondent that stores patient data by implementing industry best security practices at the locations where patient data is stored to ensure the confidentiality, integrity, and availability of the patient data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondent, where applicable, must maintain backup policies and procedures of Data containing patient's personal data, image repositories, and provisioned environments. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The backup storage infrastructure must be located in physically protected, limited-access facilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Prior to disposal of any Device returned or discontinued by or on behalf of the Participating Health Service, Respondents must securely wipe or destroy all patient data consistent with industry standards, such as NIST 800-88, to ensure that no patient data is retrievable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Upon request of the Participating Health Service, Respondent must provide confirmation of patient's personal data wipe or destruction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11 |  |  | General Requirements
|  | a |  | Tendered System must provide the capability to acquire, process, store /retrieve, and distribute ECG and related data results in a fast, reliable, and accurate manner |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | System must be capable of receiving ECG's from multiple sources and formats (support Multi-vendor and Multi-modality) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | System must be capable of receiving and analysing ECG acquired at the patient bedside by patient monitors |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | System must be capable of receiving and analysing ECG acquired by Ambulatory Monitoring Systems (i.e. Holter Systems) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | System must be capable of re-analysing any ECG sent from 3rd party ECG sources that meet the standard format requirements (minimum 12 leads for 10 seconds, 500 samples per second). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | System must be capable of receiving and re-analysing ECGs from stress testing systems |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | ECG Analysis algorithm must be capable of analysing 12, 16 and 18 lead ECGs acquired at patient side |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Bidirectional communication with ECG machine, where patient demographics must be made available with or without a work order. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The system must be able to accept data in one of the following formats:
|  |  | i | Raw ECG data |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | .pdf |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | .xml |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | HL7 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | DICOM |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | The Solution must be capable of performing interpretation and a comparison of current and previous data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The interpretative comparisons must be performed on the:
|  |  |  | • waveform; |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • statements; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • measurements. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | It must be possible to configure the workflow to direct a specific test or patient type to a specific Cardiologist or group of Cardiologists |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Allow confirmation of ECG by authorised users; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Allow generation of reports. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | The Solution must provide a range of Physician Report review tools. Tenderer to list all available tools. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | The Solution must allow for role based customisation of the GUI for viewing and interpretation of data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | The Solution must provide an electronic signature for any final report generated by the system. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11.1 |  |  | Optional Configurations
|  | a |  | Clinicians should be able to access the patient ECG from anywhere and on any mobile device without the use of additional applications that depended on device operating system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | Data Management
|  | a |  | Access to data repository must only be by system authorised users. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12.1 |  |  | Data Recording
|  | a |  | Web-enabled formats are preferred. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12.2 |  |  | Data Transmission and Interfacing
|  | a |  | Data transmission must ensure patient privacy. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Clinical systems are being integrated to other patient information systems including HIS, EMR etc. Systems must be using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | System must allow bidirectional data flow |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12.3 |  |  | Data Hardware and Software
|  | a |  | Authorised users must be able to upgrade software, troubleshoot and configure the system settings. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Printouts must be of an appropriate diagnostic quality. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The ability to print to non-thermal paper is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All connecting devices must have standard TCP/IP connections. Cabling installation and management may become a requirement of the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Permissions and Participating Health Services’ protocols must be adhered to. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Vendor supplied networking equipment and servers should be located in the Participating Health Services’ ICT areas, such as communications rooms and server farms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | For Vendor supplied networks, all connecting cables must be provided and installed by the successful Respondent(s) (including fibre optics, where applicable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | For Hospital supplied networks, all patch cables must be provided by the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Patch cables must be green, unless specifically requested to be another colour by a Participating Health Service. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Upon purchase of any ECG Management system, successful Respondents must provide to Participating Health Services:
|  |  | i | technical diagrams of the installation, including a network diagram |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a technical address spreadsheet with all IP, subnet masks and MAC addresses (where appropriate locations should be indicated). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | System must be scalable to increasing workload by means of adding additional servers which operate in parallel whilst providing business continuity without having to replace the ECG server. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | list any known incompatible ECG machines |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 4-Ambulatory
Category 4: Ambulatory Monitoring
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 4.0 |  |  |  | Ambulatory Monitoring | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Holter Monitoring Systems
|  |  | i | Database for electrocardiogram acquired from Holter monitoring; downloaded and stored results from Holter monitors |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Holter Monitor, Continuous
|  |  | i | A Holter monitor is a small, wearable device that records the heart’s rhythm Continuously. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Holter Monitor, Event
|  |  | i | A Holter monitor is a small, wearable device that records the possible Heart Rhythm disturbances. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Ambulatory Blood Pressure Monitoring Systems
|  |  | i | Database for blood pressure readings acquired from Ambulatory Blood Pressure Monitoring; downloaded and stored results from ambulatory blood pressure monitors |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Ambulatory Blood Pressure Monitors
|  |  | i | The device that records the blood pressure readings over a defined period of time |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 4.01&4.04 - Holter Monitoring Systems / Ambulatory Blood Pressure Monitoring Systems
| 2 |  |  | Standards and Compliance
|  | a |  | Tendered System must only collect, store, and access the minimum amount of patient's personal data, necessary to perform its function. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List the patient's personal data to be collected, stored and accessed
|  |  | ii | Stored patient's personal data must be encrypted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If Tendered Systems are cloud based, then the location of the cloud-based servers must be within the boarders of Australia. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services may prefer that the cloud based servers are located within the State of Victoria. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must maintain and ensure its permitted third-parties, maintain the physical and virtual security of any facilities owned, managed, licensed, or controlled by Respondent that stores patient data by implementing industry best security practices at the locations where patient data is stored to ensure the confidentiality, integrity, and availability of the patient data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Respondent, where applicable, must maintain backup policies and procedures of Data containing patient's personal data, image repositories, and provisioned environments. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | The backup storage infrastructure must be located in physically protected, limited-access facilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Prior to disposal of any Device returned or discontinued by or on behalf of the Participating Health Service, Respondents must securely wipe or destroy all patient data consistent with industry standards, such as NIST 800-88, to ensure that no patient data is retrievable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Upon request of the Participating Health Service, Respondent must provide confirmation of patient's personal data wipe or destruction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | General Requirements
|  | a |  | Must be a PC based Workstation offering digital recording, review, analysis, reporting and archival of ambulatory monitored studies |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The Workstation must as a minimum include the following capabilities:
|  |  | i | The ability to network to a site provided Laser printer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Digital recorder configuration and study download |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Full Disclosure analysis, interpretation and reporting |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | A patient study and report database |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must include Full Disclosure analysis, interpretation and reporting for a minimum of 12 hours. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Respondent to state pricing for options reporting 24 hours and 48 hours in the TRW |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Workstation database must offer the minimum capacity to store the complete study (Full Disclosure) and report for 500 patients |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tendered system must be suitable for use with Adult population |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered system should be suitable for use with Paediatric population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | General Requirements - Holter Monitoring Systems
|  | a |  | The Workstation must as a minimum include the following capabilities:
|  |  | i | ST analysis, interpretation and reporting |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | QT analysis, interpretation and reporting |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Heart rate variability / RR interval analysis, interpretation and reporting |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Pacemaker analysis, interpretation and reporting |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must include user selectable 2 leads of ECG Full Disclosure analysis, interpretation and reporting for a minimum of 12 hours. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Respondent to state pricing for options reporting 24 hours and 48 hours in the TRW |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | ICT
|  | a |  | Does the tendered product need to have connectivity onto the hospital's IT network infrastructure? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.1 |  |  | Data Transmission and Interfacing
|  | a |  | Data transfer from the ambulatory recorder to the management system must be via industry standard connections. For example: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state wi-fi version
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | iii | USB flash drive |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | memory cards |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state the type of card
|  |  | v | Other, please specify
|  | b |  | Data transmission via Bluetooth (of any version) is not permitted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Data transmission must be encrypted to comply with privacy requirements, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | state the cryptographic method employed.
|  | d |  | Data transmission to other systems must be using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | ECG rhythm printouts must be of an appropriate diagnostic quality. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Printouts should be on non-thermal paper. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 4.02-4.03 - Holter Monitor, Real-time / Holter Monitor, Retrospective
| 7 |  |  | General Requirements
|  | a |  | Tendered system must be suitable for use with Adult population |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered system should be suitable for use with Paediatric population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Full Disclosure digital recording for at least:
|  |  | i | 24 hours of monitoring |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | 2 ECG channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Recorded annotation must include:
|  |  | i | patient event marker |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | time |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Arrhymia detection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Pacemaker detection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.1 |  |  | Power Supply
|  | a |  | Must be a battery powered device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | monitor should use commonly available battery type (alkaline) that does not require frequent replacement |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | runtime on battery must be greater than 24 hours |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | depleted rechargeable battery (where available) must fully recharge in under 16 hours |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | must have a low battery alert |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Patient type |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Number of patient lead wires |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | full 12-lead |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | derived 12-lead |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | type of storage and capacity, (MB) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | size of capacity (hr) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | frequency response |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | sampling rate (samples/sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Resolution |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 4.05 - Ambulatory Blood Pressure Monitors
| 9 |  |  | General Requirements
|  | a |  | Tendered system must be suitable for use with Adult population |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered system should be suitable for use with Paediatric population. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered device measures blood pressure noninvasively for at least a 24 hour period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tender device must have Full Disclosure of all measurements |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Recorded annotation must include:
|  |  | i | patient event marker |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | time |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | alarm type |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.1 |  |  | Power Supply
|  | a |  | Must be a battery powered device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | monitor should use commonly available battery type (alkaline) that does not require frequent replacement |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | runtime on battery must be greater than 24 hours |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | depleted rechargeable battery (where available) must fully recharge in under 16 hours |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Recorder must have a low battery indicator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 10 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Patient type |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | systolic blood pressure range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Accuracy |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Overpressure cut-off (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Pulse Rate range (bpm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | measurement frequency intervals |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Measurement duration (hr) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Cuff inflation rate (mmHg/sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Cuff deflation rate (mmHg/sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | list all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 5-Intrapartum
Category 5: Intrapartum Monitoring
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 5.0 |  |  |  | Intrapartum Monitoring | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | FHR Dopplers
|  |  | i | Foetal Heart Rate (FHR) Dopplers, also known as foetal heartbeat monitors, are devices used to detect and monitor the heartbeat of a foetus |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | CTG Machines
|  |  | i | Known as a cardiotocograph, is used to monitor the foetal heartbeat and uterine contractions during pregnancy and labour |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Wireless Transducer Patch Systems
|  |  | i | Device to record intrapartum maternal/foetal monitor that noninvasively measures and displays foetal heart rate (FHR), maternal heart rate (MHR), and uterine activity (UA) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Obstetric Management Systems
|  |  | i | Database for cardiotocography acquired from CTG machines during pregnancy and labour, allowing annotations on CTG trace. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 5.01 - FHR Dopplers
| 2 |  |  | General Requirements
|  | a |  | Tendered device must :
|  |  | i | include a speaker with user selectable volume of the FHR |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | monitor FHR rates of at least 50 to ≥200 bpm |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | include power on/off controls |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | include an audio output socket for headphone use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.1 |  |  | Power Supply
|  | a |  | Battery-powered devices must operate or at least twelve (12) hours, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | include a low battery indicator. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (g) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | probe frequency/ies (MHz) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Ultrasound intensity (mw/cm2) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | interchangeable probe (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Display size, H x W (mm) (where available) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • List displayed variables shown on display |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | FHR range (bpm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | IPX rating of probe/s (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | IPX rating of unit (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | List all accessories included in the base configuration, e.g. case/bag |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 5.02 - CTG Machines
| 4 |  |  | General Requirements
|  | a |  | The minimum base unit configuration for CTG machine must include the following:
|  |  | i | ultrasound transducer to measure foetal heart rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | toco transducer to monitor uterine activity |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | integrated recorder (chart speed 1cm/min) recording all monitored patient parameters for both foetus and mother |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | identify and distinguish between both the maternal and foetal heart rates |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | patient event switch |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | CTG machine should be able to connect to an Obstetric Management System (OMS) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List all compatible OMS
|  | c |  | If the CTG machine is offered with a cart, the cart must be easily transportable and meet all OH&S manual handling requirements. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | If a cart is offered, a braking feature must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Portable CTG machines must not exceed a total weight (including any auxiliary equipment) of seven (7) kilograms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | A cart must be included for CTG machines that exceed a total weight (including any auxiliary equipment) of seven (7) kilograms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | When US transducer is connected, numeric FHR must be displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | When toco transducer is connected, numeric uterine activity must be displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Each of the physiologic parameters monitored by the tendered device, must comply with the relevant requirements stated in Part E: Specifications for Optional Physiological Parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5 |  |  | Data Management
| 5.1 |  |  | Data Recording and Management
|  | a |  | CTG machines must incorporate date and time recording of patient observations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Additional configurable parameters are desirable, such as temperature, weight, pain etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | CTG machines must record and store all events for the entire patient episode, along with patient details. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | CTG machines must have an incorporated paper recorder that continuously records FHR and UA waveforms at 1 cm/min |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | CTG machines must have an incorporated paper recorder that captures the sequence of events during use on a patient (including the times of events, waveforms etc.). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Patient data that is stored on the internal memory should be encrypted to an open standard (i.e. non-proprietary). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.2 |  |  | Data Transmission
|  | a |  | CTG machines must be able to transfer data via one of the following connection types: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state Wi-Fi technology used
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | iii | serial (e.g. RS232) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Data transmission must be encrypted for privacy requirements, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | state the cryptographic method.
|  | c |  | All data output ports must be activated and ready for use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preference may be given to CTG machines capable of bidirectional communication. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Data transmission to other systems is preferred, using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free upon purchase. Details of all conditions and subsequent annual costs must also be provided pre-purchase. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | suitable probe frequency/ies (MHz) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of waveforms able to be displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • parameters able to have waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum number of numerical fields monitoring physiologic parameters |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | List all physiological parameters that are monitored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state the accuracy of each monitored physiologic parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | list all parameters able to be transmitted wirelessly (e.g. telemetry) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | IPX rating of wired transducers (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | IPX rating of wireless transducers (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | IPX rating of unit (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 5.03 - Wireless Transducer Patch Systems
| 7 |  |  | General Requirements
|  | a |  | Requirements for each physiological measurement must be met, as outline in 'Part E: Specifications for Physiological Parameters'. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered wireless transducers must:
|  |  | i | work with a tendered CTG machine/system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | must be single patient use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | attach to the patient without the use of a belt or strap |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Product Information
|  | a |  | Must include a product brochure in ITS response for each tendered device | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each wearable device, Respondent must submit as part of their tender response, a complete set of operator user manuals in portable document format (i.e., .pdf). | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 5.04 - Obstetric Management Systems
| 9 |  |  | Standards and Compliance
|  | a |  | Tendered System must only collect, store, and access the minimum amount of patient's personal data, necessary to perform its function. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List the patient's personal data to be collected, stored and accessed
|  |  | ii | Stored patient's personal data must be encrypted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If Tendered Systems are cloud based, then the location of the cloud-based servers must be within the boarders of Australia. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services may prefer that the cloud based servers are located within the State of Victoria. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must maintain and ensure its permitted third-parties, maintain the physical and virtual security of any facilities owned, managed, licensed, or controlled by Respondent that stores patient data by implementing industry best security practices at the locations where patient data is stored to ensure the confidentiality, integrity, and availability of the patient data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondent, where applicable, must maintain backup policies and procedures of Data containing patient's personal data, image repositories, and provisioned environments. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The backup storage infrastructure must be located in physically protected, limited-access facilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Prior to disposal of any Device returned or discontinued by or on behalf of the Participating Health Service, Respondents must securely wipe or destroy all patient data consistent with industry standards, such as NIST 800-88, to ensure that no patient data is retrievable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Upon request of the Participating Health Service, Respondent must provide confirmation of patient's personal data wipe or destruction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 10 |  |  | General Requirements
|  | a |  | Tendered System must provide the capability to acquire, process, store /retrieve, and distribute CTG and related data results in a fast, reliable, and accurate manner |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Each and every CTG machine must connect to the system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All data from each and every CTG machine must be captured by the system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | System must interface with vendor neutral / all brand of CTG machines |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Scalable / Modular System; Respondents must allow system to be scalable at deployment to fit into Health Services allocated budget. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Customisable Application; Respondents must provide the required resources for Health Services to customise the app, configure for site unique dashboards, charting data, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | System must have the ability to admit, discharge and transfer patients |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Flexible client access; e.g. virtual, remote desktop, mobile |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | All systems must have the capability for remote connection by successful Respondent's service provider to access. It's the Participating Health Services' discretion whether connection is enabled. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free upon purchase. Details of all conditions and subsequent annual costs must also be provided upon purchase. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | A summary report should be downloadable from each patient episode in a universal format, so that it may become a component of hospital records. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Patient episode information must be stored on Participating Health Services’ information systems. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Patient data that is stored on the internal memory should be encrypted to an open standard (i.e. non-proprietary). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | Data must be stored in such a manner that it is incorruptible and can be retrieved at a later date for use in any potential legal analysis. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11 |  |  | ICT
|  | a |  | Access to data repository must only be by system authorised users. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11.1 |  |  | Data Recording
|  | a |  | Web-enabled formats are preferred. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The entire FHR and maternal UA must be recorded electronically and must be uploadable to a computer. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11.2 |  |  | Data Transmission and Interfacing
|  | a |  | Data transfer from the CTG Machine to the management system must be via industry standard connections. For example: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state wi-fi version
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | iii | Other, please specify
|  | b |  | Data transmission via Bluetooth (of any version) is not permitted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Data transmission must be encrypted to comply with privacy requirements, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | state the cryptographic method employed.
|  | d |  | Data transmission to other systems must be using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Participating Health Services may preference CTG Machines that communicate wirelessly with the Obstetric Management System. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11.3 |  |  | Data Hardware and Software
|  | a |  | Authorised users must be able to upgrade software, troubleshoot and configure the system settings. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Printouts must be of an appropriate diagnostic quality. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The ability to print to non-thermal paper is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All CTG machines should have standard TCP/IP connections. Cabling installation and management may become a requirement of the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Permissions and Participating Health Services’ protocols must be adhered to. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Vendor supplied networking equipment and servers should be located in the Participating Health Services’ ICT areas, such as communications rooms and server farms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | For Vendor supplied networks, all connecting cables must be provided and installed by the successful Respondent(s) (including fibre optics, where applicable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | For Hospital supplied networks, all patch cables must be provided by the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Patch cables must be green, unless specifically requested to be another colour by a Participating Health Service. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Upon purchase of any obstetric management system, successful Respondents must provide to Participating Health Services:
|  |  | i | technical diagrams of the installation, including a network diagram |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a technical address spreadsheet with all IP, subnet masks and MAC addresses (where appropriate locations should be indicated). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | list any known incompatible CTG machines |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 6-Sleep&Neuro
Category 6: Neurological, Sleep, and Sleep Related Monitoring
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 6.0 |  |  |  | Neurological, Sleep, and Sleep Related Monitoring | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Cerebral Function Monitors [Amplitude-integrated Electroencephalography (aEEG)]
|  |  | i | An amplitude-integrated EEG (aEEG) monitor to continuously monitor the electrocortical activity of patients in an ICU setting, more so in NICUs, for monitoring at risk preterm and term infant. Has fewer channels and monitored for longer periods than traditional EEG monitoring. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Electroencephalography (EEG) Units
|  |  | i | A device that records the electroencephalogram (EEG), the electrical activity in the brain, by having small sensors (electrodes) attached to the scalp to detect the electrical signals produced by the brain. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Can include video telemetry, video EEG. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Polysomnography (PSG) Units
|  |  | i | Sleep Studies. The monitoring of EEG and other vital signs, generally overnight, to assess sleep. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Can include video telemetry. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Electromyograph (EMG) / Evoked Potential (EP) Systems - Nerve Conduction Study (NCS)
|  |  | i | Electromyography (EMG) units measure the muscle response, or electrical activity, in response to a nerve’s stimulation of the muscle. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Evoked potential units measure the brain bioelectric response to either a visual, auditory, and/or somatosensory external stimuli. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Nerve conduction study devices measure how fast and how well the nerve responds to stimuli, by first applying the stimuli and then measuring the response. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Clinical Neurology Management Systems
|  |  | i | Database for electroencephalogram acquired from one or more of the above type listed machines |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Data Management
| 2.1 |  |  | Data Recording and Management
|  | a |  | Devices with data recording incorporated must time and date stamp all recorded patient observations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Additional configurable patient parameters, such as age, weight, etc. are desirable as manual entry |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Devices should record and store multiple events |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Devices must be able to printout any captured sequence of events during use on a patient (including the times of events, waveforms etc.). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | It is desirable that device be capable of printing to an A4 printer directly. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Patient's personal information that is stored on internal memory should be encrypted to an open standard (i.e. non-proprietary). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.2 |  |  | Data Transmission
|  | a |  | If a data storage device is present, data must be able to be downloaded to a Data Management System when required. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Device must be able to transfer data via one of the following connection types: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state Wi-Fi technology used
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | iii | RS232 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | IR connection. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Bluetooth |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • the operator should be able to disabled/enabled the Bluetooth connection in settings on the device |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | USB flash drive |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | memory cards |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state the type of card
|  |  | viii | Other, please specify
|  | c |  | Data transmission must ensure that the privacy of the patient's personal information is not breached. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All data output ports must be activated and ready for use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Preference may be given to devices capable of bidirectional communication. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 6.01 - Cerebral Function Monitors
| 3 |  |  | General Requirements
|  | a |  | Electroencephalograph (EEG) acquisition system intended to be used in a hospital environment to record, collect, display and facilitate manual marking of aEEG recordings. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The aEEG must be obtained from at least a pair of parietal electrodes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Able to analyse and identify background patterns in aEEG, including |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | continuous and discontinuous activity, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | burst suppression, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | low voltage, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | inactive patterns. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Include detection algorithm that identifies sections of EEG/aEEG that may correspond to electrographic seizures |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display modes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | List all parameters that are monitored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Number of channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Resolution (bits) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Sampling rate (Hz) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Frequency range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | CMRR (dB) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | Trending memory length |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | IPX rating of unit (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 6.02 - Electroencephalograph (EEG) Units
| 5 |  |  | General Requirements
|  | a |  | Tendered device must at least be a 32 channel digital EEG system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | A dedicated review workstation is required. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.1 |  |  | EEG Amplifiers
|  | a |  | Must have configurable Low frequency filters setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must have configurable High frequency filters setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must have a CMRR of better than at least 110 dB @ 50 Hz |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Preamplifier input impedance at least more than 10 MΩ |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Sensitivity control range must be at least from 1 to 300 μv/mm |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Must have calibration signals |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.2 |  |  | Analysis Software
|  | a |  | Must include automatic spike and seizure detection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The automatic detection sensitivity must be user configurable |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Display of any group of at least 16 of all channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Event markers and annotation to be inserted in signal window and be added to the waveforms in the form of free text |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Community used event markers must be able to be selected through small dropdown menu or something similar without interrupting the acquisition or the display. Please describe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | It must be possible to manually enter Patient demographics |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The electrodes used must be displayed next to each respective channel |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Automatic alarms must occur when EEG values exceed or fall below the established limits |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The system must be able to produce:
|  |  | i | Free text report in soft and hard copy |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Polygraphy; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Real time EEG as displayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Event Searching Utility must be provided to allow the user to find, view, or delete specific categories of scored events please describe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | If Brain Mapping Frequency Analysis is offered, then the unit must display all channels with full brain mapping functions providing colour coded graphical displays |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.3 |  |  | EEG Digital Videometry
|  | a |  | A camera with the following features accessories is required: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Minimum 1080P video resolution. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Day/night switch |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Autofocus |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Auto brightness |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Remote Pan/Tilt/Zoom controlled remotely |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | A microphone; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | A camera mount |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The review software must have at least the following video editing functions:
|  |  | i | Copy, cut and delete; |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Marking; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Crop |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The system must be able to synchronise the EEG waveforms and video images: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | throughout both normal playing and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | fast forward/reverse recording |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The system must allow the operator to select a review in any of these formats: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Video image and waveforms in one screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Video images only; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Waveforms only |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Video Editing must allow
|  |  | i | Removal of entire video recording, retaining EEG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Independent removal of combined segments of Video & EEG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Trim sections of Video whilst keeping all or selected sections of EEG; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Keep all Video and EEG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Digital video must be recorded using MPEG4 compression. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Please describe an alternate method of compression.
|  | g |  | It must be possible to resize the Video Image, i.e. change the size and position of the video image during recording and review |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.4 |  |  | EEG Review Workstation
|  | a |  | The review workstation must allow the reporting physician to:
|  |  | i | Select portions of the waveforms and save it to a separate file |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | The file must be able to be archived. Please state method. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Append/edit a text report |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Annotate and reannotate the waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Include time and date automatically with the ability to be corrected by the operator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Review and edit the video images; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Produce a printout of the above |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.5 |  |  | Ambulatory EEG Monitoring
|  | a |  | Must have a 24 referential channel recorder system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The Ambulatory EEG Monitoring offer must have the following features:
|  |  | i | CMMR of at least 110 dB |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Patient Event Marker |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Review at variable time base |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | The unit must use commonly available batteries |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | The unit must display battery level |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | An ambulatory recorder that will record a minimum of 24 hours of continuous recording and up to 96 hours of data |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• What length of recording is available
|  |  | vii | Sampling rate up to 256 Hz |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Import studies into central patient database |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Must be compatible with the EEG software installed on the main system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Removable Storage Media minimum 5GB |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Must have a purpose built lockable case |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Ambulatory Recorder must not be able to be accidentally switched off |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display modes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | List all parameters that are monitored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Number of channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Resolution (bits) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Sampling rate (Hz) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Frequency range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | CMRR (dB) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | Trending memory length |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | Dedicated workstation offline and sharing database or online |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | List EEG amplifier Low frequency setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | List EEG amplifier High frequency setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xix | EEG amplifier sensitivity range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xx | List display modes offered in the Analysis Software package |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxi | List all Patient personal information/demographics fields in Analysis Software package |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxii | Brain Mapping Frequency Analysis offered |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxiii | State the video resolution of camera |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxiv | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxv | IPX rating of unit (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxvi | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 6.03 - Polysomnography (PSG) Units
| 7 |  |  | General Requirements
|  | a |  | At the time of offer, PSG units must comply with the latest (most current) version of the Polysomnography (PSG) Guidelines as published by the Australasian Sleep Association (ASA). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The system should be able to monitor multiple beds simultaneously |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered system should be at least 32 referential channels / 6 differential channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The amplifier must have a Common Rejection Ratio of at least 100dB |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The sensitivity control must cover the range of 0.5 to 2,000 µv/cm over 8 steps, or better |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Must have configurable Low frequency filters setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Must have configurable High frequency filters setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Both common and differential mode impedance's must exceed 50 MΩ. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | A notch filter centred at 50 Hz that can be enabled/disabled when reviewing waveforms is required |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Is the tendered product a Type (Level) 1 respiratory diagnostic sleep study system? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• If yes to the above question, then compete all questions in clause (k):
|  | k |  | As a minimum, Type 1 respiratory diagnostic sleep study systems must monitor and measure: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | two(2) or more Electroencephalogram (EEG) channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | bilateral electro-oculograms (EOG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | submental electromyography (EMG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | electrocardiography (ECG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | bilateral anterior tibial muscle activity |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | pulse oximetry (SpO2) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | sound/audio recording |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | respiratory thoraco-abdominal movements |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Respiratory Inductance Plethysmography (RIP), or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Piezoelectric (PE) belt |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | nasal pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | oronasal thermal airflow |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | body position |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Tendered Type 1 respiratory diagnostic sleep study systems must be able to receive the following signals from CPAP machines: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | PAP set pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | PAP flow |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | PAP leak |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Is the tendered product a Type (Level) 2 respiratory diagnostic sleep study system? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• If yes to the above question, then compete all questions in clause (n):
|  | n |  | As a minimum, Type 2 respiratory diagnostic sleep study systems must monitor and measure: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | two(2) or more Electroencephalogram (EEG) channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | bilateral electro-oculograms (EOG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | chin electromyography (EMG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | electrocardiography (ECG) or heart rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | oronasal thermal airflow |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | respiratory thoraco-abdominal movements |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | pulse oximetry (SpO2) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | Is the tendered product a Type (Level) 3 respiratory diagnostic sleep study system? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• If yes to the above question, then compete all questions in clause (p):
|  | p |  | As a minimum, Type 3 respiratory diagnostic sleep study systems must monitor and measure at least four (4) of the following variables: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | pulse oximetry (SpO2) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | respiratory effort; thoracic and/or abdominal |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | airflow; (oral or nasal) pressure or oronasal thermistor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | position; head or body |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | electrocardiography (ECG) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | tonometry |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | actigraphy |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | sound/audio recording |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | Is the tendered product a Type (Level) 4 respiratory diagnostic sleep study system? (yes/no) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• If yes to the above question, then compete all questions in clause (r):
|  | r |  | As a minimum, Type 4 respiratory diagnostic sleep study systems must monitor and measure at least one (1) or more of the following variables: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | pulse oximetry (SpO2) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | heart rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | airflow; (oral or nasal) pressure or oronasal thermistor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | s |  | The system must allow users to select between automatic and manual analysis |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | t |  | The user must be able to add annotations to indicate events when reviewing the waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | u |  | Automatic analysis must be made by the equipment using criteria setup by the operator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | v |  | Type 1 systems must allow for real time manual analysis on the record |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | w |  | Parameters produced in a PSG report must meet ASA/AASM guidelines |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | x |  | It must be possible to add Event Markers to the waveforms in the form of free text |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | What free text Event Markers are available
|  | y |  | The list of common event markers should be revised to match AASM recommended calibrations
|  | z |  | It must be possible to manually enter Patient demographics |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | aa |  | Manual Entry must be an alternative |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Annotations alongside the relevant piece of the waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | ab |  | Time and date must be automatically entered along with the waveform |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Manual setting of time and date must also be possible |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 7.1 |  |  | PSG Digital Videometry
|  | a |  | A camera with the following features accessories is required: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Minimum 1080P video resolution. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Day/night switch |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Autofocus |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Auto brightness |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Remote Pan/Tilt/Zoom controlled remotely |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | An appropriate mounted microphone |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | A camera mount (the location of which must be agreed with end-users) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The review software must have at least the following video editing functions:
|  |  | i | Copy, cut and delete; |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Marking; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Crop |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The system must be able to synchronise the EEG or PSG waveforms and video images: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | throughout both normal playing and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | fast forward/reverse recording |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The system must allow the operator to select a review in any of these formats: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Video image and waveforms in one screen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Video images only; and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Waveforms only |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Digital video must be recorded using MPEG4 compression. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Please describe an alternate method of compression.
| 8 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display modes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | List all parameters that are monitored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Number of channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | state CRR |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | amplifier sensitivity range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | List amplifier Low frequency setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | List amplifier High frequency setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | state the input impedance |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | List all Patient personal information/demographics fields |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | State the video resolution of camera |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xix | IPX rating of unit (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xx | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxi | State the Type (Level) of sleep study capabilities of the tendered PSG system; (1 – 4) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 6.04 - Electromyography (EMG) / Evoked Potential (EP) Systems - Nerve Conduction Study (NCS)
| 9 |  |  | General Requirements
|  | a |  | The system offered must be capable of receiving the following information electronically from a Departmental Information System: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Patient Demographics, i.e. Name, DOB, Gender, Address |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Patient Weight |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Patient Height |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Physician Name |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Technician Name |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Hospital Name |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Patient Medical Record Number (MRN) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Study Request Information |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Other Please describe
|  | b |  | Manual Entry must be an alternative |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The system must continuously display patient information, test name and nerve being tested |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The EMG system must have inbuilt speaker |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The system must have foot switch for start/stop/save function |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.1 |  |  | EP/EMG Amplifier
|  | a |  | The system must provide up to 8 Independent channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Sensitivity must be adjustable from 10 mV to 100mV/division over 22 steps |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Frequency range must be at least 0.2 to 20,000 Hz |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The low cut and high cut filter must be user adjustable |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The noise level from internal components must be <2 µVp-p |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The Common Mode Rejection Ratio (CMRR) must be ≥ 110 dB at 50 Hz |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The input impedance must be ≥ 100 MΩ what is the impedance offered (MΩ) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | The Sweep speed must be at least between 1 ms to 50 s |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The amplifier box must have an ON/OFF switch |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | For EP's, the head box must have 22 electrode inputs configured according to the 10/20 EEG electrode layout |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Users must be able to relabel each electrode input using a writable overlay |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.2 |  |  | EP/EMG A/D Converter
|  | a |  | Analogue/Digital Converter resolution must be >24 bits |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Sample rate per channel must be > 48 kHz |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.3 |  |  | EP/EMG Electrical Stimulator
|  | a |  | The unit must have constant current and constant voltage mode |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The duration must be at least between 0.02 to 1 ms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The intensity must be between 0 and 100mA or between 0 and 400V |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The stimulus mode can be set to either monophasic or biphasic stimulation using: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Single |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Double (Pair) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Train |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The stimulus rate can be set to nonrecurrent or recurrent stimulation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The stimulus rate can be varied between: 0.06 and 200 stimuli/sec (Hz) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Stimulus control and basic functions must be available from the stimulator probe, with programmable buttons |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.4 |  |  | EP/EMG Auditory Stimulator
|  | a |  | Noise masking insert earphones, or isolation headphones must be provided |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The intensity must range from:
|  |  | i | 0 to 139 dB Sound Pressure Level (SPL), or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | -31 to 109 NHL |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must have pre-set click duration setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Pulse rate must be at least between 5 and 50 pps |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Click, Tone Pip, and Tone Burst must be available |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Must have numerous pre-set tone frequency setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Output frequencies can be set to various tone envelopes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.5 |  |  | Visual Stimulator
|  | a |  | A Visual Evoked Potential (VEP) monitor for: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Black and white simulation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Pattern reversal simulation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Checkerboard simulation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | State the video signal formats available, e.g. VGA, DVI, HDMI |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Reversal rate must be at least between 1 50 reversals/sec |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | LED Goggles for Flash Visual Evoked Potential (FVEP) must be offered |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 9.6 |  |  | EP/EMG Software Capabilities
|  | a |  | The following Software functionality is required:
|  |  | i | Motor Nerve Conduction (MNC) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Sensory Nerve Conduction (SNC) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Combined Sensory Index |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Combined Motor and Sensory Nerve Conduction |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Inching Studies |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | FWave |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | HReflex |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Blink Reflex |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Repetitive Nerve Stimulation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Needle EMG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | MultiMUP Analysis |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Single Fibre EMG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | Macro EMG |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | Turns and Amplitude |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | AEP |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | SEP |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | VEP |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | P300/CNV |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xix | RR Interval |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xx | Sympathetic Skin Response (SSR) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxi | Galvanic Skin Response (GSR) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxii | MUNIX and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxiii | Tremor |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Free run EMG data and sound can be recorded for nominally:
|  |  | i | Up to 960 seconds for 2 channels, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Up to 360 seconds for 8 channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Stored data can be:
|  |  | i | re-analysed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | digitally filtered |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | smoothed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | inversed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | summed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | replayed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | displayed as trends |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | displayed in plots |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | frequency analysis |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Other (Please describe)
|  | d |  | Roll Back and Roll Forward features will automatically store previous responses ensuring that the best response is available eliminating unnecessary stimulations |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Up to four replications are available allowing the user to quickly verify a small pathological biological response with an easy way of selecting what result to report how many replications are available |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Signal Enhancer highlights clinically relevant data to simplify analysis and measurements |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | In SNC, it reduces stimulus artifacts to yield a better baseline |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | In F-waves it will stabilise the base line after the MWave during the time of the Fresponse making it easier to identify the response and place markers |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | This feature is non-destructive and can be turned on or off by the user |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Reporting online result views give a compact clinical overview with links back to the raw data |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 10 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iiii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | size of tendered device, L x W x H (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | weight of tendered device (kg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Display size, H x W (mm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display type e.g., OLED, LCD, LED |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • display modes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | List all parameters that are monitored |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • state measurement range of each parameter |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Number of channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | amplifier sensitivity selectable settings |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | frequency range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | CMRR (dB) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | sweep speed range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xv | noise level of offered system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvi | A/D converter resolution |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xvii | A/D converter sample rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xviii | Electrical stimulator duration range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xix | Electrical stimulator output |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xx | Electrical stimulator stimulus rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxi | describe auditory stimulator noise mask |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxii | Auditory stimulator entity range (SPL or NHL) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxiii | List auditory stimulator pre-set click duration settings |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxiv | Auditory stimulator pulse rate range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxv | List auditory stimulator pre-set tone frequency setting selections |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxvi | List auditory stimulator frequency tone envelopes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxvii | Software records EMG data for (sec) for number of channels |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxviii | List all alarms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxix | IPX rating of unit (IPxxx) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xxx | Battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • operating time on battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • time to recharge fully depleted battery (hrs) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • estimated battery lifecycle (in years) under normal use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 6.05 - Clinical Neurology Management Systems
| 11 |  |  | Standards and Compliance
|  | a |  | Tendered System must only collect, store, and access the minimum amount of patient's personal data, necessary to perform its function. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List the patient's personal data to be collected, stored and accessed
|  |  | ii | Stored patient's personal data must be encrypted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If Tendered Systems are cloud based, then the location of the cloud-based servers must be within the boarders of Australia. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services may prefer that the cloud based servers are located within the State of Victoria. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must maintain and ensure its permitted third-parties, maintain the physical and virtual security of any facilities owned, managed, licensed, or controlled by Respondent that stores patient data by implementing industry best security practices at the locations where patient data is stored to ensure the confidentiality, integrity, and availability of the patient data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondent, where applicable, must maintain backup policies and procedures of Data containing patient's personal data, image repositories, and provisioned environments. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The backup storage infrastructure must be located in physically protected, limited-access facilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Prior to disposal of any Device returned or discontinued by or on behalf of the Participating Health Service, Respondents must securely wipe or destroy all patient data consistent with industry standards, such as NIST 800-88, to ensure that no patient data is retrievable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Upon request of the Participating Health Service, Respondent must provide confirmation of patient's personal data wipe or destruction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | General Requirements
|  | a |  | Tendered System must provide the capability to acquire, process, store /retrieve results in a fast, reliable, and accurate manner |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | System must be capable of receiving data from multiple devices and formats (support Multi-vendor and Multi-modality) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | System must be capable of receiving and analysing data acquired from the patient |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Solution must provide a range of Physician Report review tools. Tenderer to list all available tools.
|  | e |  | The Solution must allow for role based customisation of the GUI for viewing and interpretation of data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The Solution must provide an electronic signature for any final report generated by the system. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12.1 |  |  | Optional Configurations
|  | a |  | Clinicians should be able to access from anywhere and on any mobile device without the use of additional applications that depended on device operating system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13 |  |  | ICT
|  | a |  | Access to data repository must only be by system authorised users. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13.1 |  |  | Data Recording
|  | a |  | Web-enabled formats are preferred. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13.2 |  |  | Data Transmission and Interfacing
|  | a |  | Data transmission must ensure patient privacy. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Clinical systems are being integrated to other patient information systems including HIS, EMR etc. Systems must be using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | System must allow bidirectional data flow |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 13.3 |  |  | Data Hardware and Software
|  | a |  | Authorised users must be able to upgrade software, troubleshoot and configure the system settings. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Printouts must be of an appropriate diagnostic quality. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The ability to print to non-thermal paper is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All connecting devices must have standard TCP/IP connections. Cabling installation and management may become a requirement of the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Permissions and Participating Health Services’ protocols must be adhered to. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Vendor supplied networking equipment and servers should be located in the Participating Health Services’ ICT areas, such as communications rooms and server farms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | For Vendor supplied networks, all connecting cables must be provided and installed by the successful Respondent(s) (including fibre optics, where applicable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | For Hospital supplied networks, all patch cables must be provided by the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Patch cables must be green, unless specifically requested to be another colour by a Participating Health Service. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Upon purchase of any Clinical Neurology Management system, successful Respondents must provide to Participating Health Services:
|  |  | i | technical diagrams of the installation, including a network diagram |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a technical address spreadsheet with all IP, subnet masks and MAC addresses (where appropriate locations should be indicated). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | System must be scalable to increasing workload by means of adding additional servers which operate in parallel whilst providing business continuity without having to replace the server. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 14 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | list any known incompatible devices |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 7-GasDevices
Category 7: Basic Gas Delivery Devices
Clause Identifiers |  |  |  | Text | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 7.0 |  |  |  | Basic Gas Delivery Devices | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | General Requirements
|  | a |  | Tendered device must either connect to a medical gas wall-outlet or medical gas bottle-outlet. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Basic gas delivery devices must be designed to be used with standard Australian gas fittings. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All devices must meet the requirements of ISO 32, | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | particularly those related to colour coding. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All accessories must be colour-coded to the type of gas they are intended to be used with. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Connections for medical gas cylinders, yokes or regulators must be pin indexed; as per ANZCA PS54. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Non-interchangeable gas hose connectors must be present on all gas inlet and outlet sockets to prevent incorrect gas supply connections; as per ANZCA PS54. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | All hand wheels must be easy to transfer between products and of a sufficient diameter to increase ease of use. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Respondents must complete Category 18 for any service related services that they may provide | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 1.1 |  |  | Regulators
|  | a |  | Regulators must be easily connected and disconnected from the gas source without a tool. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Regulators must be provided with a safety relief valve in case of over-pressure. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Suction regulators may require other options, such as filters or reservoirs. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 1.2 |  |  | Flow Meters
|  | a |  | Flow meters must provide an accuracy of ± 5%. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Flow meters must provide clear indication of flow. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondents may offer ‘step’-type flow meters. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Flow meters must provide a fail-safe system in case of excessive gas pressure. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tapered outlet fittings are preferred. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Control knobs must be accessible and large enough to be used by a variety of operators. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Low flow flowmeters are considered as delivering flows of, or less than 3 LPM. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | High flow flowmeters are considered as delivering flows greater than 15 LPM | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 1.3 |  |  | Hoses
|  | a |  | Respondents are encouraged to offer reusable accessories such as three-way adaptors and extension hoses. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All hoses must comply with AS 2902; this includes colour-coding. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All crimp-on connectors must meet the requirements of AS 2902. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Product Information
|  | a |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Tender Response Worksheet:
|  |  | i | Manufacturer | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Dimensions (mm) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Weight (g) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | For tendered regulators
|  |  |  | • cylinder connector type | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • cylinder gauge (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • output pressure, preset or adjustable (PSI) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | For tendered regulator/dial flow meters
|  |  |  | • cylinder connector type | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • cylinder gauge (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • preset pressure (PSI) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • flow settings | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • flow accuracy | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • flow meter outlet | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | For tendered flow tube & ball flow meters
|  |  |  | • flow range | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • graduation subdivisions | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • accuracy | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • set pressure (PSI) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • flow meter outlet connection | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | For tendered suction/vacuum regulators
|  |  |  | • digital/analogue | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • vacuum range (mmHg) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • accuracy | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • modes (cont./off/inter.) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • maximum vacuum (mmHg) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • overflow safety trap (yes/no) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | - trap volume | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | For tendered adapters
|  |  |  | • T or Y piece | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiii | For tendered medical gas hoses
|  |  |  | • hose length | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • hose colour | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • hose connections on each end | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xiv | ARTG number | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 8-Suction
Category 8: Portable Suction Units
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 8.0 |  |  |  | Portable Suction Units | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Portable Suction Devices, Gas-Driven (e.g. Twin-o-Vac®)
|  |  | i | Portable Suction Devices, Gas-Driven, are medical devices used to clear the airway or remove unwanted fluids from parts of the body |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Manual Suction Units
|  |  | i | Portable Suction Devices, hand or foot activated, are medical devices used to clear the airway or remove unwanted fluids from parts of the body |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Aspirators, Airway, mains powered
|  |  | i | Portable Suction Devices, driven by AC power are medical devices used to clear the airway or remove unwanted fluids from parts of the body |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Aspirators, Airway, battery powered
|  |  | i | Portable Suction Devices, driven by DC power are medical devices used to clear the airway or remove unwanted fluids from parts of the body |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 8.01 - Portable Suction Devices, Gas-Driven (e.g. Twin-o-Vac®)
| 2 |  |  | General Requirements
|  | a |  | The gas drive must be O2, unless indicated otherwise. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Portable suction devices must include a bacteria filter. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The receiver reservoir must be a minimum of 100mL. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Portable suction devices must have variable suction regulation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Preference may be given to devices that can be configured to either low or high suction (low is considered below 200mmHg of suction). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The efficiency of drive gas will be considered in the evaluation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Respondents may offer optional suction gauges. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Respondents may offer optional O2 outlets. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 8.02 - Manual Suction Units
| 3 |  |  | General Requirements
|  | a |  | Vacuum must be manually generated by use of either the hand or foot |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Hand operated devices must be light weight, <0.5kg with an empty cannister |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Product Information
|  | a |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | cannister capacity |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | vacuum range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | peak flow rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 8.03-8.04 - Airway Aspirators
| 5 |  |  | General Requirements
|  | a |  | Tendered device must consist of a:
|  |  | i | vacuum pump |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | regulator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | vacuum gauge |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | collection canister |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | carry case/handle |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must have a vacuum range of at least 0 to ≥400 mmHg |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Must have a flow rate ≥25 L/min |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Vacuum gauge must have an accuracy of within ±10% |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Must be able to regulate/control vacuum level |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Collection canister must:
|  |  | i | have a capacity ≥1,000 mL |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | provide overflow protection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | have a disposable liner |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | All patient side tubing must be single use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.1 |  |  | Battery Powered Airway Aspirators
|  | a |  | Must operate at maximum vacuum for more than 30 minutes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Must have a low battery indicator, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | should also include an audible signal. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | should have a battery indicator (battery gauge) showing the remaining battery capacity |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Must have a battery charging indicator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The total weight of battery powered airway aspirator (including all available optional configurations, batteries and carry bag) must not exceed 7 kilograms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | pump type (e.g. diaphragm, piston, rotary vane, etc.) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Collection canister capacity (mL) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Filter type on canister |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Vacuum range (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | Time to 300 mmHg (sec) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | Vacuum gauge diameter (cm) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | Vacuum gauge scale (mmHg) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 9-N2O_O2
Category 9: Nitrous Oxide and Oxygen Delivery Systems
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 9.0 |  |  |  | Nitrous Oxide and Oxygen Delivery Systems | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Analgesia Units
|  |  | i | Gas Analgesia Units are medical devices used to administer a mixture of gases for pain relief |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | 50/50 Oxygen/Nitrous Oxide Demand Valves
|  |  | i | A device used to administer a mixture of 50/50 nitrous oxide and oxygen, for pain relief |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Base Unit
|  | a |  | The base unit configuration for a nitrous oxide and oxygen delivery system must:
|  |  | i | be designed to be used with standard Australian gas fittings |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | meet the requirements of ISO 32, particularly those related to colour coding |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | be easily connected and disconnected from the gas source without a tool |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | be provided with a safety relief valve in case of over-pressure. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Connections for medical gas cylinders, yokes or regulators must be pin indexed; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Non-interchangeable gas hose connectors must be present on all gas inlet and outlet sockets to prevent incorrect gas supply connections; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All hoses must comply with AS 2902; this includes colour-coding. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 9.01 - Analgesia Units
| 3 |  |  | General Requirements
|  | a |  | Tendered systems must delivery a minimum 30% O2. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Proportional gas mixing must be clearly indicated. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | There must be an audio alarm to indicate O2 gas supply failure. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Demand valves must be used to ensure that minimal amounts of nitrous oxide are leaked to the environment. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Nitrous oxide and oxygen delivery systems must have emergency oxygen delivery capability. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered device may use bottled medical gas. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Participating Health Services may preference device with digital interface. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 4 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | maximum flow rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • combined |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • oxygen |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Oxygen concentration range |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • accuracy |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Trigger sensitivity (cmH2O) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Oxygen flush flow rate (L/min) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | wall mount / trolley mount |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 9.02 - 50/50 Oxygen/Nitrous Oxide Demand Valves
| 5 |  |  | General Requirements
|  | a |  | The demand flow rate must be a minimum of 150mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Demand flow pressure must be -5mmHg at 100mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Exhalation resistance must be -5mmHg at 50mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The system must have appropriate gas fittings for either wall or bottle supply (if the system has fittings for bottle supply, then the system must include a dedicated regulator). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | demand flow rate |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Demand trigger |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Exhalation resistance |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 10-Anaesthetic
Category 10: Anaesthesia Workstations
Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
Cat 10.0 |  |  |  | Anaesthesia Workstations | Supporting Documentation Submitted | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Anaesthesia workstation, Low Acuity
|  |  | i | Low-acuity anaesthetic machines tend to have a smaller footprint, likely to be used in an ambulatory surgery setting, and treat patients who are less likely to be needing advanced ventilation modes. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Anaesthesia workstation, General-Purpose
|  |  | i | General-purpose anaesthetic machines are intended for use with the full range of patients, and a full range of procedures, with ventilation modes comparable to most intensive care ventilators. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Anaesthesia workstation, MRI
|  |  | i | Anaesthetic machine intended to be used in an MRI suite. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Anaesthesia workstation, Analytics Applications
|  |  | i | Database for anaesthesia care information acquired from anaesthetic machines. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | Data Management
| 2.1 |  |  | Data Recording
|  | a |  | Anaesthetic machines with optional data recording must incorporate date and time recording of patient observations and anaesthetic machine settings (when available). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Additional configurable parameters are desirable, such as weight, pain etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.2 |  |  | Data Transmission
|  | a |  | If data transmission is available, anaesthetic machines must be able to transfer data via industry standard connections. For example: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Wi-Fi |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• state Wi-Fi technology used
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | ii | RJ45 (Ethernet) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• specify the broadcasting features; bandwidth, multicast or unicast
|  |  | iii | Other
|  | b |  | Data transmission must ensure that the privacy of the patient's personal information is not breached. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | All data output ports must be activated and ready for use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Data transmission to other systems is preferred in standardised languages, such as Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Data integration with HIS (as defined by the Participating Health Service) may be offered as a purchasable option. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | If a separate annual licensing arrangement is required, a minimum of one (1) year must be included for free upon purchase. Details of all conditions and subsequent annual costs must also be provided upon purchase. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.3 |  |  | Event Reporting
|  | a |  | All anaesthetic machines with data transmission capabilities must have the capability to pass event data wirelessly or via cable to other devices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A summary report should be downloadable from the anaesthetic machine in a universal format so that it may become a component of hospital records. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Event information must be able to be stored on Participating Health Services’ information systems. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | For research and development purposes, data format should be open ECG-SCP, RCG or similar Open Same Standard format. If only a proprietary format is available, a software development kit for Microsoft Windows should be provided as a purchasable option. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Patient data that is stored on the internal memory should be encrypted to an open standard (i.e. non-proprietary). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | A data copy must be able to be stored in such a manner that it is incorruptible for use in any potential legal analysis. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.4 |  |  | Optional Data Configurations
|  | a |  | It is desirable that data management systems and central stations can include the following optional configurations upon request by Participating Health Services: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | the ability to discharge/admit patients from HIS |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | full disclosure of anaesthetic information for the case, preferably with waveforms |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | web services (for remote viewing by medical staff) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | integration with other systems |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | printer. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All connecting cables must be provided by the successful Respondent(s) (including fibre optics, where applicable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.5 |  |  | Point-of-Care Integration
|  | a |  | Point-of-care integration is an important feature for some Participating Health Services. This feature provides all patient-related information to the bedside for timely decision making for clinical staff (e.g. path results, PACS, HIS).
|  | b |  | Point-of-care integration systems must be able to:
|  |  | i | follow standard TCP/IP protocols |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | interface with Participating Health Services’ intranet. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Preference may be given to systems that have proven records of integration with path systems, PACS systems, Anaesthesia Record Keeping Applications, Anaesthesia Analytics Applications, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 10.01 - Anaesthesia workstation, Low Acuity
| 3 |  |  | General Requirements
|  | a |  | Tendered workstation all must have an electronic ventilator suitable for use on adult, paediatric and neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered workstation must include the following modes of ventilation:
|  |  | i | manual or spontaneous ventilation, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | volume controlled ventilation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered workstation must include a display monitor for the display of information/waveforms etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | display monitor must be colour |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered workstation must electronically log: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | the machine self checking process |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | errors including the ability to view retrospectively |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | alarms including the ability to view retrospectively |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tendered workstation must be able to connect to printers |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered workstation must have the ability to monitor all necessary ventilation and gas monitoring requirements with integrated alarms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Tendered workstation must be equipped with controls that are protected against accidental setting changes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Every workstation must include integral monitoring of the ventilator inspiratory and expiratory flows, volumes and airway pressure. (SI derived units) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | In the event that the castors encounter an obstruction, such as a hose on the floor, lift threshold, etc., The Anaesthetic Workstation must be stable and resist tipping when fully loaded. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | When each feature of the anaesthetic machine is enabled, any associated monitor and alarm functions must be automatically activated; as per ANZCA PS 54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | In case of total electrical and/or electronic failure, it must be possible to mechanically control the medical gases. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.01 |  |  | Power Supply
|  | a |  | An ‘on/off’ switch, if present, must be protected from unintended activation or deactivation; as per ANZCA PS54 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services will preference mechanical protection of the mains 'on/off' switch, and therefore |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must offer Participating Health Services the option to purchase, mechanical protection covers for the mains ‘on/off’ switch. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Switching ‘off’ an electronic anaesthetic machine during normal operation must require a confirmatory step and/or the machine should display a warning of imminent shut-down; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | If the anaesthetic machine requires electrical power for normal operation, a backup power supply must be a part of the machine and permit normal operation for at least 30 minutes after a mains power supply failure. An alarm must be activated at the time of the mains failure and the state of the reserve power supply must be indicated while it is in use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State length of time anaesthetic machine fully functional on backup power in the event of mains failure.
|  | d |  | Tendered workstation must indicate:
|  |  | i | whether running on mains power or battery power |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | charging status |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | quantity of remaining battery charge |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Anaesthetic machine should have power outlets to power ancillary devices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The anaesthetic machine trolley should have provisions to fit/install a UPS, to power ancillary devices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.02 |  |  | Trolley
|  | a |  | Tendered workstation must be supplied with and mounted on a trolley |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The trolley must be equipped with:
|  |  | i | a drawer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a shelf for mounting of ancillary equipment and accessories |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | a bracket for mounting a computer/keyboard/monitor arm |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | a writing surface |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | a work surface for drug preparation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | integrated mounting rails for accessories |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | an integral cable and hose management solution |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Patient Tubing Support Arm adaptable either to the Left or Right side |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Humidifier holder adaptable to either the left or right side |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | provisions for mounting infusion/syringe pumps |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | for models released from 2024 onwards, provisions allowing for the mounting of up to three (3) computer monitor displays |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Trolley castors must be greater than 125mm to provide easy access across a lift threshold, and |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | at least two castors must be fitted with brakes. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.03 |  |  | Gas Systems
|  | a |  | The anaesthetic machine must work with a wall gas supply from 280kPa–600kPa. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The Anaesthetic Workstation must be equipped with flexible supply hoses (min 3000 mm in length) for each gas fitted conforming to AS 2896 and utilising Sleeve Index System fittings |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The Anaesthetic Workstation must include connection to Medical Air, Medical Oxygen, Medical Nitrous Oxide and Medical Suction |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Anaesthetic Workstation must be equipped with Indicators for each gas fitted displaying supply pressure status |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The Anaesthetic Workstation must be equipped with a suction reservoir and suction system suitable for pharyngeal suction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The Anaesthetic Workstation must be equipped with provision for reserve Medical Oxygen supply bottles (both mounted and connected) must use the Pin Indexed System (AS 2473.3). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | A reserve supply of oxygen must be attached to the anaesthetic machine in a manner that ensures easy activation should the oxygen supply failure warning system [refer to clause 5.37.13] indicate impending failure of the external oxygen supply, as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | The Anaesthetic Workstation should be equipped with provision for reserve Nitrous Oxide supply bottle/s and (both mounted and connected) must use the Pin Indexed System (AS 2473.3). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Reserve supply bottles (both mounted and connected) must use the Pin Indexed System (AS 2473.3). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The gas supply must be arranged so that when there is a correctly working gas supply from the mains pipeline, this supply must be used in preference to any reserve gas supplies connected. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | The reserve gas supply cylinder contents, even with the cylinder turned on, must not be diminished by any intentional bleed within the system whether the Anaesthetic Workstation is turned on or off. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | When high-pressure gas supply systems are in use, an oxygen supply failure warning device must be present on the anaesthetic machine. This must: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | activate automatically when the oxygen supply pressure falls below a predetermined critical level |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | generate an alarm to warn the operator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | cut off the supply of gases other than air or oxygen to the common fresh gas outlet |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | cancel the alarm only when the oxygen supply pressure has been restored to a level above that at which the device was activated; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | If the anaesthetic machine incorporates a gas flowmeter bank, oxygen must be the last gas to enter the common gas manifold at the top of the flowmeter tubes. The oxygen flow knob must be the first from the left on the rotameter, relative to other flow knobs; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | If mechanical means are provided to mix the anaesthetic gases on the anaesthetic machine, there must be only one gas flow control knob for each gas; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | If a mechanical oxygen flow knob is provided, it must differ from the other flow control knobs so that tactile identification of the oxygen control knob is possible (for example, fluted); as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | If the anaesthetic machine is capable of delivering nitrous oxide, the machine must not deliver a hypoxic mixture. When oxygen and nitrous oxide are the only gases used, the machine must prevent delivery of a gas mixture with an oxygen concentration below that of ambient air; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | p |  | Anaesthetic gas scavenging system connections must be of a diameter that is different from the other connections used for the breathing system; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | The gas supply must be arranged so that unidirectional flow control devices prevent a contaminating flow from one gas supply into another either in normal operation or in the event of failure of any part of the system or any gas supply. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | r |  | Electronic gas flow controls must have a display that clearly indicates which gas is being adjusted and what the current flow level is for all gases. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | s |  | Direct mechanical flow controls must continuously increase gas flow when turned in an anti-clockwise direction and continuously decrease the gas flow when turned in a clockwise direction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | t |  | Flowmeters must display a graphical representation of flow for each gas. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | u |  | A separate flow indicator must be provided for each gas. The scaling must be in units of litres per minute |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | v |  | The flow measurement range for each gas must be 0 to 10 litres per minute |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | w |  | The flow measurement accuracy and resolution must be <1% |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | x |  | A high-priority alarm must be activated when high airway pressure is present; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.04 |  |  | Breathing System
|  | a |  | The Anaesthetic Workstation must be equipped with an emergency oxygen flush button with means to prevent accidental activation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The emergency oxygen flush control must be protected from accidental activation; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Oxygen flush rates should be capable of being greater than 35L/minute. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | One floating ball type oxygen flow regulated outlet delivering at least 10 L/min must be provided in a convenient location |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | A high-pressure relief valve or other means of automatically preventing dangerously high pressures in the breathing system must be present in the breathing circuit but it is not necessary for such a device to be integral to the anaesthetic machine; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Breathing systems must be easily dissembled to for cleaning purposes. Breathing systems must also be able to endure standard hospital cleaning procedures and materials. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | If a specific tool is required for the removal of the breathing system, the tool must be provided with the system. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | It must be quick and easy to change from mechanical ventilation to hand ventilation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | An Adjustable Pressure Limiting (APL) valve must be provided with a range of 0.5–70cm/H2O. Preference may be given to non-linear adjustable valves. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | A fresh gas outlet, if provided, must be 22mm outer diameter and 15mm inner diameter, visible to the operator and should be capable of being connected to the breathing system in such a way as to prevent accidental disconnection; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | A high-priority alarm must be activated when the airway pressure falls 10cm H2O below atmospheric pressure for more than one (1) second; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | The common gas outlet must have a patient safety relief valve. Specify the pressure at which the patient safety relief valve opens (and the nominal flow at which that specification is determined). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Continuous ventilation of a patient using a humidified patient circuit must be possible without damage to the ventilator system or affecting its accuracy, operation without adversely affecting ventilator performance, accuracy, operation or long term reliability. Please list compatible humidifiers. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | The delivery system must be fitted with either a refillable carbon dioxide absorbing canister having sieves and a clear polycarbonate container or a replaceable carbon dioxide absorbing canister. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | The absorber system must be autoclavable for the purpose of decontamination. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | p |  | Absorber must have Self-sealing mechanism to allow the exchange of canisters during the case without generating a leak in the system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | Provision (including conveniently located connectors) must be made to direct gases sampled from the patient circuit by the airway gas monitor either to: a) the scavenging circuit, or b) returned to the patient circuit via the ventilator, (user selectable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | r |  | An integrated mechanical airway/circuit manometer gauge is desirable near the circuit attachment points. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.05 |  |  | Ventilation
|  | a |  | Provision must be made for manual ventilation of the patient using a manual ventilation bag and manually Adjustable Pressure Limiting (APL) valve (hPa). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | In the VC mode, the delivered tidal volume must be independent of the rate of fresh gas flow (i.e. compensation provided to ensure delivered tidal volume equals selected tidal volume independent of changes in fresh gas flow). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | In the VC mode, the delivered tidal volume must have compensation provided for circuit compliance and leaks (as measured during system self test) to ensure delivered tidal volume equals selected tidal volume independent of circuit compliance and leaks). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | It must be possible to configure the default mechanical ventilation mode and settings, which are delivered at start-up. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tendered Workstation must deliver a Tidal Volume (TV) of 20 to 1400 mL in Volume Controlled Ventilation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Ventilators must be able to ventilate in a range of respiration rates from 4–60BPM. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Ventilators must be able to provide Positive End Expiratory Pressure (PEEP) of 5–30cmH2O. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Ventilators must be able to sustain a leak of <= 150mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Ventilators must provide minimum gas flow of 120mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | A high airway pressure alarm must result in the cessation of an inspiratory breath and the opening of a safety valve to relieve the pressure. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The airway pressure relief must be in addition to, and at a level above, the controlled pressure level in the time cycled pressure-controlled mode of ventilation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | All current settings of the ventilator must be shown simultaneously and continuously during operation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | Ventilator settings and measured values must be clearly differentiated from other monitoring values and waveforms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | The ventilator must incorporate an expiratory tidal volume monitoring and alarms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | The Anaesthetic Workstation must provide a means to allow switching between manual and automatic ventilation without the disconnection of hoses |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.06 |  |  | Airway Gas Monitoring
|  | a |  | A comprehensive set of alarms for the airway gas monitoring must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A side stream sampling technique must be used to provide a source for an integrated airway gas monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The sampling rate must be between 50 and 200 mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Airway gas monitoring must commence automatically and alarms must be automatically enabled either when the machine is turned on or upon the detection of the first breath. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | A water trap must be provided to protect the measuring units. Please specify if single use or reusable water traps are available. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | A connector must be provided for exhaust sample gas to be returned to the anaesthesia circuit or directed to the scavenging system. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Each waveform must have a user adjustable scale. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Verification of calibration by the user must be possible. Please describe the calibration method |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The Respondent must specify the method used for calibration. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | User initiated calibration at both 21% and 100% oxygen must be possible. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.07 |  |  | Display
|  | a |  | Operating buttons for soft-key controls must clearly align with on-screen labels without any parallax error. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | It must be possible to measure and display airway pressures, inspiratory volumes, expiratory volumes, pressure-volume loops, flow-volume loops, airway pressure waveforms, airway flow waveforms, compliance and airway resistance. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | A means of displaying ventilator flow and pressure waveforms to the patient must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All waveforms must be identified by an on-screen label. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The monitor must display, in a digital form, data derived from each waveform displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Digital values must automatically align with associated waveforms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Both inspiratory and expiratory tidal volume (TV) and minute volume (MV) values must be displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.08 |  |  | Alarms
|  | a |  | The Anaesthetic Workstation must automatically activate associated ventilator monitors and alarms when machine is turned ON |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A clear visual and audible alarm must be generated when there is:
|  |  | i | a sustained positive pressure in the breathing system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a high peak pressure in the breathing system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | sustained sub atmospheric pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | loss of mains power |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | low Battery |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | loss of mains gas pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | non-cycling |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | circuit disconnection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | low/high expiratory volume |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | apnoea. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• It is desirable for an indication of the cessation of apnoea. Please specify.
|  | c |  | Ventilator alarm limits must be displayed at all times (see below). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Ventilator alarms must be prioritised according to level of urgency. Priorities of alarms must change in both tone and visual interpretation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The ventilator ALARM SILENCE button must not silence any alarm for a period exceeding two (2) minutes. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The ventilator ALARM SILENCE button must not silence high priority alarms for a period exceeding one (1) minute (e.g.. apnoea). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The ventilator ALARM SILENCE button must not remove the visual warning of an existing alarm condition. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.09 |  |  | Vaporisers
|  | a |  | If more than one delivery system can be simultaneously mounted on the anaesthetic machine, a interlock system must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Vaporisers with mechanical adjustment dials, when used in high-pressure circuits, must increase the delivered anaesthetic vapour concentration when the dial is rotated in an anti-clockwise direction; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The minimum type of anaesthetic vaporiser that Respondents must offer is isoflurane, sevoflurane and desflurane. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Anaesthetic Workstation must deliver volatile anaesthetic agents with an accurate mechanism |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The Anaesthetic Workstation must have safe guards - to prevent simultaneous dual vapour administration and this must not be able to be bypassed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Anaesthetic Workstation equipped with vaporisers with mechanical rotary dials must increase the delivered vapour concentrations when the dial rotated in anti-clockwise direction and must have a detent for the off position. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The Anaesthetic Workstation must compensate for low gas flows and variations in room temperature to allow precise delivery of volatile anaesthetics at minimal gas, fresh gas flow rates |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | The Anaesthetic Workstation must display the volatile agent - currently in use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The Anaesthetic Workstation must facilitate easy and quick change of agent during operation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | The vaporiser must be easily removed for filling and maintenance. A secure latching mechanism must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | The vaporiser must have an agent specific filling system to prevent inadvertent use of the incorrect agent. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | The vaporiser must have a liquid level indicator clearly and indelibly marked with maximum and minimum filling levels. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Where mechanical vapour control is offered, each workstation must provide for the mounting of two anaesthetic vaporisers for ease of transfer between volatile agents. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | A safety mechanism to prevent accidental operation of the control dial must be provided |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | The vaporiser must be clearly and indelibly marked with the generic name of the anaesthetic agent and colour coded according with the appropriate Australian Standards. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 3.1 |  |  | Patient Monitor Integration
|  | a |  | The Anaesthetic Workstation must accommodate a patient monitor and integrate seamlessly. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Please list all the compatible patient monitors
|  | b |  | The Anaesthetic Workstation must support the weight of the patient monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Please specify the max weight
| 4 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | and contain a complete set of technical specifications which describe the inbuilt SPIROMETRY ventilator monitoring hardware and software performance, including measurement range, measurement accuracy and alarm capabilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Details on flushing requirement to remove traces of volatile agent in preparation for Malignant Hyperthermia patients. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 10.02 - Anaesthesia workstation, General-Purpose
| 5 |  |  | General Requirements
|  | a |  | Tendered workstation all must have an electronic ventilator suitable for use on adult, paediatric and neonatal populations. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered Workstation must be microprocessor controlled with flow and pressure sensors, and electronically controlled valves |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Tendered workstation must be capable of advanced modes of ventilation including: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | manual or spontaneous ventilation, |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | volume controlled modes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | pressure controlled modes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | pressure support modes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | CPAP, PEEP, etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | synchronise to spontaneous respiration |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered workstation must include at least two (2) display monitors for the display of information/waveforms etc. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | display monitors must be colour |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Tendered workstation must electronically log: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | the machine self checking process |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | errors including the ability to view retrospectively |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | alarms including the ability to view retrospectively |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Tendered workstation must be able to connect to printers |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Tendered workstation must have the ability to monitor all necessary ventilation and gas monitoring requirements with integrated alarms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Tendered workstation must be equipped with controls that are protected against accidental setting changes |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Every workstation must include integral monitoring of the ventilator inspiratory and expiratory flows, volumes and airway pressure. (SI derived units) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | In the event that the castors encounter an obstruction, such as a hose on the floor, lift threshold, etc., The Anaesthetic Workstation must be stable and resist tipping when fully loaded. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | When each feature of the anaesthetic machine is enabled, any associated monitor and alarm functions must be automatically activated; as per ANZCA PS 54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | In case of total electrical and/or electronic failure, it must be possible to mechanically control the medical gases. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.01 |  |  | Power Supply
|  | a |  | An ‘on/off’ switch, if present, must be protected from unintended activation or deactivation; as per ANZCA PS54 |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services will preference mechanical protection of the mains 'on/off' switch, and therefore |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must offer Participating Health Services the option to purchase, mechanical protection covers for the mains ‘on/off’ switch. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Switching ‘off’ an electronic anaesthetic machine during normal operation must require a confirmatory step and/or the machine should display a warning of imminent shut-down; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | If the anaesthetic machine requires electrical power for normal operation, a backup power supply must be a part of the machine and permit normal operation for at least 30 minutes after a mains power supply failure. An alarm must be activated at the time of the mains failure and the state of the reserve power supply must be indicated while it is in use. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | State length of time anaesthetic machine fully functional on backup power in the event of mains failure.
|  | d |  | Tendered workstation must indicate:
|  |  | i | whether running on mains power or battery power |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | charging status |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | quantity of remaining battery charge |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Anaesthetic machine should have power outlets to power ancillary devices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The anaesthetic machine trolley should have provisions to fit/install a UPS, to power ancillary devices. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.02 |  |  | Trolley
|  | a |  | Tendered workstation must be supplied with and mounted on a trolley |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The trolley must be equipped with:
|  |  |  | a drawer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | a shelf for mounting of ancillary equipment and accessories |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a bracket for mounting a computer/keyboard/monitor arm |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | a writing surface |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | a work surface for drug preparation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | integrated mounting rails for accessories |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | an integral cable and hose management solution |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | Patient Tubing Support Arm adaptable either to the Left or Right side |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | Humidifier holder adaptable to either the left or right side |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | provisions for mounting infusion/syringe pumps |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | for models released from 2024 onwards, provisions allowing for the mounting of up to three (3) computer monitor displays |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Trolley castors must be greater than 125mm to provide easy access across a lift threshold |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Trolley castors At least two castors must be fitted with brakes. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.03 |  |  | Gas Systems
|  | a |  | The anaesthetic machine must work with a wall gas supply from 280kPa–600kPa. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The Anaesthetic Workstation must be equipped with flexible supply hoses (min 3000 mm in length) for each gas fitted conforming to AS 2896 and utilising Sleeve Index System fittings |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The Anaesthetic Workstation must include connection to Medical Air, Medical Oxygen, Medical Nitrous Oxide and Medical Suction |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Anaesthetic Workstation must be equipped with Indicators for each gas fitted displaying supply pressure status |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The Anaesthetic Workstation must be equipped with a suction reservoir and suction system suitable for pharyngeal suction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The Anaesthetic Workstation must be equipped with provision for reserve Medical Oxygen supply bottle/s and (both mounted and connected) must use the Pin Indexed System (AS 2473.3). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | A reserve supply of oxygen must be attached to the anaesthetic machine in a manner that ensures easy activation should the oxygen supply failure warning system [refer to clause 5.37.13] indicate impending failure of the external oxygen supply, as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | The Anaesthetic Workstation should be equipped with provision for reserve Nitrous Oxide supply bottle/s and (both mounted and connected) must use the Pin Indexed System (AS 2473.3). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Reserve supply bottles (both mounted and connected) must use the Pin Indexed System (AS 2473.3). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | The gas supply must be arranged so that when there is a correctly working gas supply from the mains pipeline, this supply must be used in preference to any reserve gas supplies connected. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | The reserve gas supply cylinder contents, even with the cylinder turned on, must not be diminished by any intentional bleed within the system whether the Anaesthetic Workstation is turned on or off. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | When high-pressure gas supply systems are in use, an oxygen supply failure warning device must be present on the anaesthetic machine. This must: |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | activate automatically when the oxygen supply pressure falls below a predetermined critical level |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | generate an alarm to warn the operator |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | cut off the supply of gases other than air or oxygen to the common fresh gas outlet |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | cancel the alarm only when the oxygen supply pressure has been restored to a level above that at which the device was activated; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | If the anaesthetic machine incorporates a gas flowmeter bank, oxygen must be the last gas to enter the common gas manifold at the top of the flowmeter tubes. The oxygen flow knob must be the first from the left on the rotameter, relative to other flow knobs; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | If mechanical means are provided to mix the anaesthetic gases on the anaesthetic machine, there must be only one gas flow control knob for each gas; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | If a mechanical oxygen flow knob is provided, it must differ from the other flow control knobs so that tactile identification of the oxygen control knob is possible (for example, fluted); as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | p |  | If the anaesthetic machine is capable of delivering nitrous oxide, the machine must not deliver a hypoxic mixture. When oxygen and nitrous oxide are the only gases used, the machine must prevent delivery of a gas mixture with an oxygen concentration below that of ambient air; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | Anaesthetic gas scavenging system connections must be of a diameter that is different from the other connections used for the breathing system; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | r |  | The gas supply must be arranged so that unidirectional flow control devices prevent a contaminating flow from one gas supply into another either in normal operation or in the event of failure of any part of the system or any gas supply. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | s |  | Electronic gas flow controls must have a display that clearly indicates which gas is being adjusted and what the current flow level is for all gases. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | t |  | Direct mechanical flow controls must continuously increase gas flow when turned in an anti-clockwise direction and continuously decrease the gas flow when turned in a clockwise direction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | u |  | Flowmeters must display both a digital value and a bar-graph representation of flow for each gas. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | v |  | A separate flow indicator must be provided for each gas. The scaling must be in units of litres per minute |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | w |  | The flow measurement range for each gas must be 0 to 10 litres per minute |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | x |  | The flow measurement accuracy and resolution must be <1% |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | y |  | A high-priority alarm must be activated when high airway pressure is present; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.04 |  |  | Breathing System
|  | a |  | The Anaesthetic Workstation must be equipped with an emergency oxygen flush button with means to prevent accidental activation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The emergency oxygen flush control must be protected from accidental activation; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Oxygen flush rates should be capable of being greater than 35L/minute. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | One floating ball type oxygen flow regulated outlet delivering at least 10 L/min must be provided in a convenient location |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | A high-pressure relief valve or other means of automatically preventing dangerously high pressures in the breathing system must be present in the breathing circuit but it is not necessary for such a device to be integral to the anaesthetic machine; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Breathing systems must be easily dissembled to for cleaning purposes. Breathing systems must also be able to endure standard hospital cleaning procedures and materials. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | If a specific tool is required for the removal of the breathing system, the tool must be provided with the system. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | It must be quick and easy to change from mechanical ventilation to hand ventilation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | An Adjustable Pressure Limiting (APL) valve must be provided with a range of 0.5–70cm/H2O. Preference may be given to non-linear adjustable valves. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | A fresh gas outlet, if provided, must be 22mm outer diameter and 15mm inner diameter, visible to the operator and should be capable of being connected to the breathing system in such a way as to prevent accidental disconnection; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | A high-priority alarm must be activated when the airway pressure falls 10cm H2O below atmospheric pressure for more than one (1) second; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | The common gas outlet must have a patient safety relief valve. Specify the pressure at which the patient safety relief valve opens (and the nominal flow at which that specification is determined). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Continuous ventilation of a patient using a humidified patient circuit must be possible without damage to the ventilator system or affecting its accuracy, operation without adversely affecting ventilator performance, accuracy, operation or long term reliability. Please list compatible humidifiers in a separate cross referenced attachment. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | The delivery system must be fitted with either a refillable carbon dioxide absorbing canister having sieves and a clear polycarbonate container or a replaceable carbon dioxide absorbing canister. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | The absorber system must be autoclavable for the purpose of decontamination. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | p |  | Absorber must have Self-sealing mechanism to allow the exchange of canisters during the case without generating a leak in the system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | q |  | Provision (including conveniently located connectors) must be made to direct gases sampled from the patient circuit by the airway gas monitor either to: a) the scavenging circuit, or b) returned to the patient circuit via the ventilator, (user selectable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | r |  | An integrated mechanical airway/circuit manometer gauge is desirable near the circuit attachment points. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.05 |  |  | Ventilation
|  | a |  | Provision must be made for manual ventilation of the patient using a manual ventilation bag and manually Adjustable Pressure Limiting (APL) valve (hPa). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | In the VC mode, the delivered tidal volume must be independent of the rate of fresh gas flow (i.e. compensation provided to ensure delivered tidal volume equals selected tidal volume independent of changes in fresh gas flow). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | In the VC mode, the delivered tidal volume must have compensation provided for circuit compliance and leaks (as measured during system self test) to ensure delivered tidal volume equals selected tidal volume independent of circuit compliance and leaks). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | It must be possible to configure the default mechanical ventilation mode and settings, which are delivered at start-up. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Modes of ventilation must include pressure support |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Modes of ventilation must include pressure limited volume-guaranteed mode |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Ventilators with pressure mode must have a minimum volume range of 40–1400mL. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Ventilators must be able to ventilate in a range of respiration rates from 4–60BPM. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Ventilators must be able to provide Positive End Expiratory Pressure (PEEP) of 5–30cm/H2O. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Ventilators must be able to sustain a leak of <= 150mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | Ventilators must provide minimum gas flow of 120mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | A high airway pressure alarm must result in the cessation of an inspiratory breath and the opening of a safety valve to relieve the pressure. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The airway pressure relief must be in addition to, and at a level above, the controlled pressure level in the time cycled pressure-controlled mode of ventilation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | All current settings of the ventilator must be shown simultaneously and continuously during operation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | Ventilator settings and measured values must be clearly differentiated from other monitoring values and waveforms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | The ventilator must incorporate an expiratory tidal volume monitoring and alarms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | p |  | The Anaesthetic Workstation must provide a means to allow switching between manual and automatic ventilation without the disconnection of hoses |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.06 |  |  | Airway Gas Monitoring
|  | a |  | A comprehensive set of alarms for the airway gas monitoring must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A side stream sampling technique must be used to provide a source for an integrated airway gas monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The sampling rate must be between 50 and 200 mL/min. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Airway gas monitoring must commence automatically and alarms must be automatically enabled either when the machine is turned on or upon the detection of the first breath. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | A water trap must be provided to protect the measuring units. Please specify if single use or reusable water traps are available. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | A connector must be provided for exhaust sample gas to be returned to the anaesthesia circuit or directed to the scavenging system. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Each waveform must have a user adjustable scale. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | Verification of calibration by the user must be possible. Please describe the calibration method |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The Respondent must specify the method used for calibration. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | User initiated calibration at both 21% and 100% oxygen must be possible. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.07 |  |  | Display
|  | a |  | Operating buttons for soft-key controls must clearly align with on-screen labels without any parallax error. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | It must be possible to measure and display airway pressures, inspiratory volumes, expiratory volumes, pressure-volume loops, flow-volume loops, airway pressure waveforms, airway flow waveforms, compliance and airway resistance. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | A means of displaying ventilator flow and pressure waveforms to the patient must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All waveforms must be identified by an on-screen label. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The monitor must display, in a digital form, data derived from each waveform displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Digital values must automatically align with associated waveforms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | Both inspiratory and expiratory tidal volume (TV) and minute volume (MV) values must be displayed. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.08 |  |  | Alarms
|  | a |  | The Anaesthetic Workstation must automatically activate associated ventilator monitors and alarms when machine is turned ON |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A clear visual and audible alarm must be generated when there is:
|  |  | i | a sustained positive pressure in the breathing system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a high peak pressure in the breathing system |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | sustained sub atmospheric pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | loss of mains power |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | low Battery |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | loss of mains gas pressure |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | non-cycling |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | circuit disconnection |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | low/high expiratory volume |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | apnoea. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
• It is desirable for an indication of the cessation of apnoea. Please specify.
|  | c |  | Ventilator alarm limits must be displayed at all times (see below). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Ventilator alarms must be prioritised according to level of urgency. Priorities of alarms must change in both tone and visual interpretation |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The ventilator ALARM SILENCE button must not silence any alarm for a period exceeding two (2) minutes. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The ventilator ALARM SILENCE button must not silence high priority alarms for a period exceeding one (1) minute (e.g.. apnoea). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The ventilator ALARM SILENCE button must not remove the visual warning of an existing alarm condition. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.09 |  |  | Vaporisers
|  | a |  | If more than one delivery system can be simultaneously mounted on the anaesthetic machine, a interlock system must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Vaporisers with mechanical adjustment dials, when used in high-pressure circuits, must increase the delivered anaesthetic vapour concentration when the dial is rotated in an anti-clockwise direction; as per ANZCA PS54. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The minimum type of anaesthetic vaporiser that Respondents must offer is isoflurane, sevoflurane and desflurane. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The Anaesthetic Workstation must deliver volatile anaesthetic agents with an accurate mechanism |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The Anaesthetic Workstation must have safe guards - to prevent simultaneous dual vapour administration and this must not be able to be bypassed |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Anaesthetic Workstation equipped with vaporisers with mechanical rotary dials must increase the delivered vapour concentrations when the dial rotated in anti-clockwise direction and must have a detent for the off position. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | The Anaesthetic Workstation must compensate for low gas flows and variations in room temperature to allow precise delivery of volatile anaesthetics at minimal gas, fresh gas flow rates |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | The Anaesthetic Workstation must display the volatile agent - currently in use |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | The Anaesthetic Workstation must facilitate easy and quick change of agent during operation. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | The vaporiser must be easily removed for filling and maintenance. A secure latching mechanism must be provided. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | The vaporiser must have an agent specific filling system to prevent inadvertent use of the incorrect agent. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | l |  | The vaporiser must have a liquid level indicator clearly and indelibly marked with maximum and minimum filling levels. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | m |  | Where mechanical vapour control is offered, each workstation must provide for the mounting of two anaesthetic vaporisers for ease of transfer between volatile agents. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | n |  | A safety mechanism to prevent accidental operation of the control dial must be provided |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | o |  | The vaporiser must be clearly and indelibly marked with the generic name of the anaesthetic agent and colour coded according with the appropriate Australian Standards. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 5.1 |  |  | Patient Monitor Integration
|  | a |  | The Anaesthetic Workstation must accommodate a patient monitor and integrate seamlessly. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Please list all the compatible patient monitors
|  | b |  | The Anaesthetic Workstation must support the weight of the patient monitor. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Please specify the max weight
| 6 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device, and | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | contain a complete set of technical specifications which describe the inbuilt spirometry, ventilator monitoring hardware and software performance. Be sure to include measurement range, measurement accuracy and alarm capabilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Details on flushing requirement to remove traces of volatile agent in preparation for Malignant Hyperthermia patients. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 10.03 - Anaesthesia workstation, MRI
| 7 |  |  | General Requirements
|  | a |  | Tendered device must be MR safe or MR Conditional |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | MRI workstation must meet all requirements as per section:
|  |  | i | 11.01 Anaesthesia Workstation, Low-Acuity, or |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | 11.02 Anaesthesia Workstation, General-Purpose |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondent must offer any and all required MR safe accessory and/or consumable required for device to function as intended in a MRI room |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All patient leads/cables/electrodes must be almost MR safe |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | A remote slave display should be offered for use/monitor/control in the Control Room |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 8 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
Subcategory 10.04 - Anaesthesia workstation, Analytics Applications
| 9 |  |  | Standards and Compliance
|  | a |  | Tendered System must only collect, store, and access the minimum amount of patient's personal data, necessary to perform its function. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | List the patient's personal data to be collected, stored and accessed
|  |  | ii | Stored patient's personal data must be encrypted |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | If Tendered Systems are cloud based, then the location of the cloud-based servers must be within the boarders of Australia. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Participating Health Services may prefer that the cloud based servers are located within the State of Victoria. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Respondents must maintain and ensure its permitted third-parties, maintain the physical and virtual security of any facilities owned, managed, licensed, or controlled by Respondent that stores patient data by implementing industry best security practices at the locations where patient data is stored to ensure the confidentiality, integrity, and availability of the patient data and the Software Management System that stores the patient data. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Respondent, where applicable, must maintain backup policies and procedures of Data containing patient's personal data, image repositories, and provisioned environments. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The backup storage infrastructure must be located in physically protected, limited-access facilities. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Prior to disposal of any Device returned or discontinued by or on behalf of the Participating Health Service, Respondents must securely wipe or destroy all patient data consistent with industry standards, such as NIST 800-88, to ensure that no patient data is retrievable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | Upon request of the Participating Health Service, Respondent must provide confirmation of patient's personal data wipe or destruction. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 10 |  |  | General Requirements
|  | a |  | Respondent must submit as part of their tender response, a complete product brochure in portable document format (i.e., .pdf) for tendered product. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11 |  |  | ICT
|  | a |  | Access to data repository must only be by system authorised users. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11.1 |  |  | Data Recording
|  | a |  | Web-enabled formats are preferred. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11.2 |  |  | Data Transmission and Interfacing
|  | a |  | Data transmission must ensure patient privacy. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Clinical systems are being integrated to other patient information systems including HIS, EMR etc. Systems must be using the latest version of the standard interfacing and messaging standard language, Health Level Seven (HL7). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | System must allow bidirectional data flow |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 11.3 |  |  | Data Hardware and Software
|  | a |  | Authorised users must be able to upgrade software, troubleshoot and configure the system settings. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Printouts must be of an appropriate diagnostic quality. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The ability to print to non-thermal paper is desirable. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | All connecting devices must have standard TCP/IP connections. Cabling installation and management may become a requirement of the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | Permissions and Participating Health Services’ protocols must be adhered to. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | Vendor supplied networking equipment and servers should be located in the Participating Health Services’ ICT areas, such as communications rooms and server farms. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | g |  | For Vendor supplied networks, all connecting cables must be provided and installed by the successful Respondent(s) (including fibre optics, where applicable). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | h |  | For Hospital supplied networks, all patch cables must be provided by the successful Respondent(s). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | i |  | Patch cables must be green, unless specifically requested to be another colour by a Participating Health Service. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | j |  | Upon purchase of any ECG Management system, successful Respondents must provide to Participating Health Services:
|  |  | i | technical diagrams of the installation, including a network diagram |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | a technical address spreadsheet with all IP, subnet masks and MAC addresses (where appropriate locations should be indicated). |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | k |  | System must be scalable to increasing workload by means of adding additional servers which operate in parallel whilst providing business continuity without having to replace the server. |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 12 |  |  | Product Information
|  | a |  | The Respondent must submit a full and complete Instruction for Use manual, in portable document format (i.e., .pdf), for tendered device. | --Select-- | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For each tendered device in this subcategory, the Respondent must provide the following information in the Product Catalogue:
|  |  | i | Manufacturer |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of origin |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • country of final assembly |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | Warranty period |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | Year 1st manufactured |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | Estimated equipment lifecycle (use lifespan in years) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | Estimated device's end-of-sales life (state the year) |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | list any known incompatible devices |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 11-Upgrades
Category 11: Physiological Monitoring Software Upgrades
Clause Identifiers |  |  |  | Text
Cat 11.0 |  |  |  | Physiological Monitoring Software Upgrades | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Options are defined as after sales upgrades, i.e. options are additional features that can be added to or included onto tendered equipment that were not included in the original purchase | --Select--
|  |  | i | Includes software upgrades that can be added/installed on to tendered equipment that has become available since the original equipment purchase. | --Select--
| 2 |  |  | General Requirements
|  | a |  | Respondents must list all their after sales enhancements that can be added/installed on tendered equipment, including pricing, once a Health Service has purchased tendered equipment. | --Select--
| 3 |  |  | Product Information
|  | a |  | For each tendered upgrade, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | Compatible with which model tendered device | --Select--

## 12-Consumables
Category12:Physiological Monitoring Associated Consumables
Clause Identifiers |  |  |  | Text
Cat 12.0 |  |  |  | Physiological Monitoring Associated Consumables | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1.0 |  |  | Intended Purpose
|  | a |  | Associated consumable is a consumable that is required for the use with the tendered equipment, for the equipment to function as intended by the manufacturer. | --Select--
|  | b |  | It can also be a consumable that may improve the functionality of tendered equipment. | --Select--
|  |  | i | The tendered consumable does not necessarily require to be an OEM consumable. | --Select--
| 2.0 |  |  | General Requirements
|  | a |  | Tendered consumable must not be listed on any other HSV agreement | --Select--
|  | b |  | Tendered consumable must be associated (used) with only the equipment listed under this agreement. | --Select--
|  | c |  | Respondent must always have stock-on-hand of tendered consumable for the life of this agreement or until the tendered equipment's end-of-support date, or have organised a substitute. | --Select--
| 3 |  |  | Product Information
|  | a |  | For each tendered consumable, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | Compatible with which model tendered device | --Select--

## 13-SemiConsumables & Batteries
Category 13: Physiological Monitoring Semi-Consumables
Clause Identifiers |  |  |  | Text
Cat 13.0 |  |  |  | Physiological Monitoring Semi-Consumables | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1.0 |  |  | Intended Purpose
|  | a |  | A semi-consumable is a part that is used multiple times but is not expected to last the entire effective life of the equipment, such as batteries, device filters, etc. | --Select--
|  |  | i | is an item required by the tendered device to function as intended, and | --Select--
|  |  | ii | must be replaceable by user without the need of a tool. | --Select--
|  |  |  | • if a tool required, then item is considered as a spare part | --Select--
|  |  |  | - excluding batteries | --Select--
|  | b |  | An accessory is a supplementary part that enhances but does not inhibit the operations of the device, e.g. carry bag, barcode reader, etc. Device would still function as intended without the accessory. | --Select--
| 2.0 |  |  | General Requirements
|  | a |  | Respondents must list all semi-consumables and batteries that are available for use with the tendered equipment, including pricing in the Tender Response Worksheet. | --Select--
|  | b |  | Respondent must at all times have stock-on-hand of batteries for each tendered equipment for the life of this agreement or until the tendered equipment's end-of-support date. | --Select--
| 3 |  |  | Product Information
|  | a |  | For each tendered battery, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | Compatible with which model tendered device | --Select--
|  | b |  | For each tendered semi-consumable, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | Compatible with which model tendered device | --Select--

## 14 - Monitoring Products
Category 14: Monitoring Products for Physiological Monitoring
One range of Monitoring Products per column
Clause Identifiers |  |  |  | Text | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here>
Cat 14.0 |  |  |  | Monitoring Products for Physiological Monitoring | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1.0 |  |  | Intended Purpose
|  | a |  | Monitoring Products are defined as any product, accessory, or consumable required for use with the tendered equipment to undertake physiological monitoring (e.g., pulse oximetry, NIBP, Temp, etc.) of the patient. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.0 |  |  | General Requirements
|  | a |  | All tendered products within this category must work with Respondent's tendered equipment | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All and any product, accessory, or consumable, required by the tendered equipment to undertake the physiological monitoring of the patient must be listed on either: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | HSPVC2023-021 Monitoring Products, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | HPVITS2024-075 Physiological Monitoring and Anaesthetic Gas Delivery Systems, only if the product has not been awarded on HPVC2023-021 Monitoring Products. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.1 |  |  | Electrocardiograph Cables and Lead wires
|  | a |  | A range of electrocardiograph monitoring cables and compatible lead wire sets should be tendered to meet the clinical needs of ECG monitoring, if required on tendered device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Leads must be constructed in such a manner that the lead wire connector cannot make contact with electrical ground or induce hazardous potentials. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | The connector shall be constructed to prevent conductive contact with a mains outlet or a detachable power cord. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | Cable yokes and lead wires or lead wire connectors must be permanently colour coded or labelled with lead positions to minimise the risk of misconnection. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | For each tendered electrocardiograph cable and/or lead wire set, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | list the brand model defibrillator tender item works with | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | tendered item is a | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • ECG trunk cable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • ECG lead wires | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • ECG cable with nondetachable lead wires | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | number of leads on lead wire set | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | length of the cable or lead wires, in millimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | lead wire colour code option (e.g., AAMI or IEC colour coded) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | electrode connection options e.g., snap/press, grabber/scissor, alligator, banana | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | Single Patient Use or reusable leads | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | radio-translucency | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | MRI compatibility | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | Electro Surgical Unit (ESU) filtration capability | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | cleaning recommendations of reusable and single patient use items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | disinfection recommendations for reusable items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.2 |  |  | Non-Invasive Blood Pressure Cuffs/Hoses/Connectors
|  | a |  | A range of blood pressure cuffs and compatible components/parts should be tendered to meet the clinical needs for non-invasive blood pressure monitoring, if required on tendered device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | Tendered blood pressure cuffs, tubing and components should be latex free | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For each tendered cuff and/or hose, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | tendered item is a | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • NIBP hose | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • Cuff | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • bladderless cuff | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | size e.g., neonate, adult, bariatric | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | physical size of the cuff, width by length in centimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | length of tubing on the cuff, in centimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | hose length, in centimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | cleaning recommendations of reusable and single patient use items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | disinfection recommendations for reusable items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.3 |  |  | Recorder Paper
|  | a |  | Recorder paper exclusive to the tendered equipment must be tendered. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | For all recorder paper offered, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | format e.g., roll, sheet, z-fold pack etc | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | physical size of the paper/sheet of paper, in centimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | pack or roll dimensions e.g., length, width, roll length, inner core diameter in millimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | Information relating to any specific paper storage requirements necessary to maintain the integrity of a monitoring trace in the permanent patient record shall be readily available for users as printed and/or electronic information, which must include: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • the minimum storage time without significant degradation of recordings, providing manufacturer’s storage recommendations are followed | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • whether recordings are affected by optical scanning or photocopying | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • storage life of unused packs | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.4 |  |  | Haemodynamic Monitoring
|  | a |  | A range of haemodynamic monitoring cables should be tendered to meet the clinical needs for invasive blood pressure monitoring, if required on tendered device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | cleaning recommendations for cable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | disinfection recommendations for cable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.5 |  |  | Temperature Monitoring
|  | a |  | A range of Temperature monitoring cables, for use with single-use temperature probes, should be tendered to meet the clinical needs for continuous temperature monitoring, if required on tendered device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | cleaning recommendations for cable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | disinfection recommendations for cable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.6 |  |  | Pulse Oximetry
|  | a |  | Pulse oximeter probes must only be offered if they are made by the same manufacturer as tendered device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A range of pulse oximetry extension cables should be tendered to meet the clinical needs for SpO2 monitoring, if required on tendered device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For each tendered electrocardiograph cable and/or lead wire set, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | tendered item is a | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • SpO2 extension cable | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • reusable pulse oximeter probe | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • single patient use pulse oximeter probe | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | brand name | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | age range suitability i.e., adult, paediatric or neonate | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | recommended patient weight range e.g., <1kg, 10-30kg | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | sensor length, in millimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | the method of attachment e.g., clip on | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | cable length, in centimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | cleaning recommendations of reusable and single patient use items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | disinfection recommendations for reusable items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2.7 |  |  | Sensory Awareness
|  | a |  | Sensory awareness monitoring sensors must only be offered if they are made by the same manufacturer as tendered device. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | A range of sensory awareness monitoring sensors should be tendered to meet clinical needs for depth-of-anaesthesia monitoring, if required on tendered device. This includes: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | consumable (single use) and reusable sensors | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | for unilateral and bilateral placement | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | to suit adult, paediatric, and neonatal patients. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | For each sensory awareness monitoring sensor offered, Tenderers shall advise the following information in the Product Description on the Tender Response Worksheet:
|  |  | i | brand name | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | age range suitability i.e., adult, paediatric | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | sensors type i.e., unilateral, bilateral | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | brand/model of compatible monitor(s) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | MRI conditional | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | mechanical construction (e.g., exploded view showing components, construction, and materials) | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | any application requirements or recommendations related to patient age, weight, or clinical condition e.g., neonatal weight | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | skin preparation requirements, including additional products required e.g., skin cleansers | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iv | removal techniques and requirements | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | v | information regarding allergenicity, toxicology and skin testing | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vi | reactivity with commonly used skin care and skin preparation products | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | vii | effects of temperature, humidity and perspiration on performance and adhesion | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | viii | recommended wear time | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ix | any contraindications for use | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | x | shelf life of: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • unopened packs | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • opened packs | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xi | any storage requirements, e.g., temperature, humidity, light, disposal | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | xii | product cross reference table | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • equipment and sensory awareness monitoring sensor compatibility | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 15-Respiratory Products
Category 15: Respiratory Products for Physiological Monitoring
One range of Respiratory Products per column
Clause Identifiers |  |  |  | Text | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here> |  | Ordering Part Number: <enter here>
Cat 15.0 |  |  |  | Respiratory Products for Physiological Monitoring | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Respiratory Products are defined as any product, accessory, or consumable, required for use with the tendered equipment to undertake respiratory monitoring (e.g., EtCO2) of the patient. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
| 2 |  |  | General Requirements
|  | a |  | All tendered products within this category must work with Respondent's tendered equipment | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | b |  | All and any product, accessory, or consumable, required by the tendered equipment to undertake the respiratory monitoring of the patient must be listed on either: | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | i | HPVC2021-015 Respiratory Products, or | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | HPVITS2024-075 Physiological Monitoring and Anaesthetic Gas Delivery Systems, only if the product has not been awarded on HPVC2021-015 Respiratory Products. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | c |  | A range of gas sampling lines and/or airway adapters should be tendered to meet the clinical needs for end-tidal carbon dioxide monitoring, if available on the defibrillator. | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | d |  | Tendered product must be either clinically clean or sterile | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | e |  | For each tendered gas sampling line cable and/or airway adapter, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | tendered item is a | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • gas sampling line | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • reusable airway adapter | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  |  | • single patient use airway adapter | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | age range suitability i.e., adult, paediatric or neonate | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | iii | length of the gas sampling line, in centimetres | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  | f |  | The following information must be made available upon the request from Participating Health Services as printed and/or electronic information:
|  |  | i | cleaning recommendations of reusable and single patient use items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--
|  |  | ii | disinfection recommendations for reusable items | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select-- |  | --Select--

## 16-Accessories
Category 16: Physiological Monitoring Accessories
Clause Identifiers |  |  |  | Text
Cat 16.0 |  |  |  | Physiological Monitoring Accessories | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | An accessory is a supplementary part that does not alter the operations of the device but rather improves functionality, e.g. carry bag, barcode reader, etc. The device would still function as intended without the accessory. | --Select--
| 2 |  |  | General Requirements
|  | a |  | Respondents must list all available accessories for use with tendered equipment, including pricing within the Tender Response Worksheet. | --Select--
| 3 |  |  | Product Information
|  | a |  | For each tendered accessory, Respondents must provide the following information in the Product Description of the Tender Response Worksheet:
|  |  | i | Compatible with which model tendered device | --Select--

## 17-Spares
Category 17: Physiological Monitoring Equipment Spare Parts
Clause Identifiers |  |  |  | Text
Cat 17.0 |  |  |  | Physiological Monitoring Equipment Spare Parts | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Spare part is any part that is required for any preventative or corrective maintenance procedure in maintaining a tendered equipment. | --Select--
| 2 |  |  | Spare parts
|  | a |  | Respondents must list all OEM spare parts and assemblies in Tender Response Worksheet that a Health Service will require to maintain tendered machine fully functional for its entire lifecycle. | --Select--
|  | b |  | Maintenance kit is a kit of spare parts required for routine periodic maintenance | --Select--
|  |  | i | Respondents must list all maintenance kits in Tender Response Worksheet that are required to maintain tendered machine over machine's lifespan. | --Select--
|  | c |  | Respondents must list in Tender Response Worksheet all the proprietary testers, calibration kits, specialised tools, cables, dongles, etc., that are only available from the OEM, required to fully repair and maintain tendered equipment. | --Select--
|  |  | i | Generic parts, such as, mains power leads, screws, washers, fuses, etc. must not be tendered. | --Select--
|  |  |  | • Any part that is equivalent to an item readily available from RS Components and the such, must not be tendered. | --Select--
|  | d |  | For each spare part/assembly listed in Tender Response Worksheet, the Respondents must name the model equipment it is required for. | --Select--
|  |  | i | Respondents may also include spare parts for non-tendered equipment | --Select--
|  | e |  | Successful Respondent upon being requested must supply all spare part/s and/or maintenance kit/s required by Participating Health Service to maintain their tendered equipment. | --Select--

## 18-S&M Agreements
Category 18: Physiological Monitoring Service & Maintenance (S&M) Agreements
Clause Identifiers |  |  |  | Text
Cat 18.0 |  |  |  | Physiological Monitoring Service & Maintenance (S&M) Agreements | Fully comply Partially comply Does not comply Not Applicable | Please specify
| 1 |  |  | Intended Purpose
|  | a |  | Respondents must submit any and all agreements, which lists any tendered product under this HSV agreement, which Successful Respondent has entered into with any of the Participating Health Services; for the purposes of transparency for all Participating Health Services. | --Select--
|  | b |  | Services that may be offered for products tendered under this HSV agreement: | --Select--
|  | c |  | Maintenance Agreements
|  |  | i | A contract between Successful Respondent and Participating Health Service for respondent to maintain Participating Health Service’s equipment. Maintenance agreements may range anywhere from full comprehensive, to partnership maintenance arrangements with in-house Biomedical Engineering Department, to parts-only arrangements. | --Select--
|  |  | ii | Respondents must populate the Product Catalogue for all variants of their maintenance agreements. | --Select--
|  |  | iii | Respondents must submit all Service and Maintenance sales data on the template provided as part of this ITS. | --Select--
|  |  |  | • must list each part used, and its cost, in the repair process | --Select--
|  |  |  | • must list total hours taken to conduct the repair | --Select--
|  |  | iv | Copies of all new maintenance agreements for Health services executed after start of this contract that include products this tender, must be provide to HSV. | --Select--
|  |  |  | • on an annual basis (from date of contract commencement), or | --Select--
|  |  |  | • upon HSV request. | --Select--
|  |  | v | Respondents must list the exclusions in the Product Catalogue | --Select--
|  | d |  | Corrective Maintenance
|  |  | i | A fee paying service for single items where respondent identifies, isolates, and rectifies faults so that failed equipment is restored to operational condition within the manufacturer’s established tolerances. | --Select--
|  |  | ii | Respondents must populate the Product Catalogue for corrective maintenance. | --Select--
|  |  | iii | Respondents must submit all corrective maintenance sales data on the template provided as part of this ITS. | --Select--
|  |  |  | • must list each part used, and its cost, in the repair process | --Select--
|  |  |  | • must list total hours taken to conduct the repair | --Select--
|  |  | iv | Any pay-as-you-go, routine preventative maintenance services, must also be included in this subcategory | --Select--
|  | e |  | Calibration and Testing Agreements
|  |  | i | A contract between Successful Respondent and Participating Health Service for the regular assessment, calibration and testing of the medical equipment to ensure it is safe, and continues to be safe, for its intended clinical application. | --Select--
|  |  | ii | Respondents must populate the Product Catalogue for Calibration and Testing Agreements. | --Select--
|  |  | iii | Respondents must submit all Calibration and Testing sales data on the template provided as part of this ITS. | --Select--
|  |  |  | • must list all parts used, and their costs | --Select--
|  |  |  | • must list total hours taken to conduct the calibration and/or testing | --Select--
|  |  | iv | Copies of all calibration and testing agreements for Participating Health services that include any product that has been tendered within this HSV agreement, must be provided to HSV: | --Select--
|  |  |  | • on an annual basis (from date of contract commencement), or | --Select--
|  |  |  | • upon HSV request. | --Select--
|  |  | v | Respondents must list the exclusions in the Product Catalogue | --Select--
|  | f |  | Extended Warranties
|  |  | i | Respondents may offer optional extended warranties to lengthen the coverage of the manufacturer's warranty on their tendered products. | --Select--
|  |  | ii | Respondents must populate the Product Catalogue for any Extended Warranties offered. | --Select--
|  |  | iii | Respondents must submit all Extended Warranty sales data on the template provided as part of this ITS. | --Select--
|  |  | iv | Respondents must list the Extended Warranty exclusions on the Product Catalogue | --Select--
|  | g |  | Technician's / Biomedical Engineer's (BME) Training
|  |  | i | Any Technician/Biomedical Engineer service training package offered to Health Service by the Successful Respondent, including the initial free-of-charge servicing training and of any on-going annual training. | --Select--
|  |  | ii | Respondents must submit all Technician/Biomedical Engineer Training sales data on the template provided as part of this ITS, including the initial free-of-charge training. | --Select--
|  | h |  | Software and Hardware Maintenance Agreements | --Select--
|  |  | i | Software maintenance agreement include the provisions for software upgrades including code updates, version updates, code fix modifications, enhancements, and added or new functional capabilities loaded into existing software. | --Select--
|  |  |  | • Respondents must populate the Product Catalogue for Software Maintenance agreements. | --Select--
|  |  |  | • Respondents must submit all Software Maintenance sales data on the template provided as part of this ITS. | --Select--
|  |  | ii | Hardware maintenance agreement extends, replaces EOL computer hardware equipment during the useful life of the tendered medical device. | --Select--
|  |  |  | • Respondents must populate the Product Catalogue for Hardware Maintenance agreements. | --Select--
|  |  |  | • Respondents must submit all Hardware Maintenance sales data on the template provided as part of this ITS. | --Select--
|  |  | iii | Copies of all calibration and testing agreements for Participating Health services that include any product that has been tendered within this HSV agreement, must be provided to HSV: | --Select--
|  |  |  | • on an annual basis (from date of contract commencement), or | --Select--
|  |  |  | • upon HSV request. | --Select--
|  |  | iv | Respondents must list the exclusions in the Product Catalogue | --Select--
|  | i |  | Agreements
|  |  | i | A contract between Successful Respondent and Health Service where the Respondent provides the Health Service with 'equipment' at 'free-on-loan' as arrangement - where the capital costs are recouped through the consumables for a fixed length of time. | --Select--
|  |  | ii | During the term of the agreement, the Successful Respondents will be responsible for maintaining the ‘equipment’ in full working order, unless mutually agreed with the Health Service for another party to maintain equipment. | --Select--
|  |  |  | • Preference may be given to successful Respondents who provide service training to Health Service Technical Services team to maintain the equipment. | --Select--
|  |  | iii | At the end of the agreement Health Services must have the option to purchase the equipment at a minimal cost. | --Select--
|  |  | iv | Respondents must list any exclusions to agreement in the Product Catalogue. | --Select--
|  |  | v | Copies of all agreements for Health Services that include any product that has been tendered within this HSV agreement, must be provide to HSV: | --Select--
|  |  |  | • on an annual basis (from date of contract commencement), or | --Select--
|  |  |  | • upon HSV request. | --Select--
|  |  | vi | All products and services, covered within the scope of this contract, must be a contracted item on HSV contract(s) | --Select--
|  |  | vii | All products and services must be priced respectively (items wise), including any Leasing fee, if applicable | --Select--
|  |  | viii | Successful Respondents must not charge more than the HSV contracted price on any items on the agreements. | --Select--
| 2 |  |  | General Requirements
|  | a |  | Respondents must provide a full range of service agreement options for each item of equipment tendered. This may range from a full comprehensive service agreement (24 hours a day, 7 days a week) to a routine testing only agreement, or anywhere in between. Health Services reserve the right to enter into these contracts during the entire, or part of, the effective life of the equipment. | --Select--
|  |  | i | Respondents must provide unit pricing in the Product Catalogue for each service and maintenance agreement option | --Select--
|  |  | ii | Successful Respondents must apply volume discounts on all service and maintenance agreements. | --Select--
|  | b |  | Upon request by HSV, successful Respondents must submit any and all agreements they have entered into with any of the Health Services that pertain tendered products as listed in this HSV agreement. | --Select--
|  | c |  | Successful Respondents must submit Service and Maintenance sales data on the template provided as part of this ITS. | --Select--

## Appendix 1-Product List
Appendix 1 - Product List
Category |  | Subcategory
1 | Vital Signs Monitoring | 1.01 | Pulse Oximeters, Handheld
|  | 1.02 | Pulse Oximeters, Desktop
|  | 1.03 | Aneroid Sphygmomanometers, hand-held
|  | 1.04 | Aneroid Sphygmomanometers, large dial
|  | 1.05 | NIBP monitors, portable
|  | 1.06 | NIBP monitors
|  | 1.07 | End-tidal CO2 Monitors
|  | 1.08 | Vital Signs Monitors
|  | 1.09 | Wearable Devices for Continuous Monitoring
2 | Acute Care Monitoring | 2.01 | Vascular Dopplers, Blood-flow Detectors
|  | 2.02 | Bladder Scanners
|  | 2.03 | Mid-Acuity Monitors
|  | 2.04 | High-Acuity Monitors
|  | 2.05 | Patient Monitoring System Modules
|  | 2.06 | MRI Monitors
|  | 2.07 | Telemetry Units
|  | 2.08 | Central Station Monitors
|  | 2.09 | Patient Monitoring Information Management Systems
|  | 2.1 | Patient Monitoring Asset Condition Management Software
|  | 2.11 | Transcutaneous Partial Pressure Monitors
|  | 2.12 | Cardiac Output Monitors
|  | 2.13 | Regional Oximetry Monitors
|  | 2.14 | Respiratory Gas Monitors
|  | 2.15 | Depth-of-Anaesthesia Monitors
|  | 2.16 | Patient Monitoring Network
|  | 2.17 | Patient Monitoring Network, Telemetry
3 | ECG Machines | 3.01 | ECG Machines
|  | 3.02 | ECG Machines, Handheld
|  | 3.03 | PC based ECG machines
|  | 3.04 | ECG Machines, Cardiac Stress Test
|  | 3.05 | Cardiac Stress Testing Systems
|  | 3.06 | ECG Management Systems
4 | Ambulatory Monitoring | 4.01 | Holter Monitoring System
|  | 4.02 | Holter Monitor
|  | 4.03 | Ambulatory Blood Pressure Monitoring Systems
|  | 4.04 | Ambulatory Blood Pressure Monitors
5 | Intrapartum Monitoring | 5.01 | FHR Dopplers
|  | 5.02 | CTG Machines
|  | 5.03 | Wireless Transducer Patch Systems
|  | 5.04 | Obstetric Management Systems
6 | Neurological, Sleep, and Sleep Related Monitoring | 6.01 | Cerebral Function Monitors [Amplitude-integrated Electroencephalography (aEEG)]
|  | 6.02 | Electroencephalography (EEG) Units
|  | 6.03 | Polysomnography (PSG) Units
|  | 6.04 | Electromyography (EMG) / Evoked Potential (EP) Systems - Nerve Conduction Study (NCS)
|  | 6.05 | Clinical Neurology Management Systems
7 | Basic Gas Delivery Devices | 7.01 | Medical Gas Regulators - Oxygen
|  | 7.02 | Medical Gas Regulators - Medical Air
|  | 7.03 | Medical Gas Regulators - Other Medical Gases
|  | 7.04 | Regulator/Dial Flow Meter - Oxygen
|  | 7.05 | Regulator/Dial Flow Meter - Other Medical Gases
|  | 7.06 | Flow meters (flow tube & ball) - Oxygen
|  | 7.07 | Flow meters (flow tube & ball) - Oxygen, low-flow
|  | 7.08 | Flow meters (flow tube & ball) - Oxygen, high-flow
|  | 7.09 | Flow meters (flow tube & ball) - Medical Air
|  | 7.1 | Flow meters (flow tube & ball) - Other Medical Gases
|  | 7.11 | Flow Meters (Gauge type) – Medical Gases
|  | 7.12 | Flow Meters (dial, preset flow rates) - Oxygen
|  | 7.13 | Flow Meters (dial, preset flow rates) - Medical Air
|  | 7.14 | Suction Regulator - without Gauge
|  | 7.15 | Suction Regulator - with Gauge
|  | 7.16 | Suction Regulator - with Gauge & Overflow Safety Trap
|  | 7.17 | Adaptors, 3 way - Oxygen
|  | 7.18 | Adaptors, 3 way - Medical Air
|  | 7.19 | Adaptors, 3 way - Suction
|  | 7.2 | Adaptors, 3 way - Other Medical Gases
|  | 7.21 | Medical Gas Hoses - Oxygen
|  | 7.22 | Medical Gas Hoses - Medical Air
|  | 7.23 | Medical Gas Hoses - Nitrous Oxide
|  | 7.24 | Medical Gas Hoses - Suction
|  | 7.25 | Medical Gas Hoses - Other Medical Gases
8 | Portable Suction Units | 8.01 | Portable Suction Devices, Gas-Driven (Twin-o-Vac®)
|  | 8.02 | Manual Suction Units
|  | 8.03 | Aspirators, Airway, mains powered
|  | 8.04 | Aspirators, Airway, battery powered
9 | Nitrous Oxide and Oxygen Delivery Systems | 9.01 | Analgesia Units
|  | 9.02 | 50/50 Oxygen/Nitrous Oxide Demand Valves
10 | Anaesthesia workstations | 10.01 | Anaesthesia workstation, Low Acuity
|  | 10.02 | Anaesthesia workstation, General-Purpose
|  | 10.03 | Anaesthesia workstation, MRI
|  | 10.04 | Anaesthesia workstation, Analytics Applications
11 | Physiological Monitoring Software Upgrades | 11.01 | Vital Signs Monitors
|  | 11.02 | Wearable Devices for Continuous Monitoring
|  | 11.03 | Mid-Acuity Monitors, Software
|  | 11.04 | High-Acuity Monitors, Software
|  | 11.05 | MRI Monitors
|  | 11.06 | Central Station Monitors, Software
|  | 11.07 | Patient Monitoring Information Management Systems
|  | 11.08 | Transcutaneous Partial Pressure Monitors
|  | 11.09 | Cardiac Output Monitors
|  | 11.1 | Regional Oximetry Monitors
|  | 11.11 | Respiratory Gas Monitors
|  | 11.12 | Depth-of-Anaesthesia Monitors
|  | 11.13 | ECG Machines
|  | 11.14 | Cardiac Stress Testing Systems
|  | 11.15 | CTG Machines
|  | 11.16 | Electroencephalography (EEG) Units
|  | 11.17 | Polysomnography (PSG) Units
|  | 11.18 | Electromyography (EMG) / Evoked Potential (EP) Systems - Nerve Conduction Study (NCS)
|  | 11.19 | Anaesthesia Workstations
|  | 11.2 | Other Equipment/Application Upgrades for Physiological Monitoring
|  | 11.21 | Patient Monitoring System Modules, Software
|  | 11.22 | Mid-Acuity Monitors, Hardware
|  | 11.23 | High-Acuity Monitors, Hardware
|  | 11.24 | Patient Monitoring System Modules, Hardware
|  | 11.25 | Telemetry Units
|  | 11.26 | Central Station Monitors, Hardware
12 | Physiological Monitoring Associated Consumables | 12.01 | Holter/Telemetry Monitor Bags, disposable
|  | 12.02 | Transcutaneous Gas Monitoring Consumables
|  | 12.03 | Wearable wireless physiological monitoring sensors (excluding ECG and CTG machines)
|  | 12.04 | ECG Machine Spirometry Flow sensors
|  | 12.05 | Foetal Scalp Electrodes
|  | 12.06 | Foetal Reference Electrodes
|  | 12.07 | CTG Machine Consumables, Other
|  | 12.08 | CTG Machine wearable wireless monitoring sensors
|  | 12.09 | Regional Oximetry Consumables
|  | 12.1 | Cardiac Output Monitoring Consumables
|  | 12.11 | Wearable transducers, 12-lead ECG
|  | 12.12 | Analgesia Units
|  | 12.13 | Consumable, Other, for Physiological Monitoring
|  | 12.14 | Recorder Paper
|  | 12.15 | Holter ECG Electrodes
|  | 12.16 | Electrocardiograph Electrodes, MRI Monitoring
13 | Physiological Monitoring Semi-Consumables | 13.01 | Batteries, Telemetry Units
|  | 13.02 | Batteries, Monitors
|  | 13.03 | Batteries, Other
|  | 13.04 | ECG cables, Telemetry
|  | 13.05 | End-tidal CO2, mainstream sensors
|  | 13.06 | Transcutaneous Gas Monitoring Semi-Consumables
|  | 13.07 | CTG Monitors, Patient interface / adaptor cables, reusable
|  | 13.08 | CTG Monitors, transducer, cabled
|  | 13.09 | CTG Monitors, transducer, wireless
|  | 13.1 | Anaesthetic Machine, CO2 Absorber
|  | 13.11 | Anaesthetic Machine, Oxygen fuel cell sensors
|  | 13.12 | Anaesthetic Machine, flow sensors
|  | 13.13 | Semi-Consumable, Other, for Physiological Monitoring
|  | 13.14 | ECG cables, Holter
14 | Monitoring Products for Physiological Monitoring | 14.01 | Electrocardiograph Monitoring Cables and Lead wires, Bedside Monitors
|  | 14.02 | Electrocardiograph Monitoring Cables and Lead wires, Telemetry Equipment
|  | 14.03 | Electrocardiograph Monitoring Cables and Lead wires, Diagnostic ECG Machine
|  | 14.04 | Other ECG semi consumables, e.g. Alligator clips or tab electrodes
|  | 14.05 | Non-Invasive Blood Pressure Cuffs, single use
|  | 14.06 | Non-Invasive Blood Pressure Cuffs, reusable
|  | 14.07 | Non-Invasive Blood Pressure Hoses
|  | 14.08 | Non-Invasive Blood Pressure Components/Connectors
|  | 14.09 | Electrocardiograph and Monitoring Paper - Roll
|  | 14.1 | Electrocardiograph and Monitoring Paper - Z-Fold
|  | 14.11 | Intrapartum Monitoring Paper
|  | 14.12 | Haemodynamic Monitoring Cables / Hardware
|  | 14.13 | Temperature Probes, single use
|  | 14.14 | Temperature Probe Adaptor Cables / Accessories
|  | 14.15 | Pulse Oximeter Sensors, Single Patient use
|  | 14.16 | Pulse Oximeter Sensors, Reusable
|  | 14.17 | Pulse Oximeter Extension and Adaptor Cables
|  | 14.18 | Sensory Awareness Monitoring Sensors, Single-use
|  | 14.19 | Monitoring Products, Other for Physiological Monitoring
15 | Respiratory Products for Physiological Monitoring | 15.01 | Gas Sampling Lines, Anaesthesia Gases
|  | 15.02 | Gas Sampling Lines, EtCO2 Monitoring lines
|  | 15.03 | Gas Sampling Line water traps
|  | 15.04 | EtCO2 Monitoring lines, high ambient humidity
|  | 15.05 | EtCO2 Airway Adapters
|  | 15.06 | Respiratory Products, Other for Physiological Monitoring
16 | Physiological Monitoring Accessories | 16.01 | Pulse Oximeters, Desktop
|  | 16.02 | Pulse Oximeters, Handheld
|  | 16.03 | Aneroid Sphygmomanometers
|  | 16.04 | NIBP Monitors
|  | 16.05 | Respiratory Gas & EtCO2 Monitors
|  | 16.06 | Vital Signs Monitors
|  | 16.07 | Wareable Devices
|  | 16.08 | Dopplers, Vascular & FHR
|  | 16.09 | Bladder Scanners
|  | 16.1 | Mid-Acuity Monitors
|  | 16.11 | High-Acuity Monitors & Modules
|  | 16.12 | Central Station Monitors
|  | 16.13 | TCP Monitors
|  | 16.14 | Cardiac Output Monitors
|  | 16.15 | Regional Oximetry Monitors
|  | 16.16 | Depth-of-Anaesthesia Monitors
|  | 16.17 | ECG Machines
|  | 16.18 | Cardiac Stress Testing Systems
|  | 16.19 | Holter Monitors
|  | 16.2 | Ambulatory Blood Pressure Monitors
|  | 16.21 | CTG Machines
|  | 16.22 | Electroencephalography (EEG) Units
|  | 16.23 | Polysomnography (PSG) Units
|  | 16.24 | Electromyography (EMG) / Evoked Potential (EP) Systems - Nerve Conduction Study (NCS)
|  | 16.25 | Portable Suction Devices
|  | 16.26 | Aspirators
|  | 16.27 | Analgesia Units
|  | 16.28 | 50/50 Oxygen/Nitrous Oxide Demand Valves
|  | 16.29 | Anaesthesia Workstations
|  | 16.3 | Accessory, Other, for Physiological Monitoring
|  | 16.31 | Gas Device
17 | Physiological Monitoring Equipment Spare Parts | 17.01 | Pulse Oximeters, Desktop
|  | 17.02 | Pulse Oximeters, Handheld
|  | 17.03 | Aneroid Sphygmomanometers
|  | 17.04 | NIBP Monitors
|  | 17.05 | Respiratory Gas & EtCO2 Monitors
|  | 17.06 | Vital Signs Monitors
|  | 17.07 | Dopplers, Vascular & FHR
|  | 17.08 | Bladder Scanners
|  | 17.09 | Mid-Acuity Monitors
|  | 17.1 | High-Acuity Monitors & Modules
|  | 17.11 | Central Station Monitors
|  | 17.12 | TCP Monitors
|  | 17.13 | Cardiac Output Monitors
|  | 17.14 | Regional Oximetry Monitors
|  | 17.15 | Depth-of-Anaesthesia Monitors
|  | 17.16 | ECG Machines
|  | 17.17 | Cardiac Stress Testing Systems
|  | 17.18 | Holter Monitors
|  | 17.19 | Ambulatory Blood Pressure Monitors
|  | 17.2 | CTG Machines
|  | 17.21 | Electroencephalography (EEG) Units
|  | 17.22 | Polysomnography (PSG) Units
|  | 17.23 | Electromyography (EMG) / Evoked Potential (EP) Systems - Nerve Conduction Study (NCS)
|  | 17.24 | Portable Suction Devices
|  | 17.25 | Aspirators
|  | 17.26 | Analgesia Units
|  | 17.27 | Anaesthesia Workstations
|  | 17.28 | Spare Parts, Other, for Physiological Monitoring
|  | 17.29 | Telemetry Units
18 | Physiological Monitoring Service & Maintenance (S&M) Agreements | 18.01 | Physiological Monitoring Service Maintenance Agreement - Comprehensive
|  | 18.02 | Physiological Monitoring Service Maintenance Agreement - Biomed Partnership
|  | 18.03 | Physiological Monitoring Service Maintenance Agreement - Medium Level
|  | 18.04 | Physiological Monitoring Service Maintenance Agreement - Lowest Level
|  | 18.05 | Physiological Monitoring Service Maintenance Agreement - Preventative Maintenance only
|  | 18.06 | Physiological Monitoring Software Licence and Maintenance Agreement (Software Only)
|  | 18.07 | Physiological Monitoring Software Licence and Maintenance Agreement (Software & Hardware)
|  | 18.08 | Physiological Monitoring Parts Only Agreement
|  | 18.09 | Physiological Monitoring Extended Warranties
|  | 18.1 | Physiological Monitoring BME Training
|  | 18.11 | Physiological Monitoring Service Maintenance Agreement - Other
|  | 18.12 | Physiological Monitoring Service - Service Support
|  | 18.13 | Physiological Monitoring Service - Software Support
Products Out-of-Scope
Electronic Thermometers, Contact |  |  | HPVC2023-021
Electronic Thermometers, non-Contact |  |  | HPVC2023-021
Fingertip Pulse Oximeters |  |  | HPVC2023-021
Pulse Oximeter Probes/Sensors |  |  | HPVC2023-021
Blood Glucose Monitors |  |  | HPVC2018-115
INR Machines |  |  | HPVC2018-115
Test Strips, Blood Glucose/Ketone/Coagulation |  |  | HPVC2023-042
Urinalysis dipsticks |  |  | HPVC2023-042
Oesophageal intubation detectors (Colorimetric CO2 Detectors) |  |  | out-of-scope
Auroscope/Ophthalmoscopes/Diagnostic Sets |  |  | out-of-scope
Vascular Dopplers, Blood flowmeters and Laser blood flowmeters |  |  | out-of-scope
Vascular Dopplers, Laser blood flowmeters |  |  | out-of-scope

## Appendix 2-Document Reference
Appendix 2 - Document References
Standards | Standard Number | Standard Name
References to standards include any amendments, revisions or consolidations to those standards | AS/NZS IEC 60601.1 | Medical Electrical Equipment – Part 1: General requirements for basic safety and essential performance
| AS IEC 60601.1.2 | Medical electrical equipment – General requirements for basic safety and essential performance – Collateral Standard: Electromagnetic disturbances – Requirements and tests
| AS IEC 60601.1.6 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral standard: Usability
| AS IEC 60601.1.8 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems
| AS IEC 60601.1.9 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: Requirements for environmentally conscious design
| AS IEC 60601.1.11 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment
| AS IEC 60601.1.12 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: Requirements for medical electrical equipment and medical electrical systems intended for use in the emergency medical services environment
| AS 60601.2.10 | Medical electrical equipment - Particular requirements for the basic safety and essential performance of nerve and muscle stimulators
| ISO 80601-2-13 | Medical electrical equipment - Part 2-13: Particular requirements for basic safety and essential performance of an anaesthetic workstation
| AS 60601.2.23 | Medical electrical equipment - Particular requirements for the basic safety and essential performance of transcutaneous partial pressure monitoring equipment
| AS/NZS IEC 60601.2.25 | Medical electrical equipment - Particular requirements for the basic safety and essential performance of electrocardiographs
| AS/NZS IEC 60601.2.27 | Medical electrical equipment - Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment
| AS 80601.2.30 | Medical electrical equipment - Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometers
| IEC 60601-2-34 | Medical electrical equipment - Part 2-34: Particular requirements for the basic safety and essential performance of invasive blood pressure monitoring equipment
| IEC 60601-2-37 | Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment
| IEC 80601-2-49 | Medical Electrical Equipment - Part 2.49: Particular Requirements for the basic safety and essential performance of patient monitors
| AS/NZS 4059 | Anaesthetic machines - Non-electrical - For use with humans
| AS 2902 | Medical gas systems - Low pressure flexible hose assemblies
| AS 2473.3 | Valves for compressed gas cylinders - Outlet connections for medical gases (including pin-indexed yoke connections)
| ISO 32 | Gas cylinders for medical use - Marking for identification of content
| AS 2896 | Medical gas systems - Installation and testing of non-flammable medical gas pipeline systems
| AS/NZS 4187 | Reprocessing of reusable medical devices in health service organizations
Standards for Associated Services | Standard Number | Standard Name
| AS/NZS 3551 | Management programs for medical equipment
| AS/NZS IEC 60601.1 | Medical Electrical Equipment – Part 1: General requirements for basic safety and essential performance
| AS/NZS 2500 | Guide to the safe use of electricity in patient care
Legislation | Name
References to legislation includes any amendments, revisions or consolidations to those references | Electricity Safety Act 1998
Occupational Health and Safety Act 2004
Building Act 1993
Guidelines and Other References | Issuing Body | Title
References to guidelines include any amendments, revisions or consolidations to those guidelines | Australian and New Zealand College of Anaesthetists (2013) | PS54: Statement on the Minimum Safety Requirements for Anaesthetic Machines and Workstations for Clinical Practice
| Australasian Sleep Association (2017) | Guidelines for sleep studies in adults
| Australian Resuscitation Council (2010) | ARC Guidelines
| NHMRC (2010) | Australian Guidelines for the Prevention and Control of Infection in Healthcare, Commonwealth of Australia
| Therapeutic Goods Administration (1991) | Therapeutic Goods Order No. 37: General Requirements for Labels for Therapeutic Goods
| Therapeutic Goods Administration (2004) | Uniform Recall Procedure for Therapeutic Goods, Commonwealth of Australia
| WorkSafe Victoria (2000) | Code of Practice for Manual Handling
| WorkSafe Victoria (2013) | Code of practice: The storage and handling of dangerous goods
| WorkSafe Victoria (2009) | Transferring People Safely
| WorkSafe Victoria (2007) | A Guide to Designing Workplaces for Safer Handling of People

## Appendix 3-Tables
Appendix 3 - Tables
Delivery
Order type | Maximum delivery time | Time prior to installation
Standard | 90 days | 7 days
Accessories | 90 days
Consumables / semi-consumables | 5 days
Batteries | 5 days
Spare Parts (common & essential) | 5 days
Preventative Maintenance Kits | 5 days
Urgent Deliveries | > 24 business hours (metro) > 48 business hours (regional & rural)
Service & Maintenance
Device Type | 1st Level Support | 2nd Level Support [1] | 3rd Level Support | Spare Parts Delivery 'common' and 'essential'
Portable Equipment | < 2 hours via phone | Courier to service base and Returned to PHS within 14 days | Loan Unit | < 5 days
Anaesthetic Machines, and fixed equipment | < 2 hours via phone | On-site within 24 business hours (metro) 2 business days (regional & rural) | Loan Unit within 60 business hours (metro) 4 business days (regional) | < 5 days
ICT Systems, and Central Stations et al. | < 2 hours via phone | Remote Access within 4 hours | On-site within 12 business hours (metro) 24 business hours (regional & rural) | 24 business hours
[1] Unless resolved by phone support