## IMPORTANT- Instructions
Part 4 Statement of Requirements - Instructions
HSVITS2026-060 Beds, Patient Support, Treatment Surfaces, and Mobility Solution
IMPORTANT - Do not insert new rows in between the worksheets. - Do not alter formatting and structure of this xls file. - Incomplete data will result in non-award of the tendered item.
Instructions
· Respondents are to review each worksheet and provide a response where required.
· For cells requesting "--Select--" - Respondents to respond by selecting Fully Comply, Partially Comply, Does Not Comply, or Not Applicable. For a response:
· "Fully Comply" - it is optional for Respondents to add further information in the 'Please Specify' column
· "Partially Comply" - Respondents to provide further information in the 'Please Specify' column
· "Does Not Comply" - no response is required in 'Please Specify' column, however if Respondents offers a valid alternative then provide this information in the 'Please Specify' column for consideration in the Evaluation
· "Not Applicable" - no response is required in 'Please Specify' column, however further information provided by Respondent in the 'Please Specify' column will assist the Evaluation
· Please note the specifications advised in this document are the minimum specification requirements only, any equipment that exceeds these requirements are fully compliant.
· HSV request Respondents advise specifications in the unit of measurement outlined in this document for easy comparison between different manufacturers.
· Where specified for Respondents to advise specific ranges/specifications, Respondents to provide these details in the 'Fully comply, Partially comply, Does not comply' column unless specified elsewhere.
Part A: Introduction
· In addition to responding to the HSV Procurement Portal questions and Tender Response Worksheet (Tender Response Worksheet), Total Cost of Ownership (TCO), this xls file outlines the scope of this Invitation to Supply (ITS) submission.
Part B: Service, Delivery & Support Part C: General Requirements Part D: Product Specification
· Responses are limited to Fully comply, Partially comply, Does not comply, Not Applicable
· Please provide any further information in the 'Please Specify’ column.
Category Tabs
· Each equipment device/ product listed within submitted Tender Response Worksheet (TRW) to be listed within this xls file too.
· Items can only be listed in one category/subcategory. Supplier to determine best fit category.
· Ensure contract holder part numbers and pricing aligns with TRW and product catalogue. Any products listed in this file and not TRW, will not be evaluated and will not be awarded.
· Only one code per column i.e. different configurations to be split across columns due to different codes, pricing and product information.
· Responses are limited to Fully comply, partially comply, Does not comply, Not Applicable
· Where you have listed partially comply or does not comply, Please provide supporting information in the 'Please Specify’ column.

## Index
Index
Introduction | Part A - Introduction
Part B - Service, Delivery & Support
Part C - Product Requirements
Hospital Beds, Acute Care | Category 1.01 - General Ward Beds, Single
Category 1.02 - General Ward Beds, King Single
Category 1.03 - General Ward High-Low Beds, Single
Category 1.04 - General Ward High-Low Beds, King Single
Category 1.05 - Bariatric Beds
Category 1.06 - Critical Care Beds
Category 1.07 - Stand-Up Beds
Category 1.08 - Homecare Beds
Hospital Beds, Mental Health | Category 2.01 - Behavioural Healthcare Beds
Category 2.02 - Seclusion Room Beds
Category 2.03 - Behavioural Healthcare Mattresses
Neonate & Paediatric Beds | Category 3.01 - Bassinets
Category 3.02 - Cots
Category 3.03 - Youth Beds
Category 3.04 - Infant Incubators
Category 3.05 - Infant Incubators, Transport
Category 3.06 - Infant Incubator/Radiant Warmers
Category 3.07 - Infant Resuscitation Trolleys
O&G Beds | Category 4.01 - Birthing Beds
Category 4.02 - Gynaecology Chairs
Category 4.03 - Gynae Exam Couches
Category 4.04 - Hospital Stretchers, Obstetric
Procedure Chairs | Category 5.01 - Haemodialysis / Chemotherapy Chairs
Category 5.02 - Procedure Chairs
Category 5.03 - Podiatry Chairs
Category 5.04 - Ophthalmology / ENT Chairs
Category 5.05 - Blood Collection Phlebotomy Chairs
Examination Couches | Category 6.01 - Examination Plinths
Category 6.02 - Examination Tables, 2-section
Category 6.03 - Examination Tables, 3-section
Category 6.04 - Podiatry Examination Couches
Category 6.05 - Tilt Tables
Category 6.06 - Traction Tables
Category 6.07 - Ultrasound Exam Tables
Category 6.08 - Ultrasound Exam Tables, Echocardiogram
Category 6.09 - Ophthalmology / ENT Examination Couches
Patient Stretchers | Category 7.01 - Hospital Stretchers, Manual
Category 7.02 - Hospital Stretchers, Powered
Category 7.03 - Hospital Stretchers, Motorised-Drive
Category 7.04 - Hospital Stretchers, Radiographic
Category 7.05 - Hospital Stretchers, MRI Conditional
Category 7.06 - Hospital Stretcher Chairs
Category 7.07 - Hospital Stretchers, Ophthalmology / ENT
Category 7.08 - Stretcher Mattresses
Operating Tables | Category 8.01 - Universal Operating Tables
Category 8.02 - Spinal Surgery Operating Tables
Category 8.03 - 'Operating Table Mattresses
Category 8.04 - Patient Positioning Pads & Mattress Pads
Patient Movers | Category 9.01 - Bed movers
Category 9.02 - Patient Lifters, Ceiling-mounted / Mobile-gantry
Category 9.03 - Patient Lifters, Mobile
Category 9.04 - Patient Lifters, Air-Assisted or other
Category 9.06 - Standing Transfer Aids, Mobile
Category 9.07 - Full-body Transfer Board
Category 9.08 - Air-Assisted Transfer Devices
Body Support Surfaces | Category 10.01 - Hospital Bed Mattresses
Category 10.02 - Hospital Bed Mattresses, Hybrid
Category 10.04 - Mattress Systems, Alternating-Pressure
Category 10.05 - Mattress Systems, Alternating-Pressure, Hybrid
Category 10.06 - Mattress Systems, Low Air Loss
Category 10.07 - Mattress Systems, Lateral Rotation/Percussion
Spare Parts | Category 13.0 - Spare Parts
Service & Maintenance Arrangements | Category 14.0 - Service & Maintenance Arrangements
Appendices | Appendix 1 - Categories
Appendix 2 - Document References
Appendix 3 - Tables

## Definitions
Definitions
Term | Definition
Accessory | a non-essential peripheral component that provides supplementary functionality or user convenience, without being required by the equipment to operate as intended
Agreement | means the Agreement entered into by HSV and a Respondent for the provision of [enter ITS Name]. It comprises the Master Supply Agreement, the Module and all Schedules, Annexures of any kind and any attachments.
ARTG | Australian Register of Therapeutic Goods under the control of TGA
Authorised User | Any person that is authorised by a Participating Health Service to perform certain functions (e.g. settings and software configuration) on Ventilator systems (e.g. Biomedical Engineers).
Business Day | Any weekday that is not gazetted as a public holiday in Melbourne, Victoria.
Business Hours | The Hours During The Day In Which Business Is Commonly Conducted (8:00 – 17:00)
CPR | Cardiopulmonary resuscitation
cold start | Starting from mains power being off with configurations and program processing beginning from initial values
Consumable | A part that is intended to be used for a limited time period or limited number of uses on a single patient
proprietary consumable | a product that is only available from one supplier and only functions with one brand of device
associated consumable | a product that is available from one or more suppliers that only functions with one brand of device
generic consumable | a product that is available from multiple suppliers and operates on multiple brand of devices
Deliverables | The goods and/or services to be supplied by a successful Respondent pursuant to an Agreement.
DIFOT | Delivery in full, On time
Effective Life | The minimum lifespan of a piece of equipment as defined by the Department of Health’s Medical Equipment Asset Management Framework (MEAMF).
eMR | Electronic Medical Records
equipment | a fixed asset, a piece of hardware, an apparatus, a piece of machinery
firmware | Software required by a device, for the device hardware to perform device’s intended purpose.
Fixed | Items of equipment that are not intended to be moved by end users for normal use.
FOC | free of charge
foetus | unborn baby
GS1 | Global Standards One
GTIN | Global Trade Identification Number
Handheld | Items of equipment that are designed to be transported by, and operated in, one hand.
HIS | Hospital Information System
HSV | HealthShare Victoria
Infant | a patient from 4 weeks of age from term to weighing 10kg
IPXX | 1.4
ITS | this Invitation To Supply
May | Indicates an optional element; it is at the Respondent’s discretion to either meet or not meet this element, and failure to meet this element will not have an impact during evaluation.
Mobile Equipment | Items of equipment that can be transported with transportation aides (such as trolleys, roll-stands etc.).
Must | Indicates a mandatory requirement; failure to meet this requirement will have a significant negative impact during evaluation and may result in non-award. HSV may choose to amend such requirements at any stage during the tendering stage, based on either, the deviation requests received from tenderers, or otherwise.
NEHTA | National E-Health Transition Authority
Neonate | newborn baby less than 4 weeks old, from term
Normal Use | Means that the item has undergone use for which it was manufactured and intended and shows no signs of physical damage other than regular wear and tear.
NPC | National Product Catalogue
OEM | Original Equipment Manufacturer
Paediatric | a patient greater than 10kg and less than 18 years old
Participating Health Services (PHS) | Public Hospitals and other Health or Related Services, as those terms are defined in Section 3 of the Health Services Act 1988 (Vic), that are described in PART 6: REPORTING GUIDELINES AND HOSPITAL PARTICIPATION.
PC | Personal Computer
Portable | Items of equipment that are designed to be lifted and carried by hand.
RFI | Request for Information
RFID | Radio Frequency Identification
Respondent | Any person, company or organisation responding to this ITS and, unless the context otherwise requires, includes those who may access the ITS for the purpose of creating and submitting a Response.
Reusable | A device designed or intended by the manufacturer as suitable for reprocessing and reuse.
Semi-Consumable | An item that is used multiple times, has a longer lifespan than consumable, not expected to last the lifespan of the equipment; such as ECG trunk cable, ECG lead-wires
Should | Indicates a highly desirable element; unless justifiable reason exists, not meeting this element may impact evaluation.
SLA | Service Level Agreement
Single Use | A device that is intended to be used on an individual patient, during a single procedure, and then discarded.
Single-Patient Use | A device that can potentially undergo more than one episode of use on one patient only. The device may need to undergo some form of reprocessing between each use (in accordance with manufacturers’ instructions).
Spare Part | Any item required for any preventative or corrective maintenance process in maintaining tendered machine
common | a part that has a Failure Rate (MTBF) of <12 months, or
purchased more than once in a 12 month period by the sum all PHSs on this agreement.
essential | a part required in the equipment for it to function as prescribed in Intended Use, e.g. motherboard, power supply, etc.
generic | a part that is available from multiple Suppliers
infrequent | a part that needs to shipped to Australia from overseas
proprietary | a part that is only supplied by Respondent
Supplier | A successful Respondent pursuant to an Agreement entered into for supply and delivery.
Supported Life | The duration after purchase in which the successful Respondent provides ongoing service and/or parts for the equipment (the minimum supported life is specified in clause 5.11.
TCP/IP | Transmission Control Protocol/Internet Protocol
TGA | Therapeutic Goods Administration
Train-The-Trainer | Any person that is authorised by a Participating Health Service to provide internal training to end users on the use of Ventilator systems and associated products.
Transportable | Items of equipment that are designed to be moved only with transportation aides (such as roll stands) or that are only movable from bracket to bracket.
TRW | Tender Response Worksheet
UPS | Uninterruptable Power Supply
URPTG | TGA's Uniform Recall Procedure for Therapeutic Goods (URPTG) process
Will | Indicates an anticipated future condition.

## A - Introduction
return to Index page
Part A: Introduction
| Clause Identifiers |  |  |  | Text
Part A
|  | 1 |  |  | Purpose
|  |  | a |  | The purpose of this Part 4 – Statement of Requirements, is to:
|  |  |  | i | Detail the scope and range of products sought under this Invitation to Supply (ITS)
|  |  |  | ii | Specify the requirements that Respondents and/or their offered products must meet. These requirements also form part of any resulting Agreement between HPV and any successful Respondent(s).
|  | 1.1 |  |  | Participating Health Services
The Participating Health Services for this ITS are:
All ‘Public Health Services’ (as legislatively defined) referred to in Schedule 1 and Schedule 5 of the Health Services Act 1988; and
Other relevant participating health and health related organisations as follows:
Department of Health (DH) (Vic)
Gelantipy District Bush Nursing Centre Incorporated
Harrow Bush Nursing Centre
Mallee Track Health and Community Service
Mercy Hospitals Victoria Ltd
Mercy Palliative Care Ltd
Orbost Regional Health
St. Vincent's Hospital Melbourne
Woomelang and District Bush Nursing Centre
|  | 2 |  |  | Scope
|  |  | b |  | The scope of this Invitation to Supply includes:
|  |  |  | i | supply of Beds Mattresses Patient Trolleys and Treatment Chairs, including optional features
|  |  |  | ii | supply of all relevant consumables and accessories
• including all proprietary, associated and generic
|  |  |  | iii | the servicing and maintenance of item/s Respondent has offered
|  |  |  | iv | delivery of a training package, including relevant training materials.
|  |  | c |  | HPV is seeking responses for various supply solutions, including but not limited to:
|  |  |  | i | outright purchase of equipment
leasing of equipment (via a Managed Equipment Services, etc.)
|  |  |  | ii | supply of equipment only
supply of consumables and/or semi-consumables only
|  |  |  | iii | whole-of-life support/managed supply environment, including but not limited to: supply of equipment and all associated consumables maintenance (including scheduled inspection and unscheduled repairs) warranty services disposal services.
|  |  |  | iv | Supply and rental of all equipment within this contract
|  |  | d |  | Respondents may offer products in one, some or all categories listed in Appendix 1 – Product List.
|  | 3 |  |  | Product Categories
|  |  | a |  | A complete range of Beds, Mattresses, Patient Trolleys and Treatment Chairs is required for treatment of patients across Victorian Public Health Services
|  |  | b |  | The categories required include:
|  |  |  | i | 1.0 - Hospital Beds, Acute Care
|  |  |  | iii | 2.0 - Hospital Beds, Mental Health
|  |  |  | iv | 3.0 - Neonate & Paediatric Beds
|  |  |  | v | 4.0 - Obstetrics & Gynaecology Beds
|  |  |  | vi | 5.0 - Procedural Chairs
|  |  |  | vii | 6.0 - Examination Couches
|  |  |  | viii | 7.0 - Patient Stretchers
|  |  |  | ix | 8.0 - Operating Tables
|  |  |  | x | 9.0 - Patient Movers
|  |  |  | xi | 10.0 - Body Support Surfaces
|  |  |  | xii | 11.0 - Rental
|  |  |  | xiii | 12.0 - Accessories
|  |  |  | xiv | 13.0 - Spare Parts
|  |  |  | xv | 14.0 - Service & Maintenance Arrangements
|  |  | c |  | The Respondent may offer products in one, some or all categories. (DO NOT DUPLICATE PRODUCTS ACROSS CATEGORIES/SUBCATEGORIES) Ensure all items tendered in SOR are in TRW with pricing and matching part numbers
|  |  | d |  | HSV reserves the right not to consider any additional products offered.
|  |  | e |  | For a full list of products refer to Appendix 1 – Product List.
|  | 4 |  |  | Product Offering
|  |  | a |  | HSV may not consider any product that is subject to a current HSV Agreement, other than those listed below:
|  |  |  | i | HPVCS2019-060 Beds Mattresses Patient Trolleys and Treatment Chairs
|  |  | b |  | Respondent will ensure that each product is offered in only one subcategory. It is the Respondent’s responsibility to ensure that each product is submitted in the most appropriate subcategory.
|  | 5 |  |  | Reference Sites
|  |  | a |  | Respondent must provide a minimum of one reference for equipment tendered that is not awarded under the current HPVC2019-060 Beds contract with sales. Each reference must include the referee's position title, phone number, health service email address and a declaration of no conflict of interest, detail the product in question including part number and clinical findings and be on the HSV approved reference template.
A desktop evaluation will not suffice. Referees must be from a public sector background.
References are required for all equipment, consumables and third party consumables not on current contract with health service usage. References are NOT required for accessories, options or spare parts.
|  |  | b |  | Respondents will be given preference where reference sites are from the Australasian health service list in the Tender Response Worksheet.
• Categories 1, 5-7, 9-10
- References from a Royal Australasian College of Physicians accredited training site offering Level 3 Principal Training Program. List of hospitals are provided in the Tender Response Worksheet from the drop down list.
https://www.racp.edu.au/trainees/accredited-settings/basic-training-accredited-settings
• Category 2
- References from a Royal Australasian College of Physicians accredited training site offering Level 3 Principal Training Program where treatment for acute mental health inpatients is offered. List of hospitals are provided in the Tender Response Worksheet from the drop down list.
https://www.racp.edu.au/trainees/accredited-settings/basic-training-accredited-settings
• Categories 3 & 4
- References from a Royal Australasian College of Physicians accredited training site offering either Perinatal or General Paediatrics training. List of hospitals are provided in the Tender Response Worksheet from the drop down list.
https://www.racp.edu.au/trainees/accredited-settings/advanced-training-accredited-settings
• Category 8.01
- References from General Surgeons Australia accredited training hospital, Australian Orthopaedic Association accredited training hospital, Royal Australasian College of Physicians accredited training hospital offering Level 3 Principal Training Program. List of hospitals are provided in the tender response worksheet from the drop-down list.
https://generalsurgeons.com.au/wp-content/uploads/2024/08/ACCREDITATION_List-of-Accredited-Hospital-Posts.pdf
https://aoa.org.au/docs/default-source/training-(public)/accreditation/aoa-sccredited-training-sites-and-positions-sept-2024.pdf?sfvrsn=b6dfad04_4
https://www.racp.edu.au/trainees/accredited-settings/advanced-training-accredited-settings
• Category 8.02
- References from a Neurosurgical Society of Australasia Post Fellowship Education and Training Committee administered program offering positions in spinal surgery. List of hospitals are provided in the tender response worksheet from the drop-down list.
https://nsa.org.au/common/Uploaded%20files/PFET/PEFT_Positions.pdf
|  |  |  | i | that currently use, or have trialled and evaluated, the offered deliverable.
|  |  | c |  | HSV reserves the right to validate the reference with the nominated referee and seek additional feedback on findings of the trial conducted and the clinical acceptability of the deliverables.
|  |  | d |  | Respondents must obtain referee permission prior to listing them as a referee
|  |  | e |  | For products tendered without a reference, i.e. new product, it is at the discretion of HSV to non-award. Where an item tendered is accompanied by a reference from a site not listed from the Australasian health service list in the Tender Response Worksheet, it is at the discretion of HSV to determine clinical acceptability and whether or not there are alternative suitable contracted items available on contract.
|  | 6 |  |  | Clinical Trials
|  |  | a |  | The Participating Health Service may, at their discretion, research or trial new technology or use non-contracted products to perform clinical trials at any time during the Term of any resulting Agreement.
|  |  | b |  | The Participating Health Service may request trials before creating a Statement of Demand which may also provide the Supplier time to do a site inspection.
|  |  | c |  | Trials at each Participating Health Service site will last between one and four weeks, as mutually agreed upon by the Participating Health Service and the Supplier.
|  |  | d |  | The Supplier should be prepared to demonstrate their offered deliverable within four weeks of request.
|  |  | e |  | All products required to trial the offered deliverable will be provided at no cost for the trial period.
Where a clinical trial has been conducted during the life of the contract and a reference (VIC public health service/ from approved reference list) can validate the trial and health service demand, the supplier must notify HSV and request an addition to contract prior to The Participating Health Service raising a purchase order. Purchase orders must only be raised once HSV has approved the addition to contract.

## B - Service, Delivery & Support
return to Index page
Part B: Service, Delivery and Support
| Clause Identifiers |  |  |  | Text |  |  |  | No extra columns to be added by Respondent
| Part B |  |  |  | Service, Delivery & Support | Response Type | Compliance | Respondent's Response
|  | 1 |  |  | Delivery and Installation
|  |  | a |  | Offered equipment must be delivered to the location(s) specified by the Participating Health Service as quickly as possible, and in no event later than 90 days (three calendar months) from the purchase order date, unless a longer timeframe is agreed to in writing by the Participating Health Service. | Select Compliance | --Please Select--
|  |  |  | i | The successful Respondent must immediately notify both HSV and the Participating Health Service if equipment delivery is expected to exceed 90 days from the purchase order date. | Select Compliance | --Please Select--
|  |  | b |  | The successful Respondent must install and configure the offered equipment and/or system at the Participating Health Service, ensuring it is ready for clinical use. (see Appendix 3 Tables - Delivery timeframes) | Select Compliance | --Please Select--
|  |  |  | i | At the request of a Participating Health Service, the Respondent must collaborate on installation, coordinating with the health service's biomedical engineering and ICT departments (see Part B, Section 7, Biomedical Equipment Maintenance and Repair). | Select Compliance | --Please Select--
|  |  | c |  | Consumables and semi-consumables must be delivered to Participating Health Services within five business days of the purchase order date, unless otherwise agreed by the Participating Health Service. | Select Compliance | --Please Select--
|  |  |  | i | If delivery will exceed five business days, the Supplier must notify the purchase requestor in writing (email) immediately upon becoming aware of the delay. | Select Compliance | --Please Select--
|  |  | d |  | Following receipt of a purchase order, the Supplier must deliver all common and essential spare parts for the offered equipment to the Participating Health Service within the delivery timelines specified in Appendix 3 Tables, Service & Maintenance. | Select Compliance | --Please Select--
|  |  |  | i | If the Supplier anticipates a delay, they must notify the Participating Health Service before the timelines expire. | Select Compliance | --Please Select--
|  |  |  | ii | In the event of equipment failure and the Supplier's inability to deliver said spare parts within the specified timelines, they should provide a comparable loan unit to the Participating Health Service. | Select Compliance | --Please Select--
|  |  | e |  | For delivery timeframes, please refer to Appendix 3, Tables – Delivery | Info to Respondent | --Please Select--
|  |  | f |  | Equipment offered under this ITS may be procured through outright purchases or rental arrangements. | Info to Respondent | --Please Select--
|  |  | g |  | For outright purchases, the contract price must be free into store (FIS). This means the cost of delivery is included in the total price, and no additional delivery fees will be charged. | Select Compliance | --Please Select--
|  |  | h |  | Rental equipment will have a separate pricing structure from standard FIS outright purchases. The pricing is split by the rental period and delivery zone: | Info to Respondent | --Please Select--
|  |  |  | i | Rental Periods: Daily, weekly, fortnightly, monthly, and Hospital in the Home (HITH).
|  |  |  | ii | Delivery Zones: Metropolitan, Regional, and HITH
|  |  |  | iii | The rental contract price will include the applicable zoned delivery fee.
|  |  |  | iv | The supplier shall deliver rental equipment to the Participating Health Service during business hours.
Note: Rental equipment provided under this ITS is for hospital inpatients only; patients in a Hospital in the Home (HITH) program are considered hospital inpatients. Rental equipment for patients who have been discharged from the hospital is not within the scope of this tender. The Participating Health Service must be the entity invoiced for all rental equipment provided. These requirements will be reflected in the Tender Response Worksheet (TRW) structure.
|  | 1.1 |  |  | Urgent Delivery
|  |  | a |  | For the purposes of this section, urgent deliveries refers to urgent requests placed by an individual Participating Health Service, and does not include emergency situations. | Info to Respondent | --Please Select--
|  |  | b |  | The Respondent must be able to receive and action urgent delivery requests during normal business hours. | Select Compliance | --Please Select--
|  |  | c |  | Urgent delivery purchase orders that are placed prior to Respondent’s order entry deadline, should be received by Participating Health Services within timelines as prescribed in Appendix 3 Tables – Delivery. | Select Compliance | --Please Select--
|  |  | d |  | Any additional costs associated with Urgent Deliveries must be clearly outlined in the User Guide tab of the Tender Response Worksheet and must be discussed and agreed with the ordering health service prior to acceptance of any urgent purchase order. | Select Compliance | --Please Select--
|  |  | e |  | The Respondent must maintain sufficient stock in country of the following for the duration of this agreement, or until the end-of-support date for the related equipment, whichever occurs first: | Select Compliance | --Please Select--
|  |  |  | i | for all tendered consumables
|  |  |  | ii | for all tendered semi-consumables
|  |  |  | iii | for all tendered batteries
|  |  |  | iv | for all tendered preventative maintenance service kits
|  |  |  | v | for all proprietary spare parts, both common and essential, required to maintain offered equipment
|  | 2 |  |  | Support
|  |  | a |  | The offered equipment should have a functional lifespan of at least ten (10) years from the purchase date. "Functional lifespan" is the length of time the equipment is fully operational and performs all intended functions as specified by the manufacturer. | Select Compliance | --Please Select--
[5 empty] | Offered non-powered redistribution support surfaces should have a functional lifespan of at least two (2) years from the purchase date. | Select Compliance | --Please Select--
|  |  | b |  | The Respondent must provide comprehensive support services for the offered equipment for a period of no less than ten (10) years from the release date of this ITS. | Select Compliance | --Please Select--
|  |  |  | i | "Support services" shall include, but not limited to: • On-site maintenance and repair • Spare parts availability • Software updates and patches • Telephone and online technical support • Training Participating Health Service's personnel | Info to Respondent | --Please Select--
|  |  | c |  | The Supplier must provide maintenance, spare parts, and support services to Participating Health Services for a minimum of seven years following final payment of offered equipment. | Select Compliance | --Please Select--
|  |  | d |  | The successful Respondent must provide written notification to both HealthShare Victoria and each affected Participating Health Service at least six (6) months prior to the announced end-of-sales date for any equipment offered under this agreement. | Select Compliance | --Please Select--
|  |  | e |  | The successful Respondent should provide comprehensive support services for the offered equipment for at least seven years following the announcement of the end-of-sale date. | Select Compliance | --Please Select--
|  | 2.1 |  |  | Customer Service
|  |  | a |  | The Respondent must provide customer support to Participating Health Services during business hours. | Select Compliance | --Please Select--
|  |  |  | i | Participating Health Services may preference Respondents who offer product support via telephone outside of business hours. | Info to Respondent | --Please Select--
|  |  | b |  | The Supplier must provide Participating Health Services with representatives that are: | Select Compliance | --Please Select--
|  |  |  | i | inherently familiar with the contracted products
|  |  |  | ii | appropriately qualified
|  |  |  | iii | technically/clinically knowledgeable about the contracted products
|  |  |  | iv | available to respond to Participating Health Services’ queries during business hours
|  |  | c |  | It is desirable that nominated Representatives have a clinical background or experience. | Info to Respondent | --Please Select--
|  |  | d |  | The level of customer service and support required of Representatives is expected to include (but is not limited to): | Info to Respondent | --Please Select--
|  |  |  | i | liaising with clinicians to recommend products and solutions
|  |  |  | ii | promptly answering clinicians’ queries (including after hours)
|  |  |  | iii | liaising with various hospital departments (for example: operating theatre, Nurse Unit Managers)
|  |  |  | iv | providing on-site clinical support during cases (if requested)
|  |  |  | v | providing informational materials
|  |  |  | vi | providing education and in-service training upon request.
|  |  | e |  | Supplier's representatives must comply with Participating Health Services’ local policies regarding engagement with Participating Health Service staff. | Select Compliance | --Please Select--
|  | 2.2 |  |  | Operator Training
|  |  | a |  | Upon request by a Participating Health Service, the successful Respondent must provide complete details of their training package for the offered equipment, including materials, trainee numbers, and session duration, to support implementation. | Select Compliance | --Please Select--
|  |  | b |  | Operator (User) training must: | Select Compliance | Fully Comply
|  |  |  | i | be provided FOC to Participating Health Services throughout the equipment's operational life, ensuring that all new employees receive the necessary training to operate the equipment correctly,
|  |  |  | ii | be provided by experienced and qualified personnel,
|  |  |  | iii | be provided on-site and in-person, unless the Participating Health Service requests otherwise,
• to accommodate relevant staff, as determined by Participating Health Service, across a range of shifts, include nights and weekends
|  |  |  | iv | upon request, be provided by videotelephony during business hours,
|  |  |  | v | upon request, be provided as a web-based training package for night shift and weekend workers,
|  |  |  | vi | include directions for the correct cleaning procedure and appropriate cleaning/disinfectant products, and
|  |  |  | vii | include complementary face-to-face superuser training for at least two employees per Health Service site, if relevant.
|  |  |  | viii | be designed and delivered in accordance with WorkSafe Victoria Occupational Health & Safety requirements to prevent workplace injuries.
|  |  | c |  | Educational and training materials must be provided at no cost and in digital format. | Select Compliance | --Please Select--
|  |  | d |  | Training aids directly related to the Deliverables must be tendered and priced separately. | Select Compliance | --Please Select--
|  |  |  | i | All training aid accessories (e.g., chargers, batteries) must also be tendered. | Select Compliance | --Please Select--
|  | 2.3 |  |  | Instruction For Use Manuals
|  |  | a |  | For each shipped piece of equipment, the Supplier must include a complete digital copy of the Instructions for Use manual. | Select Compliance | --Please Select--
|  |  | b |  | Instructions for Use manuals must be in English and must include but not be limited to: | Select Compliance | --Please Select--
|  |  |  | i | Easy-to-understand instructions to ensure the safe and intended use of the equipment,
|  |  |  | ii | Safety and precautions requirements,
|  |  |  | iii | Recommended cleaning procedure,
|  |  |  | iv | Alarm messages (where applicable) with troubleshooting tips
|  |  | c |  | Participating Health Services should be able to download IFU manuals from the successful Respondent's website. | Select Compliance | --Please Select--
|  |  | d |  | Participating Health Services may preference Respondents offering mobile app-based learning modules or QR codes on equipment that link directly to quick user guides. | Info to Respondent | --Please Select--
|  | 2.4 |  |  | Warranty
|  |  | a |  | Respondents should offer at least 60-month warranty on the structural frame (e.g. frames of bed, chairs, stretchers, examination couches, Operating tables) of offered equipment. Participating Health Services have expressed a preference for equipment featuring a minimum 60-month warranty on its structural frame. | Select Compliance | --Please Select--
|  |  | b |  | Participating Health Services have expressed interest in extended warranty coverage. Respondents are encouraged to offer extended warranties that go beyond the manufacture's warranty, if available. The procurement of any extended warranty will be at the discretion of each Participating Health Service. | Info to Respondent | --Please Select--
|  |  | c |  | All extended warranty exclusions must clearly indicate in the Tender Response Worksheet. | Select Compliance | --Please Select--
|  | 3 |  |  | Service
|  |  | a |  | Service and maintenance must be available throughout Victoria, including Albury-Wodonga. | Select Compliance | --Please Select--
|  |  | b |  | The Respondent must maintain a repair facility within the metropolitan area of an Australian capitol city. | Select Compliance | --Please Select--
|  |  | c |  | To ensure adequate support for the equipment offered, the Respondent must meet at least one of the following criteria regarding factory-trained field service engineers:
|  |  |  | i | have them directly employed and based in Victoria | Select Compliance | --Please Select--
|  |  |  | ii | access them through a subcontract arrangement within Victoria | Select Compliance | --Please Select--
|  |  |  | iii | offer their virtual support to Participating Health Services during standard business hours | Select Compliance | --Please Select--
|  |  |  | iv | be able to deploy them to Participating Health Services to meet the service and maintenance timelines outlined in Appendix 3 | Select Compliance | --Please Select--
|  |  | d |  | If the Supplier utilises authorised third-party service providers for ongoing support, they must provide a list of those providers to Participating Health Services. | Select Compliance | --Please Select--
|  |  | e |  | The successful Respondent must adhere to the Service & Maintenance response times specified in: | Select Compliance | --Please Select--
|  |  |  | i | Appendix 3: Tables - Service & Maintenance.
|  |  | f |  | A service report detailing all works performed must be provided to the Participating Health Service with each equipment service. | Select Compliance | --Please Select--
|  |  | g |  | Respondents should provide a full spectrum of service contract options for all offered equipment and applications, encompassing everything from comprehensive 24/7 maintenance agreements down to basic agreements. | Select Compliance | --Please Select--
|  |  |  | i | Participating Health Services may enter into these service contracts at any time during the equipment's effective life. | Info to Respondent | --Please Select--
|  |  | h |  | When a cleaning service is provided for a component or complete system, all applicable terms and conditions must be explicitly documented and communicated to the Participating Health Service. This document must clearly specify courier charges, turnaround times, detailed cleaning procedures, potential service limitations, and all other relevant service details. | Select Compliance | --Please Select--
|  |  | i |  | The supplier is responsible for all maintenance of the rented mattresses, encompassing thorough cleaning and disinfection between each patient use; all necessary repairs to maintain operational integrity and safety; and all scheduled preventative maintenance and safety inspections as per regulatory standards. These services must be included in the rental fee. | Select Compliance | --Please Select--
|  | 3.1 |  |  | Repairs under Warranty
|  |  | a |  | The Respondent must bear all costs associated with warranty repairs and replacements, including, but not limited to, pickup, delivery, and on-site visits. Participating Health Services are not responsible for any such costs, except for those directly resulting from the Health Service’s employee’s accident, negligence, or malicious acts. | Select Compliance | --Please Select--
|  |  |  | i | The Respondent should also provide a suitable loan unit, free of charge, to the Participating Health Service within 2 business days of being notified of the fault. | Select Compliance | --Please Select--
|  | 3.2 |  |  | Spare Parts and Repairs
|  |  | a |  | All common and essential spare parts must be stocked in Australia. | Select Compliance | --Please Select--
|  |  | b |  | The Supplier must maintain spare parts availability for at least five (5) years after the date of final manufacture of the offered equipment. | Select Compliance | --Please Select--
|  |  |  | i | These spare parts should be available for purchase by the Participating Health Service, regardless of their staff's service training status. | Select Compliance | --Please Select--
|  |  | c |  | Upon request by a Participating Health Service, the Supplier must provide a comprehensive list of replacement parts, accessories, and assemblies, including part numbers and current purchase prices, necessary to maintain and repair the offered equipment throughout its supported life. | Select Compliance | --Please Select--
|  |  | d |  | Upon request by a Participating Health Service, the Supplier or their appointed agent must provide non-warranty repairs for the offered equipment throughout its supported life. Consumables and semi-consumables should be replaced rather than repaired whenever practical. | Select Compliance | --Please Select--
|  | 3.3 |  |  | Biomedical Equipment Maintenance and Repair
|  |  | a |  | Respondents acknowledge that the Participating Health Service may utilise its internal biomedical engineering department or an external contractor to perform the following tasks on offered equipment: safety and performance testing; acceptance testing of new equipment (prior to clinical use); warranty assessment and coordination of warranty repairs with the successful Respondent; non-warranty repairs; routine preventative maintenance; and installation of user configuration settings. | Info to Respondent | --Please Select--
|  |  | b |  | The Supplier must provide ongoing biomedical service training to the Participating Health Service's biomedical engineering staff. | Select Compliance | --Please Select--
|  |  |  | i | In the Product Catalogue, provide details on the training delivery method (e.g., annual seminars, on-site face-to-face training, virtual sessions) and associated costs.
|  |  | c |  | After biomedical service training, Supplier using software to service or configure offered equipment must provide each trainee with a copy of the software and any required accessories (e.g., cables, dongles). | Select Compliance | --Please Select--
|  |  | d |  | Respondents must provide a complete list of all tools and software required for servicing and/or configuring the offered equipment in the Tender Response Worksheet, if applicable. | Select Compliance | --Please Select--
|  |  | e |  | Biomedical service training is separate from and in addition to the training requirements specified under Part B Section 3 Training. | Info to Respondent | --Please Select--
|  |  | f |  | Upon request by the Participating Health Service, the Supplier must provide comprehensive English-language technical service manuals, including full lists of all replaceable parts with part numbers, for the equipment purchased by Health Service. | Select Compliance | --Please Select--
|  | 4 |  |  | Key Performance Indicators
|  |  | a |  | Refer to Module C Schedule 6 | Info to Respondent | --Please Select--
|  | 5 |  |  | Service Level Agreement
|  |  | a |  | Participating Health Services may enter into a Service Level Agreement (SLA) with the Supplier(s), on the terms set out in the Draft Agreement clause 4.3 Service Level Agreement. The SLA may cover the following arrangements: the provision of products on consignment; requirements for stock management and rotation; arrangements for ordering, invoicing and delivery; social procurement commitments or framework set out by a Participating Health Service that may be linked to the Social Procurement Framework set out in the Agreement; clinical support, including attendance requirements for Representatives in relation to education and training; and /or communication arrangements for product recalls & safety alerts (see Part C Section 3 Recall Process). | Info to Respondent | --Please Select--
|  |  | b |  | The SLA will be in addition to the Agreement between the Supplier(s) and HSV and will not alter any terms of the Agreement. Any SLA entered into between the Supplier and a Participating Health Service must be established in accordance with the framework of the Agreement and must not contravene or undermine the terms of the Agreement. | Select Compliance | --Please Select--
|  |  | c |  | HSV will not be responsible for monitoring compliance with any SLA. This is a matter of agreement between the parties to the SLA. | Info to Respondent | --Please Select--
|  |  | d |  | Supplier(s) must provide a copy of all Service Level Agreements/ Order Contracts to HSV within one (1) week of being finalised. | Select Compliance | --Please Select--

## C - General Requirements
return to Index page
Part C: General Requirements
| Clause Identifiers |  |  |  | Text |  | (All) Product Compliance |  | No extra columns to be added by Respondent
| Part C |  |  |  | General Requirements | Response Type | Compliance | Respondent's Response
|  | 1 |  |  | Standards and Compliance
|  |  | a |  | All medical devices offered must comply with the Australian Therapeutic Goods Administration (TGA) regulatory requirements. This means either being listed on the ARTG or having an exemption. | Select Compliance | --Please Select--
|  |  |  | i | Upon request by a Participating Health Service, the Supplier must provide the ARTG certificate for the offered medical device. | Select Compliance | --Please Select--
|  |  | b |  | The offered equipment must adhere to all relevant Australian Standards, current legislation and regulations, and applicable clinical practice guidelines. Where an Australian Standard is not available, compliance with a recognised international standard is required. | Select Compliance | --Please Select--
|  |  |  | i | Refer to Appendix 2 - Document References for a list of collective Compliance Requirements.
|  |  | c |  | The offered equipment must operate without causing electromagnetic interference to other medical devices and must comply with an applicable Electromagnetic Compatibility (EMC) standard. | Select Compliance | --Please Select--
|  |  | d |  | The Manufacturer of offered therapeutic goods must conform with AS ISO 14971 Medical devices — Application of risk management to medical devices requirements. | Select Compliance | --Please Select--
|  | 1.1 |  |  | Recall Process
|  |  | a |  | HSV will immediately suspend any offered product on its Pricing Schedule subject to a TGA Procedure for Recall, Product Alerts and Product Corrections (PRAC) Class I life-threatening notification until the matter is rectified. | Info to Respondent | --Please Select--
|  |  | b |  | The Supplier must resolve all TGA PRAC market actions (both software and hardware) for the offered equipment within six months of notification. Failure to comply will result in suspension of offered product from the HSV catalogue. | Select Compliance | --Please Select--
|  |  |  | i | Class I (life-threatening) software defects must be resolved within three (3) months of notification. Failure to comply will result in the suspension of the product from the HSV catalogue. | Select Compliance | --Please Select--
|  |  | c |  | Suppliers should accept PRAC Notice Acknowledgement Forms from Participating Health Services through the GS1 Recall Health portal. | Select Compliance | --Please Select--
• will be at the discretion of Participating Health Service whether or not to submit Acknowledge Form via GS1 Recall Health.
|  | 1.2 |  |  | Cybersecurity
|  |  | a |  | Do any of the offered products have the capability to connect to the hospital's internal network? (Yes/No) If 'Yes' then: | Select Compliance | --Please Select--
|  |  |  | i | The Respondent must have a Cyber Incident Response Program in place to assist Participating Health Services in the event of a cyber incident affecting offered equipment. | Select Compliance | --Please Select--
|  |  |  | ii | The Supplier must immediately notify the Department of Health (Digital Health team), HealthShare Victoria, and all affected Participating Health Services within 72 hours of becoming aware of any intelligence report indicating active exploitation of a cybersecurity vulnerability in offered product. | Select Compliance | --Please Select--
|  |  |  | iii | Upon request by a Participating Health Service, the Supplier must provide the latest MDS2 form for the Deliverable. | Select Compliance | --Please Select--
|  |  |  | iv | Upon request by a Participating Health Service, the Supplier must provide the Software Bill of Materials (SBOM), or, in the event that the SBOM is not available, the Software of Unknown Provenance (SOUP). | Select Compliance | --Please Select--
|  |  | b |  | Equipment incorporating an embedded operating system (OS) should be maintained with continuous security updates and patches up to its documented end-of-support date. | Select Compliance | --Please Select--
|  |  |  | i | Security patches released by the OS vendor for the embedded operating system must be assessed and implemented within 30 calendar days of their public availability, or as soon as practicable if implementation within this timeframe is not feasible, with justification documented. | Select Compliance | --Please Select--
|  |  |  | ii | The security patch shall be provided at no additional cost to the Participating Health Service.
|  |  | c |  | Within 30 calendar days of a reputable and independent third party publicly disclosing any critical or high-severity vulnerability affecting the tendered system, the Supplier must provide and deploy, at no additional cost, a security patch or corrective update that remediates the vulnerability in accordance with the Essential Eight Mitigation Strategy. | Select Compliance | --Please Select--
|  |  |  | i | Where technical constraints prevent full remediation within this timeframe, the Supplier must implement an interim risk mitigation measure and provide a remediation plan agreed with the Participating Health Service.
|  |  | d |  | Should a security vulnerability be discovered within a tendered product and confirmed to be under active exploitation, the Supplier must: | Select Compliance | --Please Select--
|  |  |  | i | Provide an immediate temporary mitigation or workaround to all Participating Health Services affected by the vulnerability within 48 hours of being notified.
|  |  |  | ii | Provide a permanent remedial solution to all Participating Health Services affected by the vulnerability within 30 calendar days of being notified.
• Where technical constraints prevent full remediation within this timeframe, the Supplier must implement an interim risk mitigation measure and provide a remediation plan agreed with the Participating Health Service.
|  |  | e |  | All Participating Health Services use endpoint protection solutions (antimalware, firewalls, EDRs). Any tendered product seeking network connectivity to the hospital’s enterprise network (excluding the isolated patient monitoring LAN) must fully support integration and communication with an industry-standard endpoint protection solution. | Select Compliance | --Please Select--
|  |  | f |  | Participating Health Services are responsible for their own endpoint protection. While this ITS does not include endpoint protection, any solution offered by the Respondent must generate logs in industry-standard formats (CEF, Syslog) to allow seamless integration with hospital’s existing security information and event management (SIEM) or log repository systems for effective threat detection and monitoring. | Select Compliance | --Please Select--
|  |  | g |  | The Respondent must allow the Participating Health Service to enable and/or install their own endpoint protection solution on the Deliverable without incurring supplementary charges. | Select Compliance | --Please Select--
|  |  | h |  | Respondents must include a list of compatible endpoint protection solutions in the Product Catalogue for the tendered products that require connection to the hospital’s enterprise network. | Data/Info Required | --Please Select--
|  |  | i |  | For enhanced security, all network ports on delivered equipment must be disabled by default. The Participating Health Service's IT team is responsible for enabling only the necessary ports during the installation. | Select Compliance | --Please Select--
Data Transmission
|  |  | j |  | The offered equipment must transmit data using industry-standard communication channels, such as Cat6 Ethernet or IEEE 802.11 (Wi-Fi), and incorporate appropriate security protocols. | Select Compliance | --Please Select--
Remote Access
|  |  | k |  | Remote access to any tendered product must adhere to the following mandatory security requirements: | Select Compliance | --Please Select--
|  |  |  | i | Secure Environment: access must be conducted exclusively through the Participating Health Service's existing, approved remote access gateway.
|  |  |  | ii | Time-Limited Access: access sessions must be automatically terminated after a defined period of inactivity and shall have configurable maximum session durations.
|  |  |  | iii | End-to-End Encryption: all data transmitted during remote access sessions must be encrypted using industry-standard protocols.
|  |  |  | iv | Unique User Authentication: access must be granted only through individually assigned user accounts with strong, unique credentials. Generic or group accounts are strictly prohibited.
|  |  |  | v | Multi-Factor Authentication (MFA): all remote access attempts must be authenticated using a robust MFA protocol, incorporating at least two distinct authentication factors.
Software Updates
|  |  | l |  | The Supplier must provide all software updates and patches at no cost. | Select Compliance | --Please Select--
|  |  |  | i | Where an update introduces significant operational changes, the Supplier must deliver complementary user training. Major version upgrades are excluded from this obligation. | Select Compliance | --Please Select--
|  |  | m |  | The Supplier must obtain prior written approval from the relevant Participating Health Service before installing any software updates or upgrades. | Select Compliance | --Please Select--
|  | 2 |  |  | Consumables and Semi-Consumables
|  |  | a |  | Consumables and semi-consumables must be available as a separate purchase option, with pricing that is distinct from any starter pack that may contain those items. | Select Compliance | --Please Select--
|  |  | b |  | The Supplier must provide the Participating Health Service with a document detailing all compatible consumables, semi-consumables, and accessories for the purchased equipment, including part numbers and descriptions. | Select Compliance | --Please Select--
|  |  |  | i | The Respondent must submit a catalogue in PDF format listing all compatible consumables, semi-consumables, and accessories for the offered equipment, including part numbers and descriptions. | Data/Info Required | --Please Select--
|  |  | c |  | The reprocessing or refurbishing of a single use (or single patient) medical device, as per the TGA, will be considered as a remanufactured medical device. Remanufactured medical devices are explicitly excluded from the scope of this Invitation to Supply (ITS) and contract. The determination and subsequent decision regarding the use of such remanufactured items shall remain solely at the discretion of the Participating Health Service. | Info to Respondent | --Please Select--
|  | 3 |  |  | Substances of Concern
|  | 3.1 |  |  | Latex
|  |  | a |  | The Participating Health Service may preference products (including their accompanying packaging) that are latex-free, unless otherwise stated. | Info to Respondent | --Please Select--
|  | 3.2 |  |  | PVC Plasticisers
|  |  | a |  | For any offered consumable that contains plastic the Respondent must: | Select Compliance | --Please Select--
|  |  |  | i | list the name of plasticisers (e.g., phthalates) added to the plastic to make it flexible, and
|  |  |  | ii | the corresponding weight percentage of phthalate in the final plasticised material.
|  |  | b |  | During normal intended use, the plasticised material must not release or migrate endocrine-disrupting phthalates listed in EU REACH Regulation Annex XIV, such as DEHP, BBP, DBP, DIBP. | Select Compliance | --Please Select--
|  |  | c |  | Participating Health Services favour consumables that either do not contain phthalates listed in EU REACH Regulation Annex XIV, which are identified as having endocrine-disrupting properties, or contain phthalates at levels within the safe limits established by the EU REACH Regulation. | Info to Respondent | --Please Select--
|  | 4 |  |  | Packaging and Labelling
|  |  | a |  | Any carton/box greater than 15kg must have a "Heavy" label externally adhered on | Select Compliance | --Please Select--
|  |  | b |  | The packaging for each offered product must: | Select Compliance | --Please Select--
|  |  |  | i | be easy to open,
|  |  |  | ii | be robust as to protect the quality of offered product, and
|  |  |  | iii | if the offered product is a consumable and supplied as sterile, then the packaging must:
• be openable in a way, that the sterile integrity of offered consumable remains unaffected,
• complies with ISO 11607-1 Packaging for terminally sterilized medical devices — Part 1: Requirements for materials, sterile barrier systems and packaging systems, and
• packaged in a manner that protects the contents from contamination during transportation, storage and handling.
|  |  | c |  | Therapeutic goods must be labelled in accordance with Essential Principle 13 of Schedule 1 of the Therapeutic Goods (Medical Devices) Regulations 2002. | Select Compliance | --Please Select--
|  |  | d |  | Offered products must be delivered in accordance with the manufacturer’s instructions. | Select Compliance | --Please Select--
|  |  | e |  | Individually packaged products should indicate (where applicable): | Select Compliance | --Please Select--
|  |  |  | i | whether the product is sterile;
|  |  |  | ii | use-by date;
|  |  |  | iii | whether the product is MRI compatible (for implantable products);
|  |  |  | iv | whether the product (or packaging) contains latex or is latex-free; and
|  |  |  | v | tracking traceable identification labels.
|  | 5 |  |  | Infection Control
|  |  | a |  | Where applicable, all items must meet the requirements of the Australian Guidelines for the Prevention and Control of Infection in Healthcare, Canberra: National Health and Medical Research Council (2019). | Select Compliance | --Please Select--
|  |  |  | i | Reusable medical devices (RMD) and agents for reprocessing RMD must meet the reprocessing standard AS 5369:2023 Reprocessing of reusable medical devices and other devices in health and non-health related facilities. | Select Compliance | --Please Select--
|  |  |  | ii | Upon request by Participating Health Service, the Supplier must provide reprocessing instructions for reusable products. | Select Compliance | --Please Select--
|  |  | c |  | The Respondent must ensure that compatible disinfectants (including wipes) are available for Participating Health Services to purchase from an Australian supplier. | Select Compliance | --Please Select--
|  |  |  | i | Participating Health Services are encouraged to select disinfectants listed on HSV's HPVC2024-194 Personal Protective Equipment agreement. | Info to Respondent | --Please Select--
|  |  | d |  | Upon request by the Participating Health Service, the Supplier must provide a complete list of all recommended cleaning and disinfection products (including wipes) suitable for use on the offered equipment. This list should include the maximum permissible concentration level of each active ingredient. | Select Compliance | --Please Select--
|  |  | e |  | To facilitate cleaning, offered products must have smooth, non-porous surfaces, with | Select Compliance | --Please Select--
|  |  |  | i | no sharp edges, or hook & loop fastener (e.g. Velcro®).
|  | 6 |  |  | Occupational Health and Safety
|  |  | a |  | Mobile equipment must be designed to minimize manual handling risks for nursing staff, considering factors such as weight, dimensions, and ergonomics, to prevent musculoskeletal injuries. | Select Compliance | --Please Select--
|  |  | b |  | Any mobile equipment exceeding 7 kilograms must be offered with a suitable transportation accessory, such as a five-legged roll-stand with an integrated cable and tubing management system, or a dedicated trolley, or a bed mounting mechanism. | Select Compliance | --Please Select--
|  |  |  | i | To maintain optimal hygiene, roll-stands and trolleys should include mounts for a hand sanitiser and/or equipment disinfectant wipe dispenser. | Select Compliance | --Please Select--
|  |  | c |  | Mobile devices equipped with castors must be designed with sufficient stability and load capacity for their intended use. Effective braking mechanisms are essential to prevent accidental movement. | Select Compliance | --Please Select--
|  |  | d |  | All equipment and cases must be designed to eliminate hazardous sharp edges and pinch points. | Select Compliance | --Please Select--
|  | 6.1 |  |  | Carry Bags
|  |  | a |  | For each portable equipment, a custom protective carry bag should be offered as an optional accessory. | Select Compliance | --Please Select--
|  |  | b |  | To maintain optimal hygiene, single use disposable carry bags are permitted for portable equipment, especially for Telemetry Units. | Info to Respondent | --Please Select--
|  |  | c |  | Carry bags should feature well-organised, quickly accessible, and secure compartments for the essential supplies. The bag's construction should provide effective dust resistance to protect contents. | Select Compliance | --Please Select--
|  |  | d |  | Carry bag handles should be positioned for optimal balance and sized for a comfortable grip by a small adult. | Select Compliance | --Please Select--
|  |  | e |  | For carry bags equipped with shoulder straps should have the straps positioned for optimal balance and adequately padded for comfort. | Select Compliance | --Please Select--
|  | 6.2 |  |  | Mounting Hardware
|  |  | a |  | Optional mounting brackets must be designed and constructed to securely hold the mounted device, with a load capacity sufficient for the weight, under all intended conditions, preventing accidental detachment or movement. | Select Compliance | --Please Select--
|  |  | b |  | Mounting brackets must allow for adjustable height, tilt. Exceptions granted for MRI-Conditional equipment mounting brackets. | Select Compliance | --Please Select--
|  |  | c |  | An integrated cable management system must be included with mounting brackets. | Select Compliance | --Please Select--
|  |  | d |  | A range of mounting bracket options must be offered. | Select Compliance | --Please Select--
|  | 7 |  |  | Human Factors Engineering
|  |  | a |  | Equipment must have switches and controls with legible labels, arranged according to established human factors principles, such as clear grouping and consistent placement. | Select Compliance | --Please Select--
|  |  |  | i | Audio prompts, if available, must be easily understood. | Select Compliance | --Please Select--
|  |  | b |  | Electronic products must incorporate a self-diagnostic system that identifies internal malfunctions and provides clear visual or auditory warnings. | Select Compliance | --Please Select--
|  |  |  | i | Additional quick diagnostic cards may be provided to supplement this. | Info to Respondent | --Please Select--
|  | 7.1 |  |  | Alarms
|  |  | a |  | Equipment must include an alarm system with tiered severity levels for all monitored physiological parameters. | Select Compliance | --Please Select--
|  |  | b |  | Equipment must allow users to set default alarm settings for adult, paediatric and neonatal patients. | Select Compliance | --Please Select--
|  |  |  | i | Alarm settings must reset to default when the equipment is powered off and on. | Select Compliance | --Please Select--
|  |  | c |  | The audible alarm must remain operational and enabled at all times. Temporary silencing is permitted. | Select Compliance | --Please Select--
|  |  | d |  | Audible alarms must have adjustable volume, with a minimum level that ensures they are clearly heard above 65dB of ambient noise. | Select Compliance | --Please Select--
|  |  | e |  | When an alarm's audio alert is silenced, a visual indicator must clearly show the persistent alarm condition. This visual indicator should also include a countdown timer, after which the audio alarm will re-trigger if the condition remains unresolved. | Select Compliance | --Please Select--
|  |  | f |  | If the alarm volume is silenced, it must reactivate upon a new critical alarm. | Select Compliance | --Please Select--
|  | 7.2 |  |  | Displays
|  |  | a |  | Displays must be legible across diverse viewing angles and lighting conditions. | Select Compliance | --Please Select--
|  |  | b |  | Displays must be of a size to ensure all information is readable at the equipment's intended viewing distances. | Select Compliance | --Please Select--
|  |  | c |  | Display features should include a configurable layout, colour capability, and auto-dimming. | Select Compliance | --Please Select--
|  | 8 |  |  | Product Information
|  |  | a |  | For each equipment model offered, the Respondent must submit: | Data/Info Required | --Please Select--
|  |  |  | i | complete operator user manuals in PDF format, and
|  |  |  | ii | comprehensive product brochures in PDF format.
|  |  | b |  | For each type of consumable offered, the Respondent must submit a copy of relevant product diagrams, specifications or brochures (Product Information) to assist in accurately identifying products offered. | Data/Info Required | --Please Select--
|  |  |  | i | All consumable product information submitted should: | Select Compliance | --Please Select--
• be in electronic format
• be in English
• be specific to the product offered
• contain the Respondent’s company name
• include the product code
• include a detailed specification of the product
• include clear diagrams/pictures of the product.
|  |  |  | ii | To assist in managing this material, all product information submitted should be labelled with the relevant HSV category and subcategory number. | Select Compliance | --Please Select--
• Electronic copies should include the HSV Category and subcategory numbers in the filename or identifying metadata.
|  |  |  | iii | HSV may not consider unlabelled submissions. | Info to Respondent | --Please Select--
|  |  |  | iv | Product information will not be evaluated but is necessary to assist in accurately identifying products offered.
|  |  |  | v | HSV reserves the right to exclude products from evaluation when the product is unidentifiable and provided information is:
• not labelled as per clause b(iii) above; or
• is incomplete as to clause b(ii).
|  |  | c |  | Product samples are not to be provided unless specifically requested by HSV. | Info to Respondent | --Please Select--
|  |  | d |  | The Respondent should not submit information relating to products that are not called for in this ITS. | Info to Respondent | --Please Select--
|  |  | e |  | Research papers should not be provided unless specifically requested by HSV. | Info to Respondent | --Please Select--

## List (hidden)
--Please Select-- | --Please Select-- | --Please Select--
Fully Comply | Yes | Acknowledged
Partially Comply | No | Not Applicable
Does Not Comply | Not Applicable
Not Applicable

## 1.01&1.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 1.0 |  |  |  | Hospital Beds, Acute Care | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 1.01 & 1.02 General Ward Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | General Ward Bed
|  |  |  | i | An electrically powered device, designed with integrated electric motors and a handheld or integrated control panel that can be operated by the patient or nursing staff; features a mattress support platform that divides into three or more pivoted sections allowing for independent positioning of the head, torso, and legs; includes a headboard and footboard, and adjustable siderails that are movable and latchable. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• may accommodate a standard single mattress with nominal dimensions of 92cm x 190cm
• may accommodate a king single mattress with nominal dimensions of 107cm x 203cm
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | General ward beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following:
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  |  | b |  | Beds equipped with removable siderails should have clearly visible and durable markers indicating the correct placement of the siderails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | If bed ends are designed for use as handholds, they must withstand the forces applied during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds with ergonomic handgrips integrated into the bedframe are encouraged, for caregivers to use to transport patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Beds equipped with motorised drives are encouraged to enhance staff safety and reduce physical strain. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds equipped with motorised drives must meet the requirements of IEC 60601-2-52, clause 201.9.4.2.4.3 regarding threshold traversal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | General Ward Beds must have a safe working load of at least 250 kg, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | accommodate a bed length extension of at least 20 cm.
|  |  | b |  | The mattress support platform must be height adjustable, with the dimensions suited for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | single mattress of at least 180 cm in length and less than 100 cm in width, or a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | king single mattress of at least 200 cm in length and more than 100 cm in width. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Beds featuring integrated mattresses must also comply with all criteria within a subcategory of Category 11. The respondent must also complete one of the subcategories with Category 11 for the offered bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Respondents should offer additional options to General Ward Beds, such as anti-pinch side rails, integrated patient weighing scales, bed exit alarms, and under-bed lighting. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The mattress support platform should be capable of achieving a low position, ideally somewhere around 30 to 35 cm from the floor. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 12.01 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | Beds must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Bed extensions must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Beds must be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | An optional self-help or patient pole with a height-adjustable handpiece must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | All optional features offered, including but not limited to orthopaedic frames, monitor shelves, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connectivity and Data Integration
|  |  | h |  | Offers for beds featuring assistive technologies such as movement and position monitoring, weight monitoring, incontinence management, bed exit alarms, and motion-activated nightlights is encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Beds with integrated assistive technologies should connect with other hospital systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Centralised Bed Monitoring Systems
|  |  | j |  | Respondents offering beds equipped with assistive technologies, are encouraged to also include centralised bed monitoring systems in their product portfolio. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Centralised bed monitoring systems are out-of-scope of this ITS.
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Articulations
|  |  | a |  | All beds must be equipped with electrically adjustable profile functions, including backrest and knee break, and stepless adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer beds with features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Side Rails
|  |  | a |  | Beds equipped with continuous, full-length side rails will be excluded from consideration. Respondents must offer beds with patient-safe side rail configurations, including, but not limited to, split or other partial-length designs. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | d |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Side rails should not increase the overall width of the bed frame, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to safely and effectively support the patient during ingress and egress from the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model General Ward Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | ii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iii | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• maximum extended mattress platform dimensions for mattress extensions, Width x Length (cm)
|  |  |  | x | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for bed to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xi | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum bed width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | list all accessories that are available offered bed (self-help pole, IV pole, integrated mattress)
|  |  |  | xvi | list all options available for offered bed
• fifth wheel (Yes/No)
• motorised drive (Yes/No)
• list all others
|  |  |  | xvii | assistive technologies (smart beds)
• list all installed smart features (scales, falls detection, incontinence monitoring, motion-activated lighting, etc.)
• bed includes connectivity features for Nurse Call, IT network, or other systems (Yes/No)
- which system(s)
- data transmitted over (ethernet, wi-fi, other network connections)
• Centralised Bed Monitoring System available from Supplier (Yes/No)
|  |  |  | xviii | state the size of the display, H x W (cm) where applicable
|  |  |  | xix | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xx | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xxi | state the 'Drop Test' height for handset
|  |  |  | xxii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | warranty period offered on bed structural frame (months)
|  |  |  | xxv | warranty period offered on bed electronics (months)
|  |  |  | xxvi | warranty period offered on bed handset & cables (months)
|  |  |  | xxvii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxviii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxix | Estimated lifespan of device (years of service)
|  |  |  | xxx | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 1.03&1.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 1.0 |  |  |  | Hospital Beds, Acute Care | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 1.03 & 1.04 General Ward High-Low Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | General Ward High-Low Bed
|  |  |  | i | A dedicated type of hospital electric bed designed with a significantly lower frame height compared to standard hospital beds. The reduced height is a key safety feature, intended to minimise the risk and severity of falls, particularly for the elderly, frail, or cognitively impaired patients who are prone to rolling out of bed. Despite the beds low profile, it still retains most of the functionalities found in a standard hospital bed. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• may accommodate a standard single mattress with nominal dimensions of 92cm x 190cm
• may accommodate a king single mattress with nominal dimensions of 107cm x 203cm
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | General ward low beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  |  | b |  | Beds equipped with removable siderails should have clearly visible and durable markers indicating the correct placement of the siderails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | If bed ends are designed for use as handholds, they must withstand the forces applied during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds with ergonomic handgrips integrated into the bedframe are encouraged, for caregivers to use to transport patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Beds equipped with motorised drives are encouraged to enhance staff safety and reduce physical strain. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds equipped with motorised drives must meet the requirements of IEC 60601-2-52, clause 201.9.4.2.4.3 regarding threshold traversal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | General Ward Low Beds must have a safe working load of at least 250 kg. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The mattress support platform must have an adjustable height ranging from at least 20 to 65 cm and be capable of accommodating either a single mattress (minimum 180 cm length, maximum 100 cm width) or a king single mattress (minimum 200 cm length, greater than 100 cm width). | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents should offer additional options to General Ward Low Beds, such as anti-pinch side rails, bed extension. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for Ultra Low model beds are encouraged, those designed where the mattress platform rests very close to the floor. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking and steering system must be single-foot operable, accessible from both side of the bed, and clearly labelled indicating its function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 12.01 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | Beds must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Bed extensions must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | For Ultra-Low beds, side rails must be available as an option | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Beds must be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | An optional self-help or patient pole with a height-adjustable handpiece must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | All optional features offered, including but not limited to orthopaedic frames, monitor shelves, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connectivity and Data Integration
|  |  | i |  | Offers for beds featuring assistive technologies such as movement and position monitoring, weight monitoring, incontinence management, bed exit alarms, and motion-activated nightlights is encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Beds with integrated assistive technologies should connect with other hospital systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Centralised Bed Monitoring Systems
|  |  | k |  | Respondents offering beds equipped with assistive technologies, are encouraged to also include centralised bed monitoring systems in their product portfolio. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Centralised bed monitoring systems are out-of-scope of this ITS.
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Articulations
|  |  | a |  | All beds must be equipped with electrically adjustable profile functions, including backrest and stepless adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Electrically powered knee articulation is encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer beds with features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Side Rails
|  |  | a |  | Beds equipped with continuous, full-length side rails will be excluded from consideration. Respondents must offer beds with patient-safe side rail configurations, including, but not limited to, split or other partial-length designs. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Side rails must comply with IEC 60601-2-52, and for Ultra-Low beds, be removable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | d |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Side rails should not increase the overall width of the bed frame, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to safely and effectively support the patient during ingress and egress from the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model General Ward Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress platform dimensions (without extensions), Width x Length (cm)
• maximum extended mattress platform dimensions for mattress extensions, Width x Length (cm)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for bed to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xii | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum bed width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual, etc.)
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe the brake mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed (self-help pole, IV pole, integrated mattress)
|  |  |  | xvii | list all options that are available for offered bed
|  |  |  | xviii | assistive technologies (smart beds)
• list all installed smart features (scales, falls detection, incontinence monitoring, motion-activated lighting, etc.)
• bed includes connectivity features for Nurse Call, IT network, or other systems (Yes/No)
- which system(s)
- data transmitted over (ethernet, wi-fi, other network connections)
• Centralised Bed Monitoring System available from Supplier (Yes/No)
|  |  |  | xix | state the size of the display, H x W (cm) where applicable
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | warranty period offered on bed structural frame (months)
|  |  |  | xxvi | warranty period offered on bed electronics (months)
|  |  |  | xxvii | warranty period offered on bed handset & cables (months)
|  |  |  | xxviii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxix | Estimated device's end-of-sales life (state the year)
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
|  |  |  | xxxii | • can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxiii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 1.05
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 1.0 |  |  |  | Hospital Beds, Acute Care | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 1.05 Bariatric Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Bariatric Bed
|  |  |  | i | An electrically powered bed specifically designed to accommodate a bariatric patient whose body mass exceeds the safe working load of standard hospital beds. These beds typically include a sturdier and wider frame than standard beds, giving extra support and facilitating the ingress and egress of patients. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• may also be a high-low configuration bed
• may also be an ultra-low configuration bed
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Bariatric beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  |  | b |  | Beds equipped with removable siderails should have clearly visible and durable markers indicating the correct placement of the siderails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | If bed ends are designed for use as handholds, they must withstand the forces applied during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds with ergonomic handgrips integrated into the bedframe are encouraged, for caregivers to use to transport patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Beds equipped with motorised drives are encouraged to enhance staff safety and reduce physical strain. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds equipped with motorised drives must meet the requirements of IEC 60601-2-52, clause 201.9.4.2.4.3 regarding threshold traversal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Bariatric Beds must have a safe working load that significantly exceeds 250kg, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | accommodate a bed length extension of at least 20 cm.
|  |  | b |  | The mattress support platform must be electrically height adjustable and should have a minimal operational width of 107cm or more. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Bariatric beds should be offered with: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | adjustable mattress platform width
|  |  |  | ii | motorised drive
|  |  |  | iii | dynamic variable bed width mattress
|  |  |  | iv | auto-regression, preventing patient slide when backrest raised
|  |  |  | v | integrated weigh scales
|  |  | d |  | Beds featuring integrated mattresses must also comply with all criteria within a subcategory of Category 11. The respondent must also complete one of the subcategories with Category 11 for the offered bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Respondents should offer additional options to Bariatric Beds, such as anti-pinch side rails, motorised drive, and integrated patient weighing scales. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 12.01 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | Beds must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Bed extensions must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Beds must be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | An optional self-help or patient pole with a height-adjustable handpiece must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | All optional features offered, including but not limited to fifth wheel, monitor shelves, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connectivity and Data Integration
|  |  | h |  | Offers for beds featuring assistive technologies such as movement and position monitoring, weight monitoring, incontinence management, bed exit alarms, and motion-activated nightlights is encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Beds with integrated assistive technologies should connect with other hospital systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Centralised Bed Monitoring Systems
|  |  | j |  | Respondents offering beds equipped with assistive technologies, are encouraged to also include centralised bed monitoring systems in their product portfolio. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Centralised bed monitoring systems are out-of-scope of this ITS.
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Articulations
|  |  | a |  | All beds must be equipped with electrically adjustable profile functions, including backrest and knee break, and stepless adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer beds with features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Side Rails
|  |  | a |  | Beds equipped with continuous, full-length side rails will be excluded from consideration. Respondents should offer beds with patient-safe side rail configurations, including, but not limited to, split or other partial-length designs. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Side rails should not increase the overall width of the bed frame, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to safely and effectively support the patient during ingress and egress from the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model General Ward Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress platform dimensions, the range in Length (cm)
• mattress platform dimensions, the range in Width (cm)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for bed to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xii | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum bed width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed (self-help pole, IV pole, integrated mattress)
|  |  |  | xvii | list all options available for offered bed
• motorised drive (Yes/No)
• adjustable width (Yes/No)
• integrated weigh scales (Yes/No)
• integrated mattress (Yes/No)
- state whether foam/air/hybrid and describe mattress features
• list all others
|  |  |  | xviii | assistive technologies (smart beds)
• list all installed smart features (scales, falls detection, incontinence monitoring, motion-activated lighting, etc.)
• bed includes connectivity features for Nurse Call, IT network, or other systems (Yes/No)
- which system(s)
- data transmitted over (ethernet, wi-fi, other network connections)
• Centralised Bed Monitoring System available from Supplier (Yes/No)
|  |  |  | xix | state the size of the display, H x W (cm) where applicable
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | warranty period offered on bed structural frame (months)
|  |  |  | xxvi | warranty period offered on bed electronics (months)
|  |  |  | xxvii | warranty period offered on bed handset & cables (months)
|  |  |  | xxviii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxix | Estimated device's end-of-sales life (state the year)
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 1.06
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 1.0 |  |  |  | Hospital Beds, Acute Care | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 1.06 Critical Care Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Critical Care Beds
|  |  |  | i | A critical care bed is a highly specialised and technologically advanced bed designed for patients in the intensive care unit (ICU) who require continuous monitoring and life support. Unlike standard hospital beds, an intensive care bed is equipped with a multitude of integrated features and functionalities crucial for managing critically ill individuals. These beds often include built-in weight scales for precise medication dosing and fluid balance monitoring, advanced pressure redistribution surfaces to prevent pressure ulcers in immobile patients, and electronic controls for a wide range of positional adjustments, including Trendelenburg, reverse Trendelenburg, and cardiac chair positions, to optimize respiratory and circulatory function. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Critical Care Beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  |  | b |  | Beds equipped with removable siderails should have clearly visible and durable markers indicating the correct placement of the siderails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | If bed ends are designed for use as handholds, they must withstand the forces applied during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds with ergonomic handgrips integrated into the bedframe are encouraged, for caregivers to use to transport patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Beds equipped with motorised drives are encouraged to enhance staff safety and reduce physical strain. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds equipped with motorised drives must meet the requirements of IEC 60601-2-52, clause 201.9.4.2.4.3 regarding threshold traversal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Critical Care Beds must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 250 kg
|  |  |  | ii | accommodate a bed length extension of at least 20 cm
|  |  |  | iii | central locking mechanism with directional steering capability
|  |  |  | iv | integrated head angle indicator
|  |  |  | v | integrated patient weighing scales
|  |  |  | iv | accessory support for 2 IV poles
|  |  | b |  | The mattress support platform must be height adjustable and designed to accommodate mattresses with a minimum length of 180 cm and a minimum width of 88 cm. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed frame or mattress support must be designed to enable the placement and operation of an x-ray cassette under the patient. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Patient and caregiver bed controls must be integrated within the bed side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Beds featuring integrated mattresses must also comply with all criteria within a subcategory of Category 11. The respondent must also complete one of the subcategories with Category 11 for offered bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Respondents should offer additional options to Critical Care Beds, such as integrated mattress with rotational support and percussion therapy, anti-pinch side rails, touchscreen controls, integrated patient weighing scales, motorised drive, cardiac chair position, bed exit alarms, and under-bed lighting. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | The mattress support platform should be capable of achieving a low position, ideally somewhere around 30 to 35 cm from the floor. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 12.01 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | Beds must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Bed extensions must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Beds must be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | All optional features offered, including but not limited to orthopaedic frames, monitor shelves, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connectivity and Data Integration
|  |  | g |  | Offers for beds featuring assistive technologies such as movement and position monitoring, weight monitoring, incontinence management, bed exit alarms, and motion-activated nightlights is encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Beds with integrated assistive technologies should connect with other hospital systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Centralised Bed Monitoring Systems
|  |  | i |  | Respondents offering beds equipped with assistive technologies, are encouraged to also include centralised bed monitoring systems in their product portfolio. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Centralised bed monitoring systems are out-of-scope of this ITS.
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Articulations
|  |  | a |  | All beds must be equipped with electrically adjustable profile functions, including backrest and knee break, and stepless adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer beds with features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Side Rails
|  |  | a |  | Beds equipped with continuous, full-length side rails will be excluded from consideration. Respondents must offer beds with patient-safe side rail configurations, including, but not limited to, split or other partial-length designs. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Side rails should not increase the overall width of the bed frame, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to safely and effectively support the patient during ingress and egress from the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model General Ward Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress platform dimensions (without extensions), Width x Length (cm)
• maximum extended mattress platform dimensions for mattress extensions, Width x Length (cm)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for bed to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xii | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum bed width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed
• orthopaedic traction frame (Yes/No)
• transport platform (Yes/No)
• list all others (self-help pole, IV pole)
|  |  |  | xvii | list all options available for offered bed
• integrated mattress (Yes/No)
- state whether foam/air/hybrid and describe mattress features
• fifth wheel (Yes/No)
• motorised drive (Yes/No)
• list all others
|  |  |  | xviii | assistive technologies (smart beds)
• list all installed smart features (scales, falls detection, incontinence monitoring, motion-activated lighting, etc.)
• bed includes connectivity features for Nurse Call, IT network, or other systems (Yes/No)
- which system(s)
- data transmitted over (ethernet, wi-fi, other network connections)
• Centralised Bed Monitoring System available from Supplier (Yes/No)
|  |  |  | xix | state the size of the display, H x W (cm) where applicable
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered equipment
|  |  |  | xxii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | warranty period offered on bed structural frame (months)
|  |  |  | xxv | warranty period offered on bed electronics (months)
|  |  |  | xxvi | warranty period offered on bed handset & cables (months)
|  |  |  | xxvii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxviii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxix | Estimated lifespan of device (years of service)
|  |  |  | xxx | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 1.07
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 1.0 |  |  |  | Hospital Beds, Acute Care | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 1.07 Stand-Up Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Stand-Up Bed
|  |  |  | i | A motorised, mains-powered standing bed assisting individuals with disabilities or bedridden patients to attain a supported standing posture. This functionality also enables patients with significant back issues to easily get in and out of bed without bending. The bed's ability to move from a horizontal to a vertical position aids in preventing adverse effects like pressure sores, hip, knee, and ankle contractures, and the deterioration of cardio-circulatory and respiratory health. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stand-Up Beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  |  | b |  | Beds equipped with removable siderails should have clearly visible and durable markers indicating the correct placement of the siderails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | If bed ends are designed for use as handholds, they must withstand the forces applied during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds with ergonomic handgrips integrated into the bedframe are encouraged, for caregivers to use to transport patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Beds equipped with motorised drives are encouraged to enhance staff safety and reduce physical strain. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds equipped with motorised drives must meet the requirements of IEC 60601-2-52, clause 201.9.4.2.4.3 regarding threshold traversal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stand-Up Beds must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 250 kg
|  |  |  | ii | integrated mattress
|  |  |  | iii | central locking mechanism with directional steering capability
|  |  |  | iv | retractable castors when bed in upright position
|  |  |  | v | integrated inclinometer
|  |  |  | vi | integrated patient weighing scales
|  |  | b |  | The mattress support platform must be height adjustable and designed to accommodate mattresses with a minimum length of 180 cm and a minimum width of 88 cm. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed frame or mattress support must be designed to enable the placement and operation of an x-ray cassette under the patient. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Patient and caregiver bed controls must be integrated within the bed side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Beds featuring integrated mattresses must also comply with all criteria within a subcategory of Category 11. The respondent must also complete one of the subcategories with Category 11 for offered bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Offers for bariatric versions are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 12.01 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | Beds must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Bed extensions must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Beds must be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | An optional self-help or patient pole with a height-adjustable handpiece must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | All optional features offered, including but not limited to orthopaedic frames, monitor shelves, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connectivity and Data Integration
|  |  | h |  | Offers for beds featuring assistive technologies such as movement and position monitoring, weight monitoring, incontinence management, bed exit alarms, and motion-activated nightlights is encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Beds with integrated assistive technologies should connect with other hospital systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Centralised Bed Monitoring Systems
|  |  | j |  | Respondents offering beds equipped with assistive technologies, are encouraged to also include centralised bed monitoring systems in their product portfolio. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Centralised bed monitoring systems are out-of-scope of this ITS.
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Articulations
|  |  | a |  | All beds must be equipped with electrically adjustable profile functions, including backrest and knee break, and stepless adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Side Rails
|  |  | a |  | Beds equipped with continuous, full-length side rails will be excluded from consideration. Respondents must offer beds with patient-safe side rail configurations, including, but not limited to, split or other partial-length designs. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Side rails should not increase the overall width of the bed frame, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to safely and effectively support the patient during ingress and egress from the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model General Ward Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress platform dimensions (without extensions), Width x Length (cm)
• maximum extended mattress platform dimensions for mattress extensions, Width x Length (cm)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for bed to return to horizontal plane from its full vertical position upon activation of the CPR release mechanism (sec)
|  |  |  | xii | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum bed width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed (self-help pole, IV pole, integrated mattress)
|  |  |  | xvii | list all options available for offered bed
• fifth wheel (Yes/No)
• motorised drive (Yes/No)
• fall prevention detection (Yes/No)
• list all others
|  |  |  | xviii | assistive technologies (smart beds)
• list all installed smart features (scales, falls detection, incontinence monitoring, motion-activated lighting, etc.)
• bed includes connectivity features for Nurse Call, IT network, or other systems (Yes/No)
- which system(s)
- data transmitted over (ethernet, wi-fi, other network connections)
• Centralised Bed Monitoring System available from Supplier (Yes/No)
|  |  |  | xix | state the size of the display, H x W (cm) where applicable
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | warranty period offered on bed structural frame (months)
|  |  |  | xxvi | warranty period offered on bed electronics (months)
|  |  |  | xxvii | warranty period offered on bed handset & cables (months)
|  |  |  | xxviii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxix | Estimated device's end-of-sales life (state the year)
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 1.08
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 1.0 |  |  |  | Hospital Beds, Acute Care | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 1.08 Homecare Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Homecare Bed
|  |  |  | i | An electrically powered bed designed for easy transportation, manoeuvrability, and rapid assembly within a home setting. Once set up, it offers adjustable height and independent back and leg sections to aid caregivers in providing care. Designed with both safety and mobility as priorities, the bed is intended for in-home patient care. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Homecare Beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, should be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end should possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Beds offered may come in a range of sizes, for example, single, king single, double, and more. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Homecare Beds must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 250 kg
|  |  |  | ii | easily transported into the home
|  |  |  | iii | dismantlable for manoeuvring through doorways and hallways
|  |  |  | iv | straightforward to assemble
|  |  |  | v | castors with locking device and directional steering
|  |  | c |  | The mattress support platform must be electrically height adjustable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Respondents should offer additional options, such as anti-pinch side rails, and self help pole. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 12.01 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | side rails must be available as an option | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Beds should be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | An optional self-help or patient pole with a height-adjustable handpiece must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Articulations
|  |  | a |  | All beds must be equipped with electrically adjustable profile functions, including backrest and knee break. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer beds with features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Side Rails
|  |  | a |  | Beds equipped with continuous, full-length side rails will be excluded from consideration. Respondents should offer beds with patient-safe side rail configurations, including, but not limited to, split or other partial-length designs. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Side rails should not increase the overall width of the bed frame, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to safely and effectively support the patient during ingress and egress from the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | CPR Release
|  |  | a |  | Where a CPR mechanism is fitted, the activation of the CPR release must trigger a rapid and safe return of the backrest to a horizontal position. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds that have a CPR release mechanism, the CPR mechanism must be clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model General Ward Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | variable height range of the mattress platform from floor level (cm)
|  |  |  | x | mattress platform dimensions, Width x Length (cm)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for bed to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xii | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum bed width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed (self-help pole, IV pole, transporting bracket)
|  |  |  | xvii | list all options available for offered bed
• fall prevention detection (Yes/No)
• list all others
|  |  |  | xviii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xix | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xx | state the 'Drop Test' height for handset
|  |  |  | xxi | manufacturer of offered equipment
• country of origin
|  |  |  | xxii | Year 1st manufactured
|  |  |  | xxiii | warranty period offered on bed structural frame (months)
|  |  |  | xxiv | warranty period offered on bed electronics (months)
|  |  |  | xxv | warranty period offered on bed handset & cables (months)
|  |  |  | xxvi | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxvii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxviii | Estimated lifespan of device (years of service)
|  |  |  | xxix | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxx | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 2.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 2.0 |  |  |  | Hospital Beds, Mental Health | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 2.01 Behavioural Healthcare Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Behavioural Healthcare Beds
|  |  |  | i | A bed specifically designed for the demanding environment of a mental healthcare unit, ensuring a high level of safety and security for both patients and staff without compromising patient comfort. Constructed with robust, tamper-resistant materials and ligature-resistant features, to minimise the potential for self-harm and accidental injury. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• may accommodate a standard single mattress with nominal dimensions of 92cm x 190cm
• may accommodate a king single mattress with nominal dimensions of 107cm x 203cm
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Behavioural Healthcare Beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Behavioural Healthcare Beds must have a safe working load of at least 200 kg, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | fitted with tamper-resistant fasteners
|  |  |  | ii | concealed wiring
|  |  |  | iii | unused accessory support mount holes covered
|  |  | b |  | The mattress support platform must be electrically height adjustable, with the dimensions suited for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | single mattress of at least 180 cm in length and less than 100 cm in width, or a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | king single mattress of at least 200 cm in length and more than 100 cm in width. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents should offer additional options to Behavioural Healthcare Beds, such as bed exit alarms, restraint attachment points, bed floor mounts | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  | b |  | Gaps in the mattress platform sections must be minimised or covered to prevent the creation of ligature points | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Mattress retainers must not present ligature points. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed must incorporate a tamper-resistant braking mechanism designed to be entirely ligature-proof. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Directional steering castors are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 2.04 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The bed should include secure attachment points to facilitate the safe application of restraints when clinically necessary. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connectivity and Data Integration
|  |  | d |  | Offers for beds featuring assistive technologies such as movement and position monitoring, weight monitoring, incontinence management, bed exit alarms, and motion-activated nightlights is encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Beds with integrated assistive technologies should connect with other hospital systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Centralised Bed Monitoring Systems
|  |  | f |  | Respondents offering beds equipped with assistive technologies, are encouraged to also include centralised bed monitoring systems in their product portfolio. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Centralised bed monitoring systems are out-of-scope of this ITS. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables (where applicable) must exhibit high crush resistance to withstand repeated bed castor loading, short no longer than 1.3 m, and be equipped with a liquid-tight seal at the equipment entry point. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Beds may powered by a portable stand-alone battery supply. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Articulations
|  |  | a |  | Offers for beds with electric backrest and knee break functionality are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer beds with features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | CPR Release
|  |  | a |  | Where a CPR mechanism is fitted, the activation of the CPR release must trigger a rapid and safe return of the backrest to a horizontal position. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds that have a CPR release mechanism, the CPR mechanism must be clearly identifiable to staff and inaccessible to patients. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends, if available, must be permanently fixed and require a tool for removal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must have solid panel infills. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Behavioural Healthcare Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Behavioural Healthcare Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | mains powered bed (Yes/No)
|  |  |  | x | remote portable battery pack operates bed (Yes/No)
|  |  |  | xi | total length of mains power cable (m)
|  |  |  | xii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | xiii | can a mattress be secured to support platform? If so, what methods are used?
|  |  |  | xiv | mattress platform dimensions (without extensions), Width x Length (cm)
|  |  |  | xv | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
|  |  |  | xvi | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xvii | describe locking mechanism (central, each caster, pedal, indicator)
• how is the bed immobilised for patient use?
|  |  |  | xviii | can bed controls (handset) be detached from patient use?
|  |  |  | xix | list all accessories that are available offered bed
|  |  |  | xx | list all options available for offered bed
• restraint attachment points (Yes/No)
- position of restraint attachment points
• fall prevention detection (Yes/No)
• list all others
|  |  |  | xxi | assistive technologies (smart beds)
• list all installed smart features (scales, falls detection, incontinence monitoring, motion-activated lighting, etc.)
• bed includes connectivity features for Nurse Call, IT network, or other systems (Yes/No)
- which system(s)
- data transmitted over (ethernet, wi-fi, other network connections)
• Centralised Bed Monitoring System available from Supplier (Yes/No)
|  |  |  | xxii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxiii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xxiv | state the 'Drop Test' height for handset
|  |  |  | xxv | manufacturer of offered equipment
• country of origin
|  |  |  | xxvi | Year 1st manufactured
|  |  |  | xxvii | warranty period offered on bed structural frame (months)
|  |  |  | xxviii | warranty period offered on bed electronics (months)
|  |  |  | xxix | warranty period offered on bed handset & cables (months)
|  |  |  | xxx | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxi | Estimated device's end-of-sales life (state the year)
|  |  |  | xxxii | Estimated lifespan of device (years of service)
|  |  |  | xxxiii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxiv | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 2.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 2.0 |  |  |  | Hospital Beds, Mental Health | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 2.02 Seclusion Room Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Seclusion Room Beds
|  |  |  | i | A bed specifically designed for seclusion rooms to prioritise patient safety and minimise the risk of self-harm or harm to others. The bed features robust, tamper-resistant materials, a streamlined design with smooth, easy-to-clean surfaces, and minimised potential ligature points. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Seclusion Room Beds offered are encouraged to be listed on the ARTG. Participating Health Services may preference ARTG-listed beds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The bed must be stable and designed to resist lifting and prevent its use as a barricade or projectile. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Seclusion room beds must not have any detachable components, which could be used for self-harm or as projectiles, and have a safe working load of at least 200kg. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Seclusion room beds can be either: | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | entirely foam-based, or
|  |  |  | ii | built with a heavy-duty frame, that can be fixed upon installation.
|  | 3.1 |  |  | Options & Accessories
|  |  | a |  | The bed can include secure attachment points to facilitate the safe application of restraints when clinically necessary. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Optional integrated load sensors are encouraged to monitor patient presence of movement, providing alerts to staff if patient leaves the bed unexpectedly. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | integrated part of the bed frame,
|  |  |  | ii | without sharp edges or gaps
|  | 3.3 |  |  | Mattress
|  |  | a |  | The bed/mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  |  | v | sufficiently firm and supportive to enable the effective administration of CPR
|  |  |  | vi | free of CFCs
|  |  |  | vii | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | viii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  | b |  | Mattresses prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Cover / Mattress Cover
|  |  | a |  | The cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | static dissipative
|  |  |  | viii | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | Mattress covers should: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be treated with an antimicrobial agent
|  |  |  | ii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  | 3.5 |  |  | Pressure Distribution
|  |  | a |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Seclusion Room Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Seclusion Room Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number, if registered with TGA | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | restraint attachment points (Yes/No)
• position of restraint attachment points
|  |  |  | ix | height from floor level of the mattress (cm)
|  |  |  | x | mattress dimensions, Width x Length x Height (cm)
• list all materials (foam types) used in the construction of the mattress core
• describe the antimicrobial treatment applied to the mattress core
• detail the microclimate management technologies integrated into the mattress
• cover moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the cover
|  |  |  | xi | list all options available for offered bed
• restraint attachment points (Yes/No)
- position of restraint attachment points
• integrated load sensor (Yes/No)
• list all others
|  |  |  | xii | manufacturer of offered equipment
• country of origin
|  |  |  | xiii | warranty period offered (months)
|  |  |  | xiv | Year 1st manufactured
|  |  |  | xv | Estimated device's end-of-sales life (state the year)
|  |  |  | xvi | Estimated lifespan of device (years of service)
|  |  |  | xvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 2.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 2.0 |  |  |  | Hospital Beds, Mental Health | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 2.03 Behavioural Healthcare Mattresses
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Behavioural Healthcare Mattress
|  |  |  | i | A mattress designed for use in an acute psychiatric setting prioritising patient safety and durability, consisting of a supportive foam core encased in a durable, waterproof, tamper-proof, seamless mattress cover. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Behavioural Healthcare Mattresses offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | date of manufacture
|  |  |  | iii | model designation
|  |  |  | iv | batch or serial number
|  |  |  | v | Safe Working Load (SWL) expressed in kilograms
|  |  |  | vi | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | vii | correct orientation and positioning of the support surface
|  |  | b |  | Participating Health Services must be permitted to write the facility ID on mattresses without voiding the warranty. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The mattress cover should include a space to write the date of purchase. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Unless contradicted by the specific requirements in clause 3.0(b), Behavioural Healthcare Mattresses must comply with all mandatory criteria within a subcategory of Category 10. Respondents are also required to complete one of the subcategories within Category 10 for their offered mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The following requirements apply to Behavioural Healthcare mattresses: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress covers must be highly cut-resistant
|  |  |  | ii | mattress covers must be zipper-free
|  |  |  | iii | mattress covers must be button-free
|  |  |  | iv | mattress covers must be fire retardant
|  |  | f |  | Offers for bariatric mattresses are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Behavioural Healthcare Mattress offered, the Respondent must submit:
|  |  |  | i | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Behavioural Healthcare Mattress offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s for both mattress core and mattress cover
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered product
• detergents that are compatible to use on the offered product
|  |  |  | v | weight of the offered mattress (kg)
|  |  |  | vi | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | vii | the external dimensions of the offered mattress, Width x Length x Height (cm)
• list all materials (foam types) used in the construction of the mattress core
• describe the mattress core (sidewalls, layers, compartments, etc.)
• describe the antimicrobial treatment applied to the mattress core
• detail the microclimate management technologies integrated into the mattress
|  |  |  | viii | mattress width uniform from end to end (Yes/No)
• if No, describe width changes
|  |  |  | ix | mattress uniform in height and density across its surface (Yes/No)
• if No, what are the differences or inconsistencies?
|  |  |  | x | castellated cells (Yes/No)
• size of castellation, Width x Length (cm)
|  |  |  | xi | hinged for backrest (Yes/No)
|  |  |  | xii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
|  |  |  | xiii | mattress cover includes handles (Yes/No)
|  |  |  | xiv | manufacturer of offered equipment
• country of origin
|  |  |  | xv | Year 1st manufactured
|  |  |  | xvi | Estimated device's end-of-sales life (state the year)
|  |  |  | xvii | warranty period of mattress (months)
|  |  |  | xviii | warranty period of mattress cover (months)
|  |  |  | xix | Estimated expected service life of mattress (years of service)
|  |  |  | xx | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 3.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 3.0 |  |  |  | Neonate & Paediatric Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 3.01 Bassinets
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Bassinet
|  |  |  | i | A bassinet is a portable sleeping enclosure with low walls, designed to provide a safe and comfortable resting space for newborns. Often featuring a clear plastic basket that fits securely into a lightweight frame with smooth-rolling casters for easy movement, the frame may also incorporate storage solutions such as shelves, cabinets, or drawers. Some bassinets offer the added convenience of electric height adjustment. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• may accommodate single infant bassinet with nominal dimensions of 42cm x 60cm
- single bassinet having a nominal 10kg SWL
• may accommodate twin infant bassinet with nominal dimensions of 68cm x 60cm
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Bassinets offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The bassinet mattress must comply with the performance requirements specified in AS/NZS 8811.1 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment offered must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | maximum patient weight capacity expressed in kilograms
|  |  |  | iii | manufacturers name
|  |  |  | iv | model designation
|  |  |  | v | serial number
|  |  |  | vi | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The offered equipment must be: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | free of protrusions that could damage the mattress or cause injury
|  |  |  | ii | designed to reduce the risk of tripping or falling.
|  |  |  | iii | constructed of materials durable and safe to neonates
|  | 3 |  |  | Product Requirements
|  |  | a |  | The bassinet frame must be sturdy and stable, designed to prevent tipping under all conditions of normal use and when exposed to common accidental forces. With the exception of stainless-steel components, all other metal surfaces are required to have a powder-coated finish that provides robust and enduring protection against corrosion. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The bassinet must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be transparent, shatterproof material, allowing for clear visibility of the neonate from all angles
|  |  |  | ii | have adequate ventilation to ensure proper air circulation around the neonate
|  |  |  | iii | easily and securely removable from the frame
|  |  |  | iv | fixed side walls with an internal depth of at least 25 cm from the top of the uncompressed mattress to the top of the sidewall
|  |  | c |  | Electric bassinets equipped with height-adjustable frames may feature a retractable or fold-down sides allowing for easy parental access to soothe their baby from the comfort of their own bed.
|  | 2.4 |  |  | Mattress
|  |  | a |  | The bassinet mattress must provide a firm and supportive sleep surface. It shall be manufactured using high-density polyurethane foam with a density of no less than 20 kg/m³. The foam should be certified low-VOC. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The mattress supplied with or intended for the bassinet must provide a firm, snug fit. When the mattress is pushed firmly to one side of the bassinet, any resulting gap on the opposite side must not exceed 2cm; this measurement applies to all sides. Furthermore, there should be no hazardous gaps created between the mattress and any part of the bassinet that could lead to entrapment. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Mattress cover must be: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable, waterproof, breathable, antistatic, flame retardant, and hypoallergenic
|  |  |  | ii | seams that are welded or heat sealed
|  |  |  | iii | cleanable with a hospital grade disinfection product
|  |  |  | iv | easily removable
Pressure Distribution
|  |  | d |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 2.4 |  |  | Castors and Braking Systems
|  |  | a |  | Bassinet must have fully enclosed castors that are non marking and anti-static. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | At least 2 casters must be fitted with easily accessible brakes. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 3.08 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following bassinet options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | trays or cabinetry
|  |  |  | ii | curtain supports / arrangements
|  |  |  | iii | ability to attach hand gel bracket
|  |  |  | iv | chart holder or designated area for patient identification
|  |  |  | v | provision for an IV pole attachment
|  | 2.3 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.x |  |  | Bassinet Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Bassinet offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Bassinet offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with (if electric)
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bassinet, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bassinet (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress dimensions, Width x Length x Height (cm)
• list all materials (foam types) used in the construction of the mattress core
• describe the antimicrobial treatment applied to the mattress core
• detail the microclimate management technologies integrated into the mattress
|  |  |  | xi | moisture vapour transmission rate (MVTR) of mattress cover (g/m2/day)
|  |  |  | xii | height of side wall (cm)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | electric bassinet (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bassinet
|  |  |  | xvii | list all options available for offered bassinet
|  |  |  | xviii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xix | state the Ingress Protection (IP) rating for both solids and liquids (if electrical)
|  |  |  | xx | manufacturer of offered equipment
• country of origin
|  |  |  | xxi | Year 1st manufactured
|  |  |  | xxii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiii | warranty period of bassinet (months)
|  |  |  | xxiv | warranty period of mattress (months)
|  |  |  | xxv | warranty period of mattress cover (months)
|  |  |  | xxvi | Estimated lifespan of bassinet (years of service)
|  |  |  | xxvii | Estimated lifespan of mattress (years of service)
|  |  |  | xxviii | Estimated lifespan of mattress cover (years of service)
|  |  |  | xxix | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 3.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 3.0 |  |  |  | Neonate & Paediatric Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 3.02 Cots
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Cot
|  |  |  | i | A specialised bed designed to ensure a safe, secure, and hygienic sleeping and resting environment for infants and young children who have not yet developed the ability to safely exit a standard bed independently, typically used from early infancy through toddlerhood. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• nominal mattress dimensions of 70cm x 140cm
• nominal 50kg SWL
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Cots offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cot mattress must comply with the performance requirements specified in AS/NZS 8811.1 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment offered must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The cot frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the cot's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cot end should possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | If cot ends are designed for use as handholds, they must withstand the forces applied during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Cots with ergonomic handgrips integrated into the cot frame are encouraged, for caregivers to use to transport patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Cots designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the cot. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | For improved usability, cot controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Participating Health Services will prioritise cots that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Cots must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 30 kg
|  |  |  | ii | adjustments for Trendelenburg and reverse Trendelenburg positioning
|  |  |  | iii | safety catches for the side rails that resist tampering or operation by a child
|  |  |  | iv | cot sides and ends that prevent climbing by the infant.
|  |  |  | v | side rails/bars appropriately spaced to prevent a child from entrapping their head or limbs
|  |  |  | vi | gaps between mattress and cot sides or ends appropriately spaced to prevent a child getting their fingers or limbs trapped
|  |  |  | vii | no holes, openings, or spaces that could trap a child's fingers or limbs
|  |  | b |  | Cots designated for NICU/PICU clinical environments must be equipped with appropriately placed and sized access ports to support ventilation tubing, physiological monitoring leads and cables, and other critical care interventions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Cots should be equipped with the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | height-adjustable mattress support platform
|  |  |  | ii | secure attachment points for an IV pole
|  |  |  | iii | compatibility for orthopaedic traction equipment
|  |  |  | iv | removable head and footboards (cot ends) for emergency access, such as for CPR
|  |  |  | v | end panels firmly attach to cot frame to prevent accidental dislodgement
|  | 3.1 |  |  | Mattress
|  |  | a |  | The cot mattress must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 30 kg
|  |  |  | ii | Deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  |  | iii | sufficiently firm and supportive to enable the effective administration of CPR
|  |  |  | iv | free of CFCs
|  |  |  | v | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | vi | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
Labelling
|  |  | b |  | Mattresses should be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | serial number
|  |  |  | iii | date of manufacture
|  |  | c |  | Participating Health Services must be permitted to write on mattresses without voiding the warranty. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The mattress cover should include a space to write the date of purchase. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mattress Cover
|  |  | e |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress core during all bed articulations
|  |  |  | viii | anti-static or static dissipative
|  |  |  | ix | resistant to the degradation from the continual use of hospital-grade disinfectants
|  |  | f |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids
|  |  |  | iii | opening to allow for the easy removal and fitting of a mattress core
if a zipper is used (waterproof preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | iv | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
Pressure Distribution
|  |  | g |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Of the four castors, at least two must incorporate a brake and at least two must incorporate directional steering; or the cot must include a central locking mechanism with directional steering capability. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 3.09 of the Tender Response Work | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Cots must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the cot's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Cots should be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All optional features offered, including but not limited to orthopaedic frames, radiopaque platform, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Cot Controls
|  |  | a |  | Bed controls must be designed to enable safe and intuitive adjustments by the caregiver only, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Cot Articulations
|  |  | a |  | Cots are encouraged to be equipped with electrically adjustable profile functions, including height adjustable mattress platform, backrest, and Trendelenburg positioning. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Cot Drop Sides
|  |  | a |  | With the cot side closed and the mattress in place, the internal depth, measured from the top of the mattress to the top edge of the cot side, must be at least 600mm. This also applies to the cot ends. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The cot drop-down sides must be capable of being securely stopped and held in place at various heights before fully lowered, and equipped with a child-resistant release mechanism. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cot drop side guides must ensure the cot's drop side moves smoothly, and are securely fastened to the cot frame. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | When the cot sides are up, the bottom edge of the cot sides and ends must be level with or below the top edge of the mattress support platform. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Cot offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Cot offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with (if electric)
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered cot, Width x Length x Height (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered cot (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress dimensions, Width x Length x Height (cm)
• list all materials (foam types) used in the construction of the mattress core
• describe the antimicrobial treatment applied to the mattress core
• detail the microclimate management technologies integrated into the mattress
• mattress cover moisture vapour transmission rate (MVTR) (g/m2/day)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
|  |  |  | xii | describe the drop sides (full length, bars, panel, etc.)
• height of cot sides from mattress platform (cm)
• bar spacing, if applicable (cm)
• cot end base securely latches to the frame (Yes/No)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | electric cot (Yes/No)
• intended clinical environment (General, NICU/PICU,
|  |  |  | xvi | list all accessories that are available offered cot
|  |  |  | xvii | list all options available for offered cot
|  |  |  | xviii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xix | state the Ingress Protection (IP) rating for both solids and liquids (if electrical)
|  |  |  | xx | manufacturer of offered equipment
• country of origin
|  |  |  | xxi | Year 1st manufactured
|  |  |  | xxii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiii | warranty period of cot (months)
|  |  |  | xxiv | warranty period of mattress (months)
|  |  |  | xxv | warranty period of mattress cover (months)
|  |  |  | xxvi | Estimated lifespan of cot (years of service)
|  |  |  | xxvii | Estimated lifespan of mattress (years of service)
|  |  |  | xxviii | Estimated lifespan of mattress cover (years of service)
|  |  |  | xxix | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxx | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 3.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 3.0 |  |  |  | Neonate & Paediatric Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 3.03 Youth Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Youth Bed
|  |  |  | i | A specialised bed that is sometimes referred to as a paediatric hospital bed, designed for toddlers, children, and adolescents requiring medical care. It features increased safety options with enhanced comfort and versatile accessibility that addresses the distinct physical and developmental needs of younger individuals. The bed serves as an intermediary solution by bridging the protective environment of a cot with the dimensions of an adult-sized hospital bed. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• nominal mattress dimensions of 80cm x 180cm
• nominal 100kg SWL
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment offered must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  |  | b |  | Beds equipped with removable siderails should have clearly visible and durable markers indicating the correct placement of the siderails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end should possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | If bed ends are designed for use as handholds, they must withstand the forces applied during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Beds with ergonomic handgrips integrated into the bed frame are encouraged, for caregivers to use to transport patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Beds designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Youth beds must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 60 kg
|  |  |  | ii | overall dimensions smaller than those of a typical hospital bed
|  |  | b |  | The mattress support platform must be height adjustable, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | lowers to a height that allows for easier and safer entry and exit for the child, minimising injury risk from falls.
|  |  | c |  | The bed should have a child-friendly appearance, which can include calming colours, appealing designs, or thematic elements to reduce anxiety and create a more welcoming environment, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | transparent panels and/or side rails, allowing for better patient visibility by caregivers and reducing a child's sense of confinement.
|  |  | d |  | Respondents should offer additional options to Youth Beds, such as anti-pinch side rails, integrated patient weighing scales, bed exit alarms, and under-bed lighting. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  |  | b |  | There should be reduced gaps between the mattress and the bedframe compared to a typical adult hospital bed to prevent entrapment and injury. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress
|  |  | a |  | Unless contradicted by the specific requirements in the clause below, clause 3.1(b), Youth Bed Mattresses must comply with all mandatory criteria within a subcategory of Category 10. Respondents are also required to complete one of the subcategories within Category 10 for their offered mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The following requirements apply to Youth Bed mattresses: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 60 kg
|  |  |  | ii | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | iii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iv | Deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  | 3.3 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Options & Accessories
|  |  | a |  | Beds featuring integrated mattresses must also comply with all criteria within a subcategory of Category 10. The respondent must also complete one of the subcategories with Category 10 for the offered bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 12.0 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | c |  | Beds must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed extensions must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Beds must be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | An optional self-help or patient pole with a height-adjustable handpiece must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | All optional features offered, including but not limited to orthopaedic frames, monitor shelves, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Controls
|  |  | a |  | Controls for bed adjustments should be intuitive and easily accessible for caregivers but out of reach or difficult for a child to operate unintentionally, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Bed Articulations
|  |  | a |  | All beds must be equipped with electrically adjustable profile functions, including backrest and knee break, and stepless adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer beds with features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | Side Rails
|  |  | a |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail locking mechanisms should be child-resistant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Side rails on Youth Beds should be higher than those on a typical adult hospital bed, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | gaps between side rail bars be smaller, or solid panels used, to prevent limb entrapment.
|  |  | f |  | Tenders are encouraged for beds featuring split side rails or side rails that do not increase the overall width of the bed frame. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.9 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Youth Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Youth Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress platform dimensions (without extensions), Width x Length (cm)
• maximum extended mattress platform dimensions for mattress extensions, Width x Length (cm)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes Fowler's positioning control button (Yes/No)
• includes patient anti-migration feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
|  |  |  | xii | describe the side rails (full, 3/4 length, split, bars/filled-in)
• height of side rail from mattress platform (cm)
• gap between side rail bars (cm), or filled-in panel
• bed frame width with the side rails (cm)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed (self-help pole, IV pole, integrated mattress)
|  |  |  | xvii | list all options available for offered bed
• fifth wheel (Yes/No)
• motorised drive (Yes/No)
• fall prevention detection (Yes/No)
• describe the child-friendly aesthetics options
• list all others
|  |  |  | xviii | state the size of the display, H x W (cm) where applicable
|  |  |  | xix | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xx | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xxi | state the 'Drop Test' height for handset
|  |  |  | xxii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | warranty period offered on bed structural frame (months)
|  |  |  | xxv | warranty period offered on bed electronics (months)
|  |  |  | xxvi | warranty period offered on bed handset & cables (months)
|  |  |  | xxvii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxviii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxix | Estimated lifespan of device (years of service)
|  |  |  | xxx | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 3.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 3.0 |  |  |  | Neonate & Paediatric Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 3.04 Infant Incubators
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Infant Incubator
|  |  |  | i | A mains-powered device designed to create a controlled microenvironment with optimal temperature and humidity for premature infants and newborns who have difficulty regulating their own body temperature. It features a transparent, easily removable hood or canopy, a comfortable mattress, a regulated air warming system, and controls for maintaining preset ambient temperature and humidity levels. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Infant Incubators offered must be included in the ARTG. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The incubator must comply with the performance requirements specified in IEC 60601-2-19 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The incubator mattress must comply with the performance requirements specified in AS/NZS 8811.1 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment offered must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | model designation
|  |  |  | iii | serial number
|  |  |  | iv | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The offered equipment must be free of protrusions that could cause injury or damage mattress, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | For infant incubators lacking hand grip attachments, the hood must be designed to securely withstand the push force applied by caregivers during transport, preventing dislodgment. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | All metal surfaces, excluding stainless steel, must be powder coated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise incubators that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Infant Incubators must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Air Mode (Manual): automatically maintain the set temperature within the enclosed hood between 23°C and 38°C, adjustable in 0.1°C increments
|  |  |  | ii | Skin Mode (Automatic): automatically adjust temperature within the enclosed hood to maintain patient’s skin temperature between 34°C and 37°C, adjustable in 0.1°C increments
|  |  |  | iii | over-temperature protection
|  |  |  | iv | includes an integrated backup thermostat
|  |  |  | v | include humidity control
|  |  |  | vi | include oxygen concentration control
|  |  |  | vii | heater power indicator
|  |  |  | viii | noise levels inside the hood should be running at or below 45 dBA, with transient noise peaks not exceeding 65 dBA
|  |  |  | ix | equipped with air inlet filters
|  |  |  | x | clear canopy with a hood access door
|  |  |  | xi | four or more hand access ports
|  |  |  | xii | four or more tubing ports
|  |  | b |  | The incubator must feature height adjustment and mattress tilt capabilities. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents should offer additional options, such as integrated patient weighing scales, pulse oximeter, double wall canopy, IV pole, and x-ray cassette tray. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress
|  |  | a |  | The mattress must be comfortable and pressure relieving. The foam should be certified low-VOC with the density of known less than 20kg/m³. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Minimum dimensions to accommodate an infant, 60cm x 37cm | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Mattress cover must be: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable, waterproof, breathable, antistatic, flame retardant, and hypoallergenic
|  |  |  | ii | seamless top
|  |  |  | iii | cleanable with a hospital grade disinfection product
|  |  |  | iv | easily removable
Pressure Distribution
|  |  | d |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | Incubator must have fully enclosed castors that are non marking and anti-static. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | At least two casters must feature easily accessible brakes, with one or more being steering casters. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories, Options, and Consumables for the offered equipment in Categories 3.10 to 3.12 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | a |  | HSV will only accept temperature probes/sensors from Respondents who also submit a corresponding equipment. The offered temperature probes/sensors must be listed on either: | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | HSPVC2023-021 Monitoring Products, or
|  |  |  | ii | HPVITS2026-060 Beds, Mattresses, Patient Trolleys, and Treatment Chairs, only if the product has not been awarded on HPVC2023-021 Monitoring Products.
|  |  | b |  | Replacement medical gas hoses should be listed in HPVC2024-075 Physiological Monitoring and Anaesthetic Gas Delivery Systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
[5 empty] | Controls must be designed to enable safe and intuitive adjustments by the operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Incubator offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b | iii | For each model Incubator offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered Incubator at min height, Width x Length x Height (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | therapeutic weight range (min/max kg)
|  |  |  | ix | adjustable height range from floor level of the mattress tray (cm)
|  |  |  | x | mattress dimensions, Width x Length x Depth (cm)
• vertical hood-to-mattress distance (cm)
• hood access door dimensions, Height x Width (cm)
|  |  |  | xi | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiii | equipped with caregiver handgrips (Yes/No)
|  |  |  | xiv | network connectivity (Yes/No)
|  |  |  | xv | list all accessories that are available offered bed (IV pole, storage tray)
|  |  |  | xvi | list all options available
• phototherapy (Yes/No)
- max spectral irradiance (μW/cm2/nm) of the Phototherapy light measured at (distance in cm)
• integrated scales (Yes/No)
• x-ray tray (Yes/No)
• integrated pulse oximeter (Yes/No)
- pulse oximetry technology (Masimo, Nellcor)
• list all others
|  |  |  | xvii | Performance:
• adjustable temperature setting in air-mode (manual), range and resolution (°C)
• adjustable temperature setting in skin-mode (servo), range and resolution (°C)
- temperature measurement, range and accuracy (°C)
• humidity setting, range and resolution (% RH)
• oxygen setting, range and resolution (%)
- oxygen measurement, range and accuracy (%)
• scales weight measurement, range and interval (g)
|  |  |  | xviii | state the size of the display, H x W (cm) where applicable
• list all monitored parameters shown on display
|  |  |  | xix | list all alarms
|  |  |  | xx | state the over temperature cutout (°C)
|  |  |  | xxi | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (operating time at full heater power)
• gauge indicating remaining battery capacity
|  |  |  | xxii | noise level inside hood (dBA)
|  |  |  | xxiii | air flow over the mattress (cm/s)
|  |  |  | xxiv | state the Ingress Protection (IP) rating for both solids and liquids
|  |  |  | xxv | manufacturer of offered equipment
• country of origin
|  |  |  | xxvi | Year 1st manufactured
|  |  |  | xxvii | warranty period offered (months)
|  |  |  | xxviii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxix | Estimated device's end-of-sales life (state the year)
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 3.05
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 3.0 |  |  |  | Neonate & Paediatric Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 3.05 Infant Incubators, Transport
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Transport Infant Incubator
|  |  |  | i | A mains-powered device designed to create a controlled microenvironment with optimal temperature and humidity for premature infants and newborns who have difficulty regulating their own body temperature for transporting infants either outside or within the healthcare facility. It typically operating on mains electricity when not in use for transportation; during transport is connected to an ambulance electrical outlet or is battery-operated. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Infant Incubators offered must be included in the ARTG. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The incubator must comply with the performance requirements specified in IEC 60601-2-19 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The incubator mattress must comply with the performance requirements specified in AS/NZS 8811.1 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment offered must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | model designation
|  |  |  | iii | serial number
|  |  |  | iv | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The offered equipment must be free of protrusions that could cause injury. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For infant incubators lacking hand grip attachments, the hood must be designed to securely withstand the push force applied by caregivers during transport, preventing dislodgment. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | All metal surfaces, excluding stainless steel, must be powder coated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise incubators that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Infant Incubators must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Air Mode (Manual): automatically maintain the set temperature within the enclosed hood between 23°C and 38°C, adjustable in 0.1°C increments
|  |  |  | ii | (desirable) Skin Mode (Automatic): automatically adjust temperature within the enclosed hood to maintain patient’s skin temperature between 34°C and 37°C, adjustable in 0.1°C increments
|  |  |  | iii | over-temperature protection
|  |  |  | iv | includes an integrated backup thermostat
|  |  |  | v | include humidity control
|  |  |  | vi | include oxygen monitor
|  |  |  | vii | heater power indicator
|  |  |  | viii | noise levels inside the hood should be running at or below 45 dBA, with transient noise peaks not exceeding 65 dBA
|  |  |  | ix | equipped with air inlet filters
|  |  |  | x | clear canopy with a hood access door
|  |  |  | xi | two or more hand access ports
|  |  |  | xii | two or more tubing ports
|  |  |  | xiii | external DC power input
|  |  | b |  | Respondents should offer additional options, such as oxygen concentration controller, double wall canopy, servo-temperature control, and humidity control | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress
|  |  | a |  | The mattress must be comfortable and pressure relieving. The foam should be certified low-VOC with the density of known less than 20kg/m³. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Minimum dimensions to accommodate an infant, 60cm x 30cm | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Mattress cover must be: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable, waterproof, breathable, antistatic, flame retardant, and hypoallergenic
|  |  |  | ii | seamless top
|  |  |  | iii | cleanable with a hospital grade disinfection product
|  | 3.2 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories, Options, and Consumables for the offered equipment in Categories 3.10 to 3.12 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | HSV will only accept temperature probes/sensors from Respondents who also submit a corresponding equipment. The offered temperature probes/sensors must be listed on either: | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | HSPVC2023-021 Monitoring Products, or
|  |  |  | ii | HPVITS2026-060 Beds, Mattresses, Patient Trolleys, and Treatment Chairs, only if the product has not been awarded on HPVC2023-021 Monitoring Products.
|  |  | c |  | Replacement medical gas hoses should be listed in HPVC2024-075 Physiological Monitoring and Anaesthetic Gas Delivery Systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Controls
[5 empty] | Controls must be designed to enable safe and intuitive adjustments by the operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Incubator offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Incubator offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of Incubator, Width x Length x Height (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | therapeutic weight range (min/max kg)
|  |  |  | ix | mattress dimensions, Width x Length x Depth (cm)
• vertical hood-to-mattress distance (cm)
• hood access door dimensions, Height x Width (cm)
|  |  |  | x | list all accessories that are available offered bed (IV pole, safety straps)
|  |  |  | xi | list all options available
• phototherapy (Yes/No)
- max spectral irradiance (μW/cm2/nm) of the Phototherapy light measured at (distance in cm)
• integrated pulse oximeter (Yes/No)
- pulse oximetry technology (Masimo, Nellcor)
• oxygen controller (Yes/No)
• oxygen connection (wall connection/flow meter)
• list all others
|  |  |  | xii | Performance:
• adjustable temperature setting in air-mode (manual), range and resolution (°C)
• adjustable temperature setting in skin-mode (servo), range and resolution (°C)
- temperature measurement, range and accuracy (°C)
• humidity setting, range and resolution (% RH)
• oxygen setting, range and resolution (%)
- oxygen measurement, range and accuracy (%)
• scales weight measurement, range and interval (g)
|  |  |  | xiii | state the size of the display, H x W (cm) where applicable
• list all monitored parameters shown on display
|  |  |  | xiv | list all alarms
|  |  |  | xv | state the over temperature cutout (°C)
|  |  |  | xvi | external DC power input, volts and current
|  |  |  | xvii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (transport time at 36.0°C)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xviii | noise level inside hood (dBA)
|  |  |  | xix | air flow over the mattress (cm/s)
|  |  |  | xx | state the Ingress Protection (IP) rating for both solids and liquids
|  |  |  | xxi | manufacturer of offered equipment
• country of origin
|  |  |  | xxii | warranty period offered (months)
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | Estimated lifespan of device (years of service)
|  |  |  | xxvii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxviii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 3.06
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 3.0 |  |  |  | Neonate & Paediatric Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 3.06 Infant Incubator/Radiant Warmers
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Infant Incubator/Radiant Warmer
|  |  |  | i | A mains-powered device that integrates the benefits of an infant incubator with the accessibility of an infant radiant warming unit. It features a centralised control and monitoring system for both functions, allowing for effortless conversion between the enclosed incubator and open-bed radiant warming modes. Ensures a stable and continuous thermal environment for infants, eliminating the need for disruptive transfers between separate units. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Offered equipment must be included in the ARTG. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | The incubator must comply with the performance requirements specified in IEC 60601-2-19 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The incubator mattress must comply with the performance requirements specified in AS/NZS 8811.1 | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment offered must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | model designation
|  |  |  | iii | serial number
|  |  |  | iv | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The offered equipment must be free of protrusions that could cause injury or damage mattress, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | For infant incubators lacking hand grip attachments, the hood must be designed to securely withstand the push force applied by caregivers during transport, preventing dislodgment. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | All metal surfaces, excluding stainless steel, must be powder coated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise incubators that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The offered equipment must provide effective thermoregulation for infants in both closed-incubator and open-radiant warmer care environments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The closed-incubator component of the device must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Air Mode (Manual): automatically maintain the set temperature within the enclosed hood between 23°C and 38°C, adjustable in 0.1°C increments
|  |  |  | ii | Skin Mode (Automatic): automatically adjust temperature within the enclosed hood to maintain patient’s skin temperature between 34°C and 37°C, adjustable in 0.1°C increments
|  |  |  | iii | over-temperature protection
|  |  |  | iv | includes an integrated backup thermostat
|  |  |  | v | include humidity control
|  |  |  | vi | include oxygen concentration control
|  |  |  | vii | heater power indicator
|  |  |  | viii | noise levels inside the hood should be running at or below 45 dBA, with transient noise peaks not exceeding 65 dBA
|  |  |  | ix | equipped with air inlet filters
|  |  |  | x | clear canopy with a hood access door
|  |  |  | xi | four or more hand access ports
|  |  |  | xii | four or more tubing ports
|  |  | c |  | For the open-warmer component, the infant can be warmed using either radiant or conductive heat (e.g. a heated mattress), or a combination of both. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The open-warmer component of the device must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | skin-mode (servo controlled): automatically control the heater output to maintain patient’s skin temperature between 35°C and 37°C, adjustable in 0.1°C increments
|  |  |  | ii | manual mode: heater run for a set duration, output manually adjustable, and an alarm sound when set time reached
|  |  |  | iii | the radiant heater, where available, have an output of more than 350Watts
|  |  |  | iv | include an examination light with a dimmable function, capable of providing a minimum illumination of 800 lux at the mattress centre
|  |  |  | v | a built-in APGAR timer with audible and visual alerts at 1, 5, and 10 minutes after birth
|  |  | e |  | The offered equipment must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a mobile device
|  |  |  | ii | feature height adjustment and mattress tilt capabilities
|  |  |  | iii | accommodate sub-infant radiographic imaging plates
|  |  |  | iv | include an integrated patient weighing scale
|  |  |  | v | ergonomic handgrips integrated into the frame
|  |  | f |  | Respondents should offer additional options, such as integrated resuscitation module, pulse oximeter, double wall canopy, phototherapy light, and drawers and shelves. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress
|  |  | a |  | The mattress must be comfortable and pressure relieving. The foam should be certified low-VOC with the density of known less than 20kg/m³. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Mattress must be radiolucent | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Minimum dimensions to accommodate an infant, 60cm x 37cm | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Mattress cover must be: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable, waterproof, breathable, antistatic, flame retardant, and hypoallergenic
|  |  |  | ii | seamless top
|  |  |  | iii | cleanable with a hospital grade disinfection product
|  |  |  | iv | easily removable
Pressure Distribution
|  |  | e |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | Incubator must have fully enclosed castors that are non marking and anti-static. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | At least two casters must feature easily accessible brakes, with one or more being steering casters. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories, Options, and Consumables for the offered equipment in Categories 3.10 to 3.12 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | HSV will only accept temperature probes/sensors from Respondents who also submit a corresponding equipment. The offered temperature probes/sensors must be listed on either: | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | HSPVC2023-021 Monitoring Products, or
|  |  |  | ii | HPVITS2026-060 Beds, Mattresses, Patient Trolleys, and Treatment Chairs, only if the product has not been awarded on HPVC2023-021 Monitoring Products.
|  |  | d |  | Replacement medical gas hoses should be listed in HPVC2024-075 Physiological Monitoring and Anaesthetic Gas Delivery Systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Infant Incubator/Radiant Warmer offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b | i | For each model Infant Incubator/Radiant Warmer offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | ii | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | UMDNS code and term
|  |  |  | iv | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | v | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | vi | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vii | the external dimensions of the offered equipment, Width x Length x Height (cm)
|  |  |  | viii | weight of the offered equipment (kg)
|  |  |  | ix | therapeutic weight range (min/max kg)
|  |  |  | x | adjustable height range from floor level of the mattress tray (cm)
|  |  |  | xi | mattress dimensions, Width x Length x Depth (cm)
• hood access door dimensions, Height x Width (cm)
• vertical hood-to-mattress distance (cm)
• radiant heater-to-mattress centre distance (cm)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | network connectivity (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed (IV pole, storage tray)
|  |  |  | xvii | list all options available
• integrated scales (Yes/No)
• phototherapy (Yes/No)
- max spectral irradiance (μW/cm2/nm) of the Phototherapy light measured at (distance in cm)
• integrated pulse oximeter (Yes/No)
- pulse oximetry technology (Masimo, Nellcor)
• integrated resuscitation module (Yes/No)
• list all others
|  |  |  | xviii | Performance - closed-incubator:
• adjustable temperature setting in air-mode (manual), range and resolution (°C)
• adjustable temperature setting in skin-mode (servo), range and resolution (°C)
- temperature measurement, range and accuracy (°C)
• humidity setting, range and resolution (% RH)
• oxygen setting, range and resolution (%)
- oxygen measurement, range and accuracy (%)
• scales weight measurement, range and interval (g)
|  |  |  | xix | Performance - open-warmer:
• radiant heat source, power (Watts) and element type
• radiant heater temperature setting in skin-mode (servo), range and resolution (°C)
- at maximum power, intensity of heater's thermal radiation (mW/cm2) at mattress centre
- distance between heater and mattress centre (cm)
- temperature measurement, range and accuracy (°C)
• prewarm mode (Yes/No)
• heated mattress output power (Watts)
• conductive heater temperature setting in skin-mode (servo), range and resolution (°C)
- temperature measurement, range and accuracy (°C)
|  |  |  | xx | state the closed-incubator over temperature cutout (°C)
|  |  |  | xxi | state the size of the display, H x W (cm) where applicable
• list all monitored parameters shown on display
|  |  |  | xxii | list all alarms
|  |  |  | xxiii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (operating time at full heater power)
• gauge indicating remaining battery capacity
|  |  |  | xxiv | closed mode, noise level inside hood (dBA)
|  |  |  | xxv | closed mode, air flow over the mattress (cm/s)
|  |  |  | xxvi | state the Ingress Protection (IP) rating for both solids and liquids
|  |  |  | xxvii | manufacturer of offered equipment
• country of origin
|  |  |  | xxviii | Year 1st manufactured
|  |  |  | xxix | warranty period offered (months)
|  |  |  | xxx | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxi | Estimated device's end-of-sales life (state the year)
|  |  |  | xxxii | Estimated lifespan of device (years of service)
|  |  |  | xxxiii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxiv | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 3.07
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 3.0 |  |  |  | Neonate & Paediatric Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 3.07 Infant Resuscitation Trolleys
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Infant Resuscitation Trolley
|  |  |  | i | A device designed for use during stabilisation and resuscitation of newborns, particularly in delivery and neonatal care settings. It typically includes a radiant warmer and may also incorporate a supplemental heating pad. Standard features often include a low-pressure suction system, an oxygen (O2) supply, a gas blender (oxygen/air), and an automated resuscitation device. The trolley may be configured as: | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• a mobile cart, with or without an integrated mattress surface
• a wall mounted unit, with or without an integrated mattress surface
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Offered equipment must be included in the ARTG. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The mattress must comply with the performance requirements specified in AS/NZS 8811.1, where applicable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment offered must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | model designation
|  |  |  | iii | serial number
|  |  |  | iv | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The offered equipment must be free of protrusions that could cause injury or damage the mattress, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
designed to reduce the risk of tripping or falling.
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Participating Health Services will prioritise incubators that facilitate easy movement and operation, where applicable. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The offered equipment must meet the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | radiant heater with an output of more than 350Watts
|  |  |  | ii | in skin-mode (servo controlled) automatically control the heater output to maintain patient’s skin temperature between 35°C and 37°C, adjustable in 0.1°C increments
|  |  |  | iii | in manual mode heater run for a set duration, output manually adjustable, and sound alarm when set time reached
|  |  |  | iv | include an examination light with a dimmable function, providing greater than 600 lux light itensity at maximum
|  |  |  | v | include an APGAR timer with audible and visual alerts at 1, 5, and 10 minutes after birth
|  |  | b |  | The offered equipment should include the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | oxygen/medical air blender
|  |  |  | ii | resuscitator
|  |  |  | iii | low pressure suction
|  |  | c |  | Respondents should offer additional options, such as oxygen monitor, pulse oximeter, heated mattress, changeover regulator for wall supply (when connected), integrated weighing scales, sub-infant x-ray plate holder, and drawers and shelves. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Medical Gas Blender
|  |  | a |  | Standalone gas blenders must be offered from the HPVC2021-076 Ventilators agreement. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The integrated blender must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | blend medical air and oxygen, sourced either from a hospital gas pipeline or cylinders, to deliver oxygen concentrations ranging from 21% to 100% at both high and low flow rates.
|  |  |  | ii | operate at gas pressures between 240 and 480kPa (35-70psi)
|  |  |  | iii | trigger an alarm when the supply pressure differential reaches 140 ±45kPa (20 ±6psi)
|  |  |  | iv | dial accuracy within ±5% of the indicated value
|  | 3.2 |  |  | Resuscitator
|  |  | a |  | Standalone resuscitators must be offered from the HPVC2021-076 Ventilators agreement. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The integrated resuscitator must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be capable of delivering T-piece resuscitation
|  |  |  | ii | equipped with an airway pressure manometer with a range of -10 to 80 cmH2O
|  |  |  | iii | user adjustable PIP
|  |  |  | iv | pressure relief valve
|  | 3.3 |  |  | Low Pressure Suction
|  |  | a |  | Standalone low pressure suction device must be offered from the HPVC2024-075 Physiological Monitoring and Anaesthetic Gas Delivery Systems. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The integrated low pressure suction device must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | vacuum no greater than 200mmHg
|  | 3.4 |  |  | Mattress
|  |  | a |  | The mattress must be comfortable and pressure relieving. The foam should be certified low-VOC with the density of known less than 20kg/m³. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Mattress must be radiolucent | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Minimum dimensions to accommodate an infant, 60cm x 37cm | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Mattress cover must be: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable, waterproof, breathable, antistatic, flame retardant, and hypoallergenic
|  |  |  | ii | seamless top
|  |  |  | iii | cleanable with a hospital grade disinfection product
|  |  |  | iv | easily removable
Pressure Distribution
|  |  | e |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.5 |  |  | Castors and Braking Systems
|  |  | a |  | Castors must be fully enclosed, non marking, and anti-static. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | At least two casters must feature easily accessible brakes, with one or more being steering casters. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories, Options, and Consumables for the offered equipment in Categories 3.10 to 3.12 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | HSV will only accept temperature probes/sensors from Respondents who also submit a corresponding equipment. The offered temperature probes/sensors must be listed on either: | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | HSPVC2023-021 Monitoring Products, or
|  |  |  | ii | HPVITS2026-060 Beds, Mattresses, Patient Trolleys, and Treatment Chairs, only if the product has not been awarded on HPVC2023-021 Monitoring Products.
|  |  | d |  | Replacement medical gas hoses should be listed in HPVC2024-075 Physiological Monitoring and Anaesthetic Gas Delivery Systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model equipment offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model equipment offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered equipment, Width x Length x Height (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | equipped with caregiver handgrips (Yes/No)
• therapeutic weight range (min/max kg)
|  |  |  | ix | Performance - open-warmer:
• radiant heat source, power (Watts) and element type
• radiant heater temperature setting in skin-mode (servo), range and resolution (°C)
- at maximum power, intensity of heater's thermal radiation (mW/cm2) at mattress centre
- distance between heater and mattress centre (cm)
- temperature measurement, range and accuracy (°C)
• prewarm mode (Yes/No)
• heated mattress output power (Watts)
• conductive heater temperature setting in skin-mode (servo), range and resolution (°C)
- temperature measurement, range and accuracy (°C)
|  |  |  | x | mobile device with castors (Yes/No)
• castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
• describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xi | includes padded mattress surface (Yes/No)
• adjustable height range from floor level of the mattress tray (cm)
• mattress dimensions, Width x Length x Depth (cm)
• radiant heater-to-mattress centre distance (cm)
|  |  |  | xii | Resuscitator (Yes/No)
• PIP range (cmH2O)
• PEEP range (cmH2O)
• pressure relief valve max setting (cmH2O)
|  |  |  | xiii | oxygen/medical air blender (Yes/No)
• oxygen monitor included (Yes/No)
• blender gas bleed rate (L/min)
|  |  |  | xiv | low pressure suction (Yes/No)
• pressure range (mmHg)
• collection container size (mL)
|  |  |  | xv | list all accessories that are available offered bed (IV pole, storage tray)
|  |  |  | xvi | list all options available
• integrated pulse oximeter (Yes/No)
- pulse oximetry technology (Masimo, Nellcor)
• infant weighing scales (Yes/No)
- scales weight measurement, range and interval (g)
• auxiliary oxygen flow meter (Yes/No)
- flow range (LPM)
• list all others
|  |  |  | xvii | network connectivity (Yes/No)
|  |  |  | xviii | state the size of the display, H x W (cm) where applicable
• list all monitored parameters shown on display
• radiant warmer skin-mode (servo controlled) temperature display range (°C)
• conductive warmer skin-mode (servo controlled) temperature display range (°C)
|  |  |  | xix | list all alarms
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (operating time at full heater power)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids
|  |  |  | xxii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxv | warranty period offered (months)
|  |  |  | xxvi | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxvii | Estimated lifespan of device (years of service)
|  |  |  | xxviii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxix | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 4.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 4.0 |  |  |  | O&G Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 4.01 Birthing Beds
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Birthing Bed
|  |  |  | i | An electrically powered device, designed to enhance comfort and safety for women during labour and delivery. It facilitates a variety of physiological birthing positions through multi-section power adjustments of the backrest, seat, and leg sections. Features integrated leg supports, a discreet afterbirth receptacle, and universal provisions for medical support accessories. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | General ward beds offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All beds must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The bed frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the bed's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the bed | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | For improved usability, bed controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Participating Health Services will prioritise beds that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Birthing Beds must have a safe working load of at least 200 kg: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | include adjustable calf supports, and
|  |  |  | ii | feature a foot section that is easily removable or retractable.
|  |  | b |  | The mattress support platform must have an adjustable height, and the dimensions for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress must be at least 180 cm in length and no more than 100 cm in width.
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg, where | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress width may exceed 90 cm.
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Mattress
|  |  | a |  | The mattress must be explicitly designed for use on a birthing bed and comprise at least two independent sections: an upper section and a removable lower (foot) section. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The mattress construction and materials must incorporate a profiling system that enables seamless adaptation to the various articulated positions of the birthing bed frame without kinking, buckling, or creating excessive pressure points. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The mattress must be engineered to effectively distribute patient weight and prevent "bottoming out". | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The mattress construction materials and methods shall ensure durability and prevent disintegration or fragmentation under normal use conditions throughout its expected lifespan. All materials in contact with the patient must be biocompatible and not have any deleterious effects on body tissue. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | For the backrest and seat sections, the mattress depth must be at least 10cm, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | with a therapeutic load capacity equivalent to or exceeding the safe working load of the birthing bed.
Labelling
|  |  | f |  | Mattresses should be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | serial number
|  |  |  | iii | date of manufacture
|  |  | g |  | Participating Health Services must be permitted to write on mattresses without voiding the warranty.
|  |  | h |  | The mattress cover should include a space to write the date of purchase.
Mattress Cover
|  |  | i |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | j |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
Pressure Distribution
|  |  | k |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.3 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 4.05 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | Beds must include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator. The installed components must not impede the bed's function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Beds must be equipped with secure and convenient attachment points for drainage collection devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All optional features offered, including but not limited to labour bar, monitor shelves, and under-bed storage, must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Birthing beds should include optional labour grips, squat bars, and footrests. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 1.3 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Bed Controls
|  |  | a |  | Patient-operated bed controls must be designed to enable safe and intuitive adjustments for a single patient, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and integrating secure caregiver lockout functionality to prevent unintended adjustments. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | To ensure patient controls are readily accessible and simple to operate, they should ideally be located on or integrated into the bed side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.7 |  |  | Bed Articulations
|  |  | a |  | Birthing beds must be equipped with electrically adjustable profile functions, including backrest and hip tilt, and stepless adjustments for Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.8 |  |  | Side Rails
|  |  | a |  | Beds equipped with continuous, full-length side rails will be excluded from consideration. Respondents must offer beds with patient-safe side rail configurations, including, but not limited to, split or other partial-length designs. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Where collapsible side rails are offered for a bed, they must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | d |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Side rails should not increase the overall width of the bed frame, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to safely and effectively support the patient during ingress and egress from the bed.
|  | 3.9 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Beds must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Birthing Bed offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Birthing Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered bed, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | SWL of offered bed (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | x | mattress platform dimensions (without extensions), Width x Length (cm)
• maximum extended mattress platform dimensions for mattress extensions, Width x Length (cm)
• describe the mattress platform sections (removable/relocatable foot section)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes pelvic tilt (Yes/No)
• includes patient perineal edge adjustment feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for bed to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xii | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum bed width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | bed frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered bed (self-help pole, IV pole, integrated mattress)
|  |  |  | xvii | list all options available for offered bed
• labour grips (Yes/No)
• squat bar (Yes/No)
• foot supports (Yes/No)
• fifth wheel (Yes/No)
• motorised drive (Yes/No)
• list all others
|  |  |  | xviii | state the size of the display, H x W (cm) where applicable
|  |  |  | xix | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xx | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• bed
• handset (if applicable)
|  |  |  | xxi | state the 'Drop Test' height for handset
|  |  |  | xxii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxv | warranty period offered on bed structural frame (months)
|  |  |  | xxvi | warranty period offered on bed electronics (months)
|  |  |  | xxvii | warranty period offered on bed handset & cables (months)
|  |  |  | xxviii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxix | Estimated lifespan of device (years of service)
|  |  |  | xxx | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 4.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 4.0 |  |  |  | O&G Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 4.02 Gynaecology Procedure Chairs
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Gynaecology Procedure Chair
|  |  |  | i | An electrically powered device, designed to provide comfortable and secure support for patients undergoing gynaecological examinations, treatments, and minor surgical procedures. It includes motorised height adjustment, head and armrests, adjustable calf supports, and a backrest capable of inclining from a vertical to a near-horizontal position. Furthermore, it offers convenient mounting points for essential equipment, such as colposcopes, monitors; a concealed instrument tray; and electronic controls for precise chair articulation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Gynaecology Procedure Chairs offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | a |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The chair must be free of protrusions that could damage the cushion or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Chairs should be designed that users cannot inadvertently come into contact with any moving "scissor action" lift mechanisms during use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For improved usability, chair controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise chairs that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Unlike a standard examination couch, the offered chair needs to be a specialised and highly adjustable chair designed for minor procedures. Gynaecology examination couches need to be offered in Category 4.03. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Chair must have a safe working load of at least 200 kg, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | height-adjustable seat base,
|  |  |  | ii | independently adjustable foot plates,
|  |  |  | iii | a clearly marked CPR release lever for rapid backrest lowering,
|  |  |  | iv | Trendelenburg positioning,
|  |  |  | v | foot controller, which should include pre-set buttons
|  |  | c |  | Non-castor chairs are preferred. Chairs equipped with casters must have their casters disengaged from the floor during use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Chairs should be equipped with an instrument tray, mounting points for equipment (e.g., colposcopes, display monitors), a patient handset, a backup battery, and a paper roll holder. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The respondent should offer optional independently adjustable calf supports. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the chair: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | The chair's backrest and seat cushions must possess a therapeutic load capacity equal to or greater than the chair's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | d |  | The chair cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | seamless upholstery
|  |  | e |  | Chair cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | f |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 4.06 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | c |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Controls
|  |  | a |  | Chair controls must be designed to enable safe and intuitive adjustments for users, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The chair must include a foot controller with switches to activate saved preset settings. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | CPR Release
|  |  | a |  | The chair controls must provide the capability to adjust the chair to a completely flat, level position. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Chairs should have a CPR release mechanism that is clearly identifiable and easily accessible. If so, upon activation the backrest must safely return to a horizontal position in a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Gynaecology Procedure Chair offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | Total Cost of Ownership (TCO) data covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Gynaecology Procedure Chair offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered chair when fully lowered, Width x Length x Height (cm)
|  |  |  | vii | the maximum floor area required for the chair, L x W (cm)
|  |  |  | viii | weight of the offered equipment (kg)
|  |  |  | ix | SWL of offered chair (kg)
• therapeutic weight range (min/max kg)
|  |  |  | x | adjustable height range from floor level of the seat base (cm)
|  |  |  | xi | seat pan dimensions, Width x Length (cm)
• seat cushion height (cm)
|  |  |  | xii | list all accessories that are available offered bed (calf supports, arm rests)
|  |  |  | xiii | list all options available for offered bed
• hand grips (Yes/No)
• CPR release mechanism (Yes/No)
• Instrument tray (Yes/No)
• paper roll holder (Yes/No)
• mounts for essential equipment (Yes/No)
- list the equipment
• list all others
|  |  |  | xiv | state the size of the display, H x W (cm) where applicable
|  |  |  | xv | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xvi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• equipment
• foot controller
• handset (if applicable)
|  |  |  | xvii | state the 'Drop Test' height for handset
|  |  |  | xviii | manufacturer of offered equipment
• country of origin
|  |  |  | xix | Year 1st manufactured
|  |  |  | xx | Estimated device's end-of-sales life (state the year)
|  |  |  | xxi | warranty period offered on chair structural frame (months)
warranty period offered on chair electronics (months)
|  |  |  | xxii | warranty period offered on chair cushion (months)
|  |  |  | xxiii | warranty period offered on chair cushion cover (months)
|  |  |  | xxiv | Estimated lifespan of chair (years of service)
|  |  |  | xxv | Estimated lifespan of chair cushion (years of service)
|  |  |  | xxvi | Estimated lifespan of chair cushion cover (years of service)
|  |  |  | xxvii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxviii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 4.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 4.0 |  |  |  | O&G Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 4.03 Gynae Exam Couches
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Gynae Exam Couch
|  |  |  | i | A table designed specifically for the comfort and efficient examination and treatment of female patients, primarily focusing on gynaecological, obstetric and sometimes urological procedures. It has a padded top cover with the stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Exam Couches must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Exam Couch must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | multi sectional
|  |  |  | v | include a foot switch
|  |  |  | vi | include independently adjustable foot plates
|  |  |  | vii | equipped with four castors that are retractable or lockable when the couch is in use
• that are foot operated and require minimal force to engage/disengage castors
|  |  |  | viii | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Exam Couches, such as handsets, breath hole, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following mandatory requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 4.06 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  |  | v | adjustable armrests
|  |  |  | vi | independently adjustable calf supports
|  |  |  | vii | instrument tray
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | e |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Exam Couch controls must be designed to enable safe and intuitive adjustments for users, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | having a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth.
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Gynae Exam Couch offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) analysis covering the first 10 years of the product's anticipated lifecycle | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Gynae Exam Couch offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered examination couch (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | patient (therapy) table dimensions, Width x Length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | ix | adjustable height range from floor level of the patient platform (cm)
|  |  |  | x | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | xi | no. of sections
• dimensions of each section (cm)
• backrest maximum incline/decline (in degrees)
• seat (pelvic section) maximum tilt, where applicable (in degrees)
• leg section maximum incline/decline, where applicable (in degrees)
|  |  |  | xii | electrically operated backrest (Yes/No)
|  |  |  | xiii | electrically operated seat tilt, where applicable (Yes/No)
|  |  |  | xiv | electrically operated leg section, where applicable (Yes/No)
|  |  |  | xv | Trendelenburg/Reverse Trendelenburg (Yes/No)
• maximum Trendelenburg/reverse Trendelenburg articulation (degrees)
|  |  |  | xvi | removable/foldaway leg section (Yes/No/Not Applicable)
|  |  |  | xvii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xviii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xix | list all consumables available for offered exam couch, e.g. disposable fitted sheets, paper roll
|  |  |  | xx | list all accessories available for offered exam couch (calf supports, arm rests)
|  |  |  | xxi | list all options available for offered exam couch
• hand grips (Yes/No)
• CPR release mechanism (Yes/No)
• Instrument tray (Yes/No)
• paper roll holder (Yes/No)
• mounts for essential equipment (Yes/No)
- list the equipment
• list all others
|  |  |  | xxii | state the size of the display, H x W (cm) where applicable
|  |  |  | xxiii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xxiv | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• equipment
• foot controller
• handset (if applicable)
|  |  |  | xxv | state the 'Drop Test' height for handset
|  |  |  | xxvi | length of handset/foot controller cable (m)
|  |  |  | xxvii | manufacturer of offered equipment
• country of origin
|  |  |  | xxviii | Year 1st manufactured
|  |  |  | xxix | Estimated device's end-of-sales life (state the year)
|  |  |  | xxx | warranty period offered on the frame (months)
|  |  |  | xxxi | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxxii | Estimated lifespan of device (years of service)
|  |  |  | xxxiii | Estimated lifespan of upholster cushions (years of service)
|  |  |  | xxxiv | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 4.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 4.0 |  |  |  | O&G Beds | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 4.04 Hospital Stretchers, Obstetric
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Obstetric Hospital Stretcher
|  |  |  | i | A hospital stretcher, offered in battery-powered or hydraulic models, designed for the safe and comfortable transport and positioning of patients throughout all stages of maternity care-from pregnancy and labour to delivery and postpartum recovery. Features adjustable leg rest that can be easily lowered, extended, or even completely removed to accommodate various birthing positions, including lithotomy positioning for vaginal deliveries. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretchers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Stretchers must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The stretcher frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the stretcher's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Stretchers must be equipped with ergonomic handgrips to enhance caregiver safety during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Stretchers designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | For improved usability, stretcher controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Participating Health Services will prioritise stretchers that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | include adjustable foot supports,
|  |  |  | ii | foot section that retracts are folds away
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator, where the installed components must not impede the stretcher's function
|  |  |  | v | an integrated IV pole mounting point
|  |  |  | vi | centrally controlled locking mechanism and directional steering
|  |  |  | vii | feature integrated, non-removable, anti-pinch side rails
|  |  | b |  | The mattress support platform must have an adjustable height, and the dimensions for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress at least 180 cm in length and 60 cm in width
|  |  | c |  | Respondents should offer additional options to Stretchers, such as secure and convenient attachment points for drainage collection devices, fifth wheel, patient belongings and medical record holders. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The stretcher's castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with stretchers must meet all mandatory criteria in Category 7.08. Additionally, Respondents must complete Category 7.08 for any offered stretchers that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 7.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | calf supports
|  |  |  | ii | integrated patient belongings storage
|  |  |  | iii | medical records and chart holder
|  |  |  | iv | radiolucent mattress platform with a mobile X-ray cassette holder allowing for full length supine imaging
radiolucent backrest with integrated upright chest X-ray cassette holder
|  |  |  | v | equipment mounting solutions
|  |  |  | vi | secured drainage collection device attachments
|  |  | e |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Stretchers should be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading, be equipped with a liquid-tight seal at the equipment entry point, and feature an integrated mains cable storage system; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Articulations
|  |  | a |  | Stretchers offered must be equipped with adjustable profile functions, including backrest, and adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Side Rails
|  |  | a |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Collapsible side rails must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The following side rail options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | split side rails
|  |  |  | ii | flush-folding side rails
|  |  |  | iii | side rails without vertical bars
|  |  |  | iv | designed to safely and effectively support the patient during ingress and egress from the stretcher.
|  | 3.7 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Stretchers must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the stretcher. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | Total Cost of Ownership (TCO) analysis covering the first 10 years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• manufacturer's recommended mattress, make, model
• describe the mattress platform sections (retractable foot section)
|  |  |  | x | no. of mattress platform sections listing each (backrest section, etc.)
• includes X-ray translucent backrest with cassette carrier (Yes/No)
• includes pelvic tilt (Yes/No)
• includes patient perineal edge adjustment feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for stretcher to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xi | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | list all accessories that are available offered stretcher (IV pole, cylinder holder, external charger for secondary battery, secondary battery)
|  |  |  | xvi | list all options available for offered stretcher
• integrated patient belongings storage (Yes/No)
• medical records and chart holder (Yes/No)
• equipment mounting solutions (Yes/No)
• secured drainage collection device attachments (Yes/No)
• labour grips (Yes/No)
• squat bar (Yes/No)
• foot supports (Yes/No)
• calf supports (Yes/No)
• fifth wheel (Yes/No)
• list all others
|  |  |  | xvii | state the size of the display, H x W (cm) where applicable
|  |  |  | xviii | stretcher connects to a 240V power outlet? (Yes/No)
|  |  |  | xix | removable secondary backup battery that staff can easily swap out (Yes/No)
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | warranty period offered on stretcher structural frame (months)
|  |  |  | xxvii | warranty period offered on stretcher electronics (months)
|  |  |  | xxviii | warranty period offered on stretcher handset & cables (months)
|  |  |  | xxix | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 5.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 5.0 |  |  |  | Procedure Chairs | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 5.01 Haemodialysis / Chemotherapy Chairs
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Haemodialysis Chair
|  |  |  | i | An electrically powered device, designed to support and position the patient in a seated or reclined posture during haemodialysis. Constructed from durable, easily cleanable materials, featuring independently adjustable, powered armrests, backrest, and footrests to optimise patient comfort, facilitate access, and enable seamless connection to the extracorporeal blood circuit. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Chemotherapy Chair
|  |  |  | i | An electrically powered device, designed to provide optimum support and customisable positioning for patients undergoing prolonged infusion treatments. Constructed from durable, easily cleanable materials, featuring independently adjustable, powered armrests, backrest, and footrests to reduce patient fatigue during the several hours required for chemotherapy infusions. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Procedure Chairs offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The chair must be free of protrusions that could damage the cushions or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Chairs should be designed that users cannot inadvertently come into contact with any moving "scissor action" lift mechanisms during use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For improved usability, chair controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise chairs that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Chair must have a safe working load of at least 200 kg, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | height-adjustable seat base,
|  |  |  | ii | arm boards as armrests,
|  |  |  | iii | a clearly marked CPR release lever for rapid backrest lowering,
|  |  |  | iv | patient handset to adjust recline functions
|  |  |  | v | both Trendelenburg and reverse Trendelenburg positioning,
|  |  |  | vi | casters.
|  |  | b |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the chair: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The chair's backrest and seat cushions must possess a therapeutic load capacity equal to or greater than the chair's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster existing cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The chair cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | seamless upholstery
|  |  | h |  | Chair cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under the maximum safe working load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking and steering system must be single-foot operable, accessible from both side of the chair, and clearly labelled indicating its function. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 5.06 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | An optional stainless steel IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | arm rests
|  |  |  | ii | patient weigh features
|  |  |  | iii | chair exit alarms
|  |  |  | iv | patient belongings storage container
|  |  |  | v | chart holders
|  |  |  | vi | secure and convenient attachment points for drainage collection devices
|  |  |  | vii | knee bend feature
|  |  |  | viii | foot stop
|  |  |  | ix | impact-resistant corner protectors
|  |  |  | x | split leg support
|  |  |  | xi | bariatric version
|  |  |  | xii | central locking device
|  |  |  | xiii | integrated tray tables for reading or meals
|  |  |  | xiv | built-in USB ports for charging devices
|  |  |  | xv | integrated heating or massage functions
|  |  | d |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Chairs with electric movement must incorporate an integrated battery backup system capable of providing power for several complete full-range adjustments in the event of a power outage. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Chair controls must be designed to enable safe and intuitive adjustments for users, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | CPR Release
|  |  | a |  | Chairs must have a CPR release mechanism that is clearly identifiable and easily accessible. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each chair offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | Acknowledged |  | Acknowledged |  | Acknowledged
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | Acknowledged |  | Acknowledged |  | Acknowledged
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | Acknowledged |  | Acknowledged |  | Acknowledged
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | Acknowledged |  | Acknowledged |  | Acknowledged
|  |  |  | v | pressure-mapping report | Data/Info Required | Acknowledged |  | Acknowledged |  | Acknowledged
|  |  | b |  | For each chair offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | Acknowledged |  | Acknowledged |  | Acknowledged
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered chair when fully lowered, Width x Length x Height (cm)
|  |  |  | vii | the maximum floor area required for the chair, L x W (cm)
|  |  |  | viii | weight of the offered equipment (kg)
|  |  |  | ix | SWL of offered chair (kg)
• therapeutic weight range (min/max kg)
|  |  |  | x | adjustable height range from floor level of the seat base (cm)
|  |  |  | xi | seat pan dimensions, Width x Length (cm)
• seat cushion height (cm)
|  |  |  | xii | list all accessories that are available offered bed (IV pole)
|  |  |  | xiii | list all options available for offered bed (USB port, etc.)
|  |  |  | xiv | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xv | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• equipment
• handset
|  |  |  | xvi | state the 'Drop Test' height for handset
|  |  |  | xvii | manufacturer of offered equipment
• country of origin
• country of final assembly
|  |  |  | xviii | Year 1st manufactured
|  |  |  | xix | Estimated device's end-of-sales life (state the year)
|  |  |  | xx | warranty period offered on chair structural frame (months)
|  |  |  | xxi | warranty period offered on chair electronics (months)
|  |  |  | xxii | warranty period offered on chair cushion (months)
|  |  |  | xxiii | warranty period offered on chair cushion cover (months)
|  |  |  | xxiv | Estimated lifespan of chair (years of service)
|  |  |  | xxv | Estimated lifespan of chair cushion (years of service)
|  |  |  | xxvi | Estimated lifespan of chair cushion cover (years of service)
|  |  |  | xxvii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxviii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 5.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 5.0 |  |  |  | Procedure Chairs | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 5.02 Procedure Chair
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Procedure Chairs
|  |  |  | i | An electrically powered device, designed to support and position the patient in a seated or reclined position, allowing healthcare providers comfortable access for examinations, treatments, and minor procedures. Its adaptable design, enhanced by accessories such as headrests, arm boards, footrests, and side rails, ensures optimal configuration for a diverse range of clinical needs. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• has a shorter wheelbase than a stretcher, of approximately 1.0m
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Procedure Chairs offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The chair must be free of protrusions that could damage the cushions or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
designed to reduce the risk of tripping or falling.
|  |  | b |  | Chairs should be designed that users cannot inadvertently come into contact with any moving "scissor action" lift mechanisms during use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For improved usability, chair controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise chairs that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Chair must have a safe working load of at least 200 kg, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | height-adjustable seat base,
|  |  |  | ii | collapsible/foldable/removable armrests,
|  |  |  | iii | a clearly marked CPR release lever for rapid backrest lowering,
|  |  |  | iv | patient handset
|  |  |  | v | both Trendelenburg and reverse Trendelenburg positioning,
|  |  |  | vi | casters.
|  |  | b |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the chair: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The chair's backrest and seat cushions must possess a therapeutic load capacity equal to or greater than the chair's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster existing cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The chair cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | seamless upholstery
|  |  | h |  | Chair cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under the maximum safe working load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | A central locking mechanism with directional steering operable via a single foot pedal is preferred. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Should individual casters be utilised, their braking and steering functions must be clearly marked and easily reachable from either side of the chair. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.06 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | All cables must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Chairs with electric movement must incorporate an integrated battery backup system capable of providing power for several complete full-range adjustments in the event of a power outage. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Chair controls must be designed to enable safe and intuitive adjustments for users, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | CPR Release
|  |  | a |  | Chairs must have a CPR release mechanism that is clearly identifiable and easily accessible. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Procedure Chair offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Procedure Chair offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered chair when fully lowered, Width x Length x Height (cm)
|  |  |  | vii | the maximum floor area required for the chair, L x W (cm)
|  |  |  | viii | weight of the offered equipment (kg)
|  |  |  | ix | SWL of offered chair (kg)
• therapeutic weight range (min/max kg)
|  |  |  | x | adjustable height range from floor level of the seat base (cm)
|  |  |  | xi | seat pan dimensions, Width x Length (cm)
• seat cushion height (cm)
|  |  |  | xii | list all accessories that are available offered bed (IV pole)
|  |  |  | xiii | list all options available for offered bed (USB port, etc.)
|  |  |  | xiv | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xv | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• equipment
• handset
|  |  |  | xvi | state the 'Drop Test' height for handset
|  |  |  | xvii | manufacturer of offered equipment
• country of origin
• country of final assembly
|  |  |  | xviii | Year 1st manufactured
|  |  |  | xix | Estimated device's end-of-sales life (state the year)
|  |  |  | xx | warranty period offered on chair structural frame (months)
warranty period offered on chair electronics (months)
|  |  |  | xxi | warranty period offered on chair cushion (months)
|  |  |  | xxii | warranty period offered on chair cushion cover (months)
|  |  |  | xxiii | Estimated lifespan of chair (years of service)
|  |  |  | xxiv | Estimated lifespan of chair cushion (years of service)
|  |  |  | xxv | Estimated lifespan of chair cushion cover (years of service)
availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxvi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 5.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 5.0 |  |  |  | Procedure Chairs | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 5.03 Podiatry Chairs
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Podiatry Chair
|  |  |  | i | An electrically powered device, designed to support a patient in a seated position, facilitating foot examination, treatment, and minor procedures. It features electrically adjustable height, reclining backrest (tilting to horizontal), electrically adjustable tilt, and footrests with lifting capabilities. The chair typically includes armrests and can be programmed for multiple standard positions. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Podiatry Chairs offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The chair must be free of protrusions that could damage the cushions or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Chairs should be designed that users cannot inadvertently come into contact with any moving "scissor action" lift mechanisms during use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For improved usability, chair controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise chairs without visible casters. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Unlike a standard examination couch, the offered chair needs to be a specialised and highly adjustable chair designed for minor procedures. Podiatry examination couches need to be offered in Category 6.04. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The chair must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Podiatry Chairs have a safe working load of at least 200kg
|  |  |  | ii | electrically height-adjustable seat base,
|  |  |  | iii | electrically tilt-adjustable seat base,
|  |  |  | iv | electrically adjustable backrest,
|  |  |  | v | adjustable leg rest
|  |  |  | vi | collapsible/foldable/removable armrests,
|  |  | c |  | Non-castor chairs are preferred. Chairs equipped with casters must have their casters disengaged from the floor during use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Chairs should be equipped with electrically adjustable leg rests and backup battery. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Podiatry Chairs must not feature split leg rests. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the chair: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The chair's backrest and seat cushions must possess a therapeutic load capacity equal to or greater than the chair's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster existing cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The chair cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | seamless upholstery
|  |  |  | viii | phthalate free
|  |  | h |  | Chair cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.06 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | All cables must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Controls
|  |  | a |  | The chair must include a foot controller with switches for programmable memory positions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Chair controls must be designed to enable safe and intuitive adjustments for users, featuring clear and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Handset, if available, must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Podiatry Chair offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Podiatry Chair offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered chair when fully lowered, Width x Length x Height (cm)
|  |  |  | vii | the maximum floor area required for the chair, L x W (cm)
|  |  |  | viii | weight of the offered equipment (kg)
|  |  |  | ix | SWL of offered chair (kg)
• therapeutic weight range (min/max kg)
|  |  |  | x | adjustable height range from floor level of the seat base (cm)
|  |  |  | xi | seat pan dimensions, Width x Length (cm)
• seat cushion height (cm)
|  |  |  | xii | list all accessories that are available for offered chair
|  |  |  | xiii | list all options that are available for offered chair
• electrically adjustable leg rests (Yes/No)
• paper roll holder (Yes/No)
• battery backup (Yes/No)
• transforms from chair to procedure table (Yes/No)
• list all others
|  |  |  | xiv | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xv | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• equipment
• foot controller
• handset (if applicable)
|  |  |  | xvi | state the 'Drop Test' height for handset
|  |  |  | xvii | manufacturer of offered equipment
• country of origin
|  |  |  | xviii | Year 1st manufactured
|  |  |  | xix | Estimated device's end-of-sales life (state the year)
|  |  |  | xx | warranty period offered on chair structural frame (months)
|  |  |  | xxi | warranty period offered on chair electronics (months)
|  |  |  | xxii | warranty period offered on chair cushion (months)
|  |  |  | xxiii | warranty period offered on chair cushion cover (months)
|  |  |  | xxiv | Estimated lifespan of chair (years of service)
|  |  |  | xxv | Estimated lifespan of chair cushion (years of service)
|  |  |  | xxvi | Estimated lifespan of chair cushion cover (years of service)
|  |  |  | xxvii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxviii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 5.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 5.0 |  |  |  | Procedure Chairs | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 5.04 Ophthalmology / ENT Chairs
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Ophthalmology / ENT Chair
|  |  |  | i | An electrically powered device, designed to precisely position patients for ophthalmic and otolaryngologic examinations and procedures. Its adaptable design, including articulated headrests, adjustable armrests/arm boards, leg supports, and side rails, ensures optimal access and stability for delicate head and neck interventions, allowing optimal configuration across diverse clinical needs, supporting patients in either seated or reclined positions. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | a |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The chair must be free of protrusions that could damage the cushions or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Chairs should be designed that users cannot inadvertently come into contact with any moving "scissor action" lift mechanisms during use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For improved usability, chair controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise chairs that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Ophthalmology / ENT chair must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | ii | height-adjustable seat base,
|  |  |  | vi | include a headrest, preferably adjustable
collapsible/foldable/removable armrests,
|  |  | b |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the chair: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The chair's backrest and seat cushions must possess a therapeutic load capacity equal to or greater than the chair's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster existing cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The chair cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | seamless upholstery
|  |  | h |  | Chair cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | Castors, if fitted, must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under the maximum safe working load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | A central locking mechanism with directional steering operable via a single foot pedal is preferred. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Should individual casters be utilised, their braking and steering functions must be clearly marked and easily reachable from either side of the chair. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 5.06 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.5 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | All cables must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
[5 empty] | Chairs with electric movement must incorporate an integrated battery backup system capable of providing power for several complete full-range adjustments in the event of a power outage. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Equipment using batteries must have a visual low-battery warning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Chair controls must be designed to enable safe and intuitive adjustments for users, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Ophthalmology / ENT Chair offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Ophthalmology / ENT Chair offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | stationary/mobile chair
|  |  |  | vi | the external dimensions of the offered chair when fully lowered, Width x Length x Height (cm)
|  |  |  | vii | the maximum floor area required for the chair, L x W (cm)
|  |  |  | viii | weight of the offered equipment (kg)
|  |  |  | ix | SWL of offered chair (kg)
• therapeutic weight range (min/max kg)
|  |  |  | x | adjustable height range from floor level of the seat base (cm)
|  |  |  | xi | backrest dimensions, Width x Length (cm)
• backrest maximum incline/decline (in degrees)
|  |  |  | xii | headrest cushion dimensions, Width x Length (cm)
• describe the headrest cushion (adjustable, height, tilt)
|  |  |  | xiii | seat pan dimensions, Width x Length (cm)
• maximum seat tilt (in degrees)
• seat cushion height (cm)
leg section dimensions, Width x Length (cm)
• leg section maximum incline/decline (in degrees)
|  |  |  | xiv | Trendelenburg/reverse Trendelenburg articulation (in degrees, N/A)
|  |  |  | xv | castor size, diameter (cm, or N/A)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
• describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xvi | list all consumables available for offered chair, e.g. disposable fitted sheets, paper roll
|  |  |  | xvii | list all accessories available for offered chair (foot control, patient handset,
|  |  |  | xviii | list all options available for offered chair (CPR Lever. O2 bottle holder, IV pole, castors)
|  |  |  | xix | state the size of the display, H x W (cm) where applicable
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• equipment
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | warranty period offered on chair structural frame (months)
|  |  |  | xxvii | warranty period offered on chair electronics (months)
|  |  |  | xxviii | warranty period offered on chair cushion (months)
|  |  |  | xxix | warranty period offered on chair cushion cover (months)
|  |  |  | xx | Estimated lifespan of chair (years of service)
|  |  |  | xxxi | Estimated lifespan of chair cushion (years of service)
|  |  |  | xxxii | Estimated lifespan of chair cushion cover (years of service)
|  |  |  | xxxiii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxiv | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 5.05
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 5.0 |  |  |  | Procedure Chairs | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 5.05 Blood Collection Phlebotomy Chairs
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Blood Collection Phlebotomy Chair
|  |  |  | i | A non-powered device designed for safe, comfortable, and efficient blood collection for both patients and healthcare professionals. Its ergonomic, robust design ensures optimal patient positioning, easy phlebotomist access, and a hygienic environment in any clinical setting. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Blood collection phlebotomy chairs offered are encouraged to be listed on the ARTG. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | All equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Chair must have a safe working load of at least 200 kg, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | non-powered
|  |  |  | ii | all metal surfaces, excluding stainless steel, powder coated
|  |  |  | iii | armrests to optimise patient positioning, ensuring easy vein access and blood draws for phlebotomist from either side
|  |  |  | iv | armrests positioned at an optimal height for phlebotomist to easily draw blood
|  |  |  | v | include a step if seat height is excessive
|  |  | b |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Chair offered, the Respondent must submit:
|  |  |  | i | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Chair offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number, if applicable | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iii | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | iv | the external dimensions of the offered chair, Width x Length x Height (cm)
|  |  |  | v | weight of the offered equipment (kg)
|  |  |  | vi | SWL of offered chair (kg)
|  |  |  | vii | seat pan dimensions, Width x Length (cm)
• seat height from floor (cm)
|  |  |  | viii | armrests width (cm)
|  |  |  | ix | foot rest dimensions (cm)
|  |  |  | x | foldaway foot rest (Yes/No)
|  |  |  | xi | list other features
|  |  |  | xii | manufacturer of offered equipment
• country of origin
|  |  |  | xiii | Year 1st manufactured
|  |  |  | xiv | Estimated device's end-of-sales life (state the year)
|  |  |  | xv | warranty period of chair (months)
|  |  |  | xvi | Estimated lifespan of chair (years of service)
|  |  |  | xvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.01 Examination Plinths
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Examination Plinth
|  |  |  | i | A table featuring power adjustable height for minor medical examinations and treatments. It has a padded top covered with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Examination Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Examination Couch must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | single section
|  |  |  | v | equipped with four castors that are retractable or lockable when the couch is in use,
• that is foot operated and requires minimal force to engage/disengage castors
|  |  |  | vi | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Examination Couches, such as handsets, breath hole, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model examination couch offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each examination couch offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered examination couch (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | x | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xi | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiii | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xiv | list all accessories that are available offered examination couch
• handset (Yes/No)
• list all others
|  |  |  | xv | list all options available for offered examination couch
• breath hole and plug (Yes/No)
• list all others
|  |  |  | xvi | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xvii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• foot controller (switch)
• handset (if applicable)
|  |  |  | xviii | state the 'Drop Test' height for handset
|  |  |  | xix | length of handset/foot controller cable (m)
|  |  |  | xx | manufacturer of offered equipment
• country of origin
|  |  |  | xxi | Year 1st manufactured
|  |  |  | xxii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiii | warranty period offered on the frame (months)
|  |  |  | xxiv | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxv | Estimated lifespan of device (years of service)
|  |  |  | xxvi | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.02 Examination Table, 2-section
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Examination Table, 2-section
|  |  |  | i | A table featuring power adjustable height and backrest for minor medical examinations and treatments. It has a padded top covered with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Examination Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Examination Couch must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | electric motor for backrest adjustments
|  |  |  | iv | safe working load of at least 200 kg
|  |  |  | v | two sections
|  |  |  | vi | equipped with four castors that are retractable or lockable when the couch is in use
• that is foot operated and requires minimal force to engage/disengage castors
|  |  |  | vii | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Examination Couches, such as handsets, breath hole, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties.
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model examination couch offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each examination couch offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered examination couch (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• backrest section length (cm)
• torso/leg section length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | x | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xi | backrest maximum incline/decline (in degrees)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xv | list all accessories that are available offered examination couch
• handset (Yes/No)
• list all others
|  |  |  | xvi | list all options available for offered examination couch
• breath hole and plug (Yes/No)
• list all others
|  |  |  | xvii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xviii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xix | state the 'Drop Test' height for handset
|  |  |  | xx | length of handset/foot controller cable (m)
|  |  |  | xxi | manufacturer of offered equipment
• country of origin
|  |  |  | xxii | Year 1st manufactured
|  |  |  | xxiii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiv | warranty period offered on the frame (months)
|  |  |  | xxv | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxvi | Estimated lifespan of device (years of service)
|  |  |  | xxvii | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxviii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.03 Examination Table, 3-section
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Examination Table, 3-section
|  |  |  | i | A table featuring power adjustable height, backrest, and leg section for minor medical examinations and treatments. It has a padded top covered with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Examination Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Examination Couch must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | electric motors for profile adjustments
|  |  |  | iv | safe working load of at least 200 kg
|  |  |  | v | three sections
|  |  |  | vi | equipped with four castors that are retractable or lockable when the couch is in use
• that are foot operated and require minimal force to engage/disengage castors
|  |  |  | vii | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Examination Couches, such as handsets, breath hole, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model examination couch offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each examination couch offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered examination couch (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• backrest section length (cm)
• torso section length (cm)
• leg section length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | x | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xi | backrest maximum incline/decline (in degrees)
|  |  |  | xii | leg section maximum incline/decline (in degrees)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xvi | list all accessories that are available offered examination couch
• handset (Yes/No)
• list all others
|  |  |  | xvii | list all options available for offered examination couch
• breath hole and plug (Yes/No)
• list all others
|  |  |  | xviii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xix | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xx | state the 'Drop Test' height for handset
|  |  |  | xxi | length of handset/foot controller cable (m)
|  |  |  | xxii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxv | warranty period offered on the frame (months)
|  |  |  | xxvi | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxvii | Estimated lifespan of device (years of service)
|  |  |  | xxviii | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxix | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.04 Podiatry Examination Couch
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Podiatry Examination Couch
|  |  |  | i | A table featuring power adjustable height, backrest, and leg section facilitating foot examination, treatment, and minor procedures. It has a padded top covered with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Examination Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Examination Couch must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | electric motor for leg rest adjustments
|  |  |  | iv | electric motor for seat tilt adjustment
|  |  |  | v | safe working load of at least 200 kg
|  |  |  | vi | at least three sections
|  |  |  | vii | equipped with four castors that are retractable or lockable when the couch is in use
• that are foot operated and require minimal force to engage/disengage castors
|  |  |  | viii | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Examination Couches, such as handsets, breath hole, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Podiatry Examination Couches must not feature split leg rests. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  |  | v | arm rests
|  |  |  | vi | foot controller which includes switches for programmable memory positions
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model examination couch offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each examination couch offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered examination couch (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• backrest section length (cm)
• torso section length (cm)
• leg section length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | x | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xi | backrest maximum incline/decline (in degrees)
|  |  |  | xii | seat tilt maximum incline/decline (in degrees)
|  |  |  | xiii | leg section maximum incline/decline (in degrees)
|  |  |  | xiv | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xvi | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xvii | list all accessories that are available offered examination couch
• handset (Yes/No)
• list all others
|  |  |  | xviii | list all options available for offered examination couch
• breath hole and plug (Yes/No)
• arm rests (Yes/No)
- extending leg rests (Yes/No)
• foot controller with switches for programmable memory positions (Yes/No)
• list all others
|  |  |  | xix | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xx | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxi | state the 'Drop Test' height for handset
|  |  |  | xxii | length of handset/foot controller cable (m)
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | warranty period offered on the frame (months)
|  |  |  | xxvii | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxviii | Estimated lifespan of device (years of service)
|  |  |  | xxix | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxx | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.05
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.05 Tilt Tables
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Tilt Table
|  |  |  | i | A table designed to gradually and safely adjust from a horizontal (completely flat) to an almost vertical (fully upright) position featuring power height adjustments and safety straps to secure the individual during the procedure. It has a padded top cover with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Tilt Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Tilt Table must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height and tilt adjustment
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | handset to control table
|  |  |  | v | feature patient safety straps
|  |  |  | vi | include foldaway foot plate
|  |  |  | vii | battery back-up
|  |  | b |  | Respondents should offer additional options to Tilt Table, such as foot plate angular adjustment, paper roll holder, inclination indicator. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | The preferred braking system is a centrally controlled locking mechanism and directional steering, operable by a single foot pedal. Failing this, each castor must be equipped with an independent brake. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | foot switch
|  |  |  | ii | inclination indicator
|  |  |  | iii | paper roll holder
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | The tilt table must have a control mechanism, such as a dedicated switch on the handset, to enable an immediate and controlled return to the neutral (supine) position. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model tilt table offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model tilt table offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the patient platform (cm)
|  |  |  | ix | tilt angle range (in degrees)
• cycle time under SWL from supine to full incline (sec)
|  |  |  | x | adjustable foot plate for dorsi/planter flexion (Yes/No)
• foot plate angle range (in degrees)
|  |  |  | xi | no. of sections
• backrest maximum incline/decline (in degrees), if relevant
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xv | list all accessories that are available offered tilt table
|  |  |  | xvi | list all options available for offered tilt table
|  |  |  | xvii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xviii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xix | state the 'Drop Test' height for handset
|  |  |  | xx | manufacturer of offered equipment
• country of origin
|  |  |  | xxi | Year 1st manufactured
|  |  |  | xxii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiii | warranty period offered on the frame (months)
|  |  |  | xxiv | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxv | Estimated lifespan of device (years of service)
|  |  |  | xxvi | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.06
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.06 Traction Tables
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Traction Table
|  |  |  | i | A table designed to support a patient while a motorised traction unit, typically attached to a mast, applies controlled traction to the spinal column and associated back muscles. The table surface features adjustable sections that can be raised and lowered to conform to the patient’s spinal curvature and has a padded top cover with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Traction Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Traction Table must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | configure for both supine and prone positioning
|  |  |  | v | feature multiple sections
|  |  |  | vi | equipped with a mounting pedestal for a traction unit
|  |  |  | vii | equipped with four castors that are retractable or lockable when the table is in use
• that are foot operated and require minimal force to engage/disengage castors
|  |  |  | ix | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Traction Table, such as breath hole, paper roll holder, pelvic tilt section. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | pelvic tilt
|  |  |  | iv | paper roll holder
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Traction Table offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv' | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model traction table offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered table (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | x | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xi | no. of sections
• dimensions of each section (cm)
• backrest maximum incline/decline (in degrees)
• leg section maximum incline/decline (in degrees)
• tilting pelvic section (Yes/No)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xv | list all accessories that are available offered tilt table
|  |  |  | xvi | list all options available for offered tilt table
|  |  |  | xvii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xviii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xix | state the 'Drop Test' height for handset
|  |  |  | xx | length of handset/foot controller cable (m)
|  |  |  | xxi | manufacturer of offered equipment
• country of origin
|  |  |  | xxii | Year 1st manufactured
|  |  |  | xxiii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiv | warranty period offered on the frame (months)
|  |  |  | xxvi | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxvi | Estimated lifespan of device (years of service)
|  |  |  | xxvii | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxviii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.07
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.07 Ultrasound Exam Tables
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Ultrasound Exam Table
|  |  |  | i | A table designed to position and support a patient during a diagnostic ultrasound imaging procedure, while also allowing sonographer to perform scan ergonomically. It has a padded top covered with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Exam Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Exam Table must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | multi sectional
|  |  |  | v | equipped with castors that are lockable
|  |  |  | vi | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Examination Couches, such as handsets, breath hole, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  |  | v | vascular arm board
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Articulations
|  |  | a |  | Exam Table must be equipped with adjustable backrest. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Exam Table must be equipped with electrically stepless adjustments for Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Exam Table offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model exam table offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered exam table (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | x | no. of sections
• dimensions of each section (cm)
• backrest maximum incline/decline (in degrees)
• leg section maximum incline/decline, where applicable (in degrees)
|  |  |  | xi | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xii | electrically operated backrest (Yes/No)
|  |  |  | xiii | electrically operated leg section, where applicable (Yes/No)
|  |  |  | xiv | table cutout for Sonographer (Yes/No)
|  |  |  | xv | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xvi | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xvii | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xviii | list all accessories available for offered exam table
• handset (Yes/No)
• arm board (Yes/No)
• list all others
|  |  |  | xix | list all options available for offered exam table
• breath hole and plug (Yes/No)
• list all others
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | length of handset/foot controller cable (m)
|  |  |  | xxiv | manufacturer of offered equipment
• country of origin
|  |  |  | xxv | Year 1st manufactured
|  |  |  | xxvi | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvii | warranty period offered on the frame (months)
|  |  |  | xxviii | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxix | Estimated lifespan of device (years of service)
|  |  |  | xxx | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.08
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.08 Ultrasound Exam Tables, Echocardiogram
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Echocardiogram Exam Table
|  |  |  | i | A table designed to position and support a patient during a diagnostic echocardiogram procedure, while also allowing sonographer to perform scan ergonomically. It has a padded top covered with a stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Exam Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Exam Table must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | multi sectional
|  |  |  | v | features a cardiac cutout
|  |  |  | vi | equipped with castors that are lockable
|  |  |  | vii | provide sufficient under-couch clearance, to facilitate the safe and unimpeded use of patient lifters / standing aids
|  |  | b |  | Respondents should offer additional options to Examination Couches, such as handsets, breath hole, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | breath hole with plug
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  |  | v | adjustable back support for patient's right side
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Articulations
|  |  | a |  | Exam Table must be equipped with adjustable backrest. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Exam Table must be equipped with electrically stepless adjustments for Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model exam table offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model exam table offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered examination couch (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
|  |  |  | x | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xi | no. of sections
• dimensions of each section (cm)
• backrest maximum incline/decline (in degrees)
• leg section maximum incline/decline, where applicable (in degrees)
|  |  |  | xii | electrically operated backrest (Yes/No)
|  |  |  | xiii | electrically operated leg section, where applicable (Yes/No)
|  |  |  | xiv | table cutout for Sonographer (Yes/No)
|  |  |  | xv | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xvi | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xvii | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xviii | list all accessories available for offered exam table
• handset (Yes/No)
• list all others
|  |  |  | xix | list all options available for offered exam table
• breath hole and plug (Yes/No)
• adjustable back support (Yes/No)
• list all others
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | length of handset/foot controller cable (m)
|  |  |  | xxiv | manufacturer of offered equipment
• country of origin
|  |  |  | xxv | Year 1st manufactured
|  |  |  | xxvi | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvii | warranty period offered on the frame (months)
|  |  |  | xxviii | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxix | Estimated lifespan of device (years of service)
|  |  |  | xxx | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 6.09
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 6.0 |  |  |  | Examination Couches | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 6.09 Ophthalmology / ENT Examination Couches
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Ophthalmology / ENT Examination Couch
|  |  |  | i | A table designed specifically to precisely position patients for ophthalmic and otolaryngologic examinations and procedures. It has a padded top cover with the stain resistant material. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Examination Couches must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Examination Couch must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mains powered
|  |  |  | ii | electric motor for height adjustment
|  |  |  | iii | electric motors for profile adjustments
|  |  |  | iv | safe working load of at least 200 kg
|  |  |  | v | three sections
|  |  |  | vi | include a headrest, preferably adjustable
|  |  | b |  | Respondents should offer additional options to Examination Couches, such as handsets, castors, armrest, paper roll holder, battery back-up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Cushions
|  |  | a |  | The following requirements apply to all cushions on the examination couch: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | ii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  |  | iii | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  | b |  | Cushion padding prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The cushion/s must possess a therapeutic load capacity equal to or greater than the examination couch's safe working load. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Suppliers are permitted to reupholster cushions. When doing so, the cushion's foam and all internal padding must be completely replaced. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | HSV classifies replacement cushions as an essential spare part. As such, the Respondent must adhere to all service and delivery requirements specified within this tender. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The upholstery on the underside of the cushion must be secured using a fastening method that creates a smooth, continuous surface free of exposed fasteners, staples, or sharp protrusions, thereby ensuring effective infection control. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Cushion Covers
|  |  | g |  | The cushion covers must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | stain resistant
|  |  |  | iii | biocompatible
|  |  |  | iv | static dissipative
|  |  |  | v | resistant to degradation by hospital-grade disinfectants
|  |  |  | vi | exhibits multi-directional stretch and is designed to minimise wrinkling
|  |  |  | vii | flame-retardant durable material
|  |  | h |  | Cushion covers should incorporate antibacterial properties. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | i |  | Within 24 months of the contract commencement date, the Respondent must submit a comprehensive pressure-mapping report for the tendered product. The report must include: | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 6.10 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | handset
|  |  |  | ii | armrests
|  |  |  | iii | battery back-up
|  |  |  | iv | paper roll holder
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must list all purchasable consumables for the offered equipment in category 6.11 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | To ensure a competitive and diverse supply chain, third-party providers are encouraged to offer a wide range of essential consumables for examination couches, such as disposable paper rolls and protective covers. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Examination Couch must be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments by the clinician, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling, and should have a foot controller cord of adequate length to facilitate unrestricted operational access from either side of the plinth. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must comply with the mechanical strength requirements specified of AS/NZS IEC 60601.1:2015, Section 15.3, which includes drop testing, stability testing, and impact testing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model examination couch offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each examination couch offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered examination couch, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered examination couch (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | base compatible for Patient Lifters (Yes/No)
• under-couch clearance for patient lifter/standing aid, floor to base underside height (cm)
|  |  |  | ix | patient (therapy) table dimensions, Width x Length (cm)
• frame dimensions to which the upholstered cushions are attached, Width x full Length (cm)
• backrest section length (cm)
• torso section length (cm)
• leg section length (cm)
|  |  |  | x | adjustable height range from floor level of the patient platform (cm)
|  |  |  | xi | headrest cushion dimensions, Width x Length (cm)
• describe the headrest cushion (adjustable, height, tilt)
|  |  |  | xii | backrest maximum incline/decline (in degrees)
|  |  |  | xiii | leg section maximum incline/decline (in degrees)
|  |  |  | xiv | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xvi | list all consumables available for offered exam table, e.g. disposable fitted sheets, paper roll
|  |  |  | xvii | list all accessories that are available offered examination couch
• handset (Yes/No)
• list all others
|  |  |  | xviii | list all options available for offered examination couch
• breath hole and plug (Yes/No)
• list all others
|  |  |  | xix | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. SWL lifts)
• gauge indicating remaining battery capacity
|  |  |  | xx | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxi | state the 'Drop Test' height for handset
|  |  |  | xxii | length of handset/foot controller cable (m)
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | warranty period offered on the frame (months)
|  |  |  | xxvii | warranty period offered on the electronics/upholstery (months)
|  |  |  | xxviii | Estimated lifespan of device (years of service)
|  |  |  | xxix | Estimated lifespan of upholstered cushions (years of service)
|  |  |  | xxx | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.01 Hospital Stretchers, Manual
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Hospital Stretcher, Manual
|  |  |  | i | A manually operated, wheeled device providing a mobile platform for transporting recumbent patients within a healthcare facility. Equipped with an adjustable backrest, side rails, and holders for IV fluid bags and oxygen delivery equipment; it may also serve as a treatment table for critical care areas like emergency, recovery or labour rooms, or as a bed in specialised medical units. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretchers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Stretchers must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The stretcher frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the stretcher's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Stretchers must be equipped with ergonomic handgrips to enhance caregiver safety during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Stretchers designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Participating Health Services will prioritise stretchers that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 250 kg
|  |  |  | ii | include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator, where the installed components must not impede the stretcher's function
|  |  |  | iii | an integrated IV pole mounting point
|  |  |  | iv | centrally controlled locking mechanism and directional steering
|  |  |  | v | feature integrated, non-removable, anti-pinch side rails
|  |  | b |  | The mattress support platform must have an adjustable height, and the dimensions for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress at least 180 cm in length and 60 cm in width
|  |  | c |  | Respondents should offer additional options to Stretchers, such as CPR release lever, secure and convenient attachment points for drainage collection devices, fifth wheel, frame-integrated restraint points, patient belongings and medical record holders. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The stretcher's castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with stretchers must meet all mandatory criteria in Category 7.08. Additionally, Respondents must complete Category 7.08 for any offered stretchers that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 7.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | c |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | integrated patient belongings storage
|  |  |  | ii | medical records and chart holder
|  |  |  | iii | radiolucent backrest with integrated upright chest X-ray cassette holder
|  |  |  | iv | equipment mounting solutions
|  |  |  | v | secured drainage collection device attachments
|  |  |  | vi | arm boards for cannulation
|  |  |  | vii | restraint (hand-cuff) attachment points
|  |  | e |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Articulations
|  |  | a |  | Stretchers offered must be equipped with adjustable profile functions, including backrest, and adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer stretchers with articulating knee Gatch (knee bend feature), and features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Side Rails
|  |  | a |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Collapsible side rails must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The following side rail options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | split side rails
|  |  |  | ii | flush-folding side rails
|  |  |  | iii | side rails without vertical bars
|  |  |  | iv | designed to safely and effectively support the patient during ingress and egress from the stretcher.
|  | 3.6 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Stretchers must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the stretcher. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | vi | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• stretcher includes a mattress (Yes/No)
• manufacturer's recommended mattress, make, model
|  |  |  | x | no. of mattress platform sections listing each (backrest section, etc.)
• includes X-ray translucent backrest with cassette carrier (Yes/No)
• includes patient anti-migration feature (Yes/No)
• includes articulating knee Gatch, knee bend feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for stretcher to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xi | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | list all accessories that are available offered bed (IV pole, cylinder holder, arm board)
|  |  |  | xvi | list all options available for offered bed
• integrated patient belongings storage (Yes/No)
• radiolucent backrest with X-ray cassette holder (Yes/No)
• medical records and chart holder (Yes/No)
• equipment mounting solutions (Yes/No)
• secured drainage collection device attachments (Yes/No)
• fifth wheel (Yes/No)
• list all others
|  |  |  | xvii | manufacturer of offered equipment
• country of origin
|  |  |  | xviii | Year 1st manufactured
|  |  |  | xix | Estimated device's end-of-sales life (state the year)
|  |  |  | xx | warranty period offered on stretcher (months)
|  |  |  | xxi | warranty period offered on stretcher structural frame (months)
|  |  |  | xxii | Estimated lifespan of device (years of service)
|  |  |  | xxiii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.02 Hospital Stretchers, Powered
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Hospital Stretcher, Powered
|  |  |  | i | A battery-powered, wheeled device providing a mobile platform for transporting recumbent patients within a healthcare facility. For improve manual handling and staff safety is equipped with powered variable height and adjustable backrest and knee break positions, also includes side rails and integrated holders for IV pole and oxygen delivery equipment. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretchers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Stretchers must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The stretcher frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the stretcher's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Stretchers must be equipped with ergonomic handgrips to enhance caregiver safety during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Stretchers designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | For improved usability, stretcher controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Participating Health Services will prioritise stretchers that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | electric motors for adjusting the mattress support platform height
|  |  |  | ii | electric motors for the manipulation of stretcher profile articulations
|  |  |  | iii | safe working load of at least 250 kg
|  |  |  | iv | include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator, where the installed components must not impede the stretcher's function
|  |  |  | v | an integrated IV pole mounting point
|  |  |  | vi | centrally controlled locking mechanism and directional steering
|  |  |  | vii | feature integrated, non-removable, anti-pinch side rails
|  |  | b |  | The mattress support platform must have an adjustable height, and the dimensions for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress at least 180 cm in length and 60 cm in width
|  |  | c |  | Respondents should offer additional options to Stretchers, such as secure and convenient attachment points for drainage collection devices, fifth wheel, frame-integrated restraint points, patient belongings and medical record holders. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The stretcher's castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with stretchers must meet all mandatory criteria in Category 7.08. Additionally, Respondents must complete Category 7.08 for any offered stretchers that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 7.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | b |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | patient weigh features
|  |  |  | ii | integrated patient belongings storage
|  |  |  | iii | medical records and chart holder
|  |  |  | iv | radiolucent backrest with integrated upright chest X-ray cassette holder
|  |  |  | v | equipment mounting solutions
|  |  |  | vi | secured drainage collection device attachments
|  |  |  | vii | knee break (Fowler position)
arm boards for cannulation
restraint (hand-cuff) attachment points
|  |  | e |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Stretchers should be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading, be equipped with a liquid-tight seal at the equipment entry point, and feature an integrated mains cable storage system; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Articulations
|  |  | a |  | Stretchers offered must be equipped with adjustable profile functions, including backrest, and adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer stretchers with articulating knee Gatch (knee bend feature), and features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Side Rails
|  |  | a |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Collapsible side rails must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The following side rail options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | split side rails
|  |  |  | ii | flush-folding side rails
|  |  |  | iii | side rails without vertical bars
|  |  |  | iv | designed to safely and effectively support the patient during ingress and egress from the stretcher.
|  | 3.7 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Stretchers must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the stretcher. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• manufacturer's recommended mattress, make, model
|  |  |  | x | no. of mattress platform sections listing each (backrest section, etc.)
• includes X-ray translucent backrest with cassette carrier (Yes/No)
• includes patient anti-migration feature (Yes/No)
• includes articulating knee Gatch, knee bend feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for stretcher to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xi | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | list all accessories that are available offered stretcher (IV pole, cylinder holder, external charger for secondary battery, secondary battery)
|  |  |  | xvi | list all options available for offered stretcher
• integrated patient belongings storage (Yes/No)
• medical records and chart holder (Yes/No)
• equipment mounting solutions (Yes/No)
• radiolucent backrest with X-ray cassette holder (Yes/No)
• secured drainage collection device attachments (Yes/No)
• fifth wheel (Yes/No)
• list all others
|  |  |  | xvii | state the size of the display, H x W (cm) where applicable
|  |  |  | xviii | stretcher connects to a 240V power outlet? (Yes/No)
|  |  |  | xix | removable secondary backup battery that staff can easily swap out (Yes/No)
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | warranty period offered on stretcher structural frame (months)
|  |  |  | xxvii | warranty period offered on stretcher electronics (months)
|  |  |  | xxviii | warranty period offered on stretcher handset & cables (months)
|  |  |  | xxix | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.03 Hospital Stretchers, Motorised-Drive
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Hospital Stretcher, Motorised-Drive
|  |  |  | i | A battery-powered hospital stretcher (Cat 8.02) with a motorised drive system, controlled by a joystick or buttons, enabling staff to effortlessly move patients across any hospital terrain. It may also include the full suite of features expected on a battery-powered hospital stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretchers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Stretchers must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  |  | b |  | Stretchers equipped with removable siderails should have clearly visible and durable markers indicating the correct placement of the siderails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The stretcher frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the stretcher's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The bed end must possess a surface or structural element compatible with the installation of a bracket for an alcohol-based hand rub dispenser. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Stretchers must be equipped with ergonomic handgrips to enhance caregiver safety during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Stretchers designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | For improved usability, stretcher controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Participating Health Services will prioritise stretchers that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | fifth wheel with an integrated electric motor
|  |  |  | ii | should have electric motors for adjusting the mattress support platform height
|  |  |  | iii | should have electric motors for the manipulation of stretcher profile articulations
|  |  |  | iv | safe working load of at least 250 kg
|  |  |  | v | include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator, where the installed components must not impede the stretcher's function
|  |  |  | vi | an integrated IV pole mounting point
|  |  |  | vii | centrally controlled locking mechanism and directional steering
|  |  |  | viii | feature integrated, non-removable, anti-pinch side rails
|  |  | b |  | The mattress support platform must have an adjustable height, and the dimensions for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress at least 180 cm in length and 60 cm in width
|  |  | c |  | Respondents should offer additional options to Stretchers, such as secure and convenient attachment points for drainage collection devices, frame-integrated restraint points, patient belongings and medical record holders. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The stretcher's castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with stretchers must meet all mandatory criteria in Category 7.08. Additionally, Respondents must complete Category 7.08 for any offered stretchers that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 7.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | c |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | patient weigh features
|  |  |  | ii | integrated patient belongings storage
|  |  |  | iii | equipment transport "shelf"
|  |  |  | iv | medical records and chart holder
|  |  |  | v | radiolucent backrest with integrated upright chest X-ray cassette holder
|  |  |  | vi | equipment mounting solutions
|  |  |  | vii | secured drainage collection device attachments
|  |  |  | viii | knee break (Fowler position)
|  |  |  | ix | arm boards for cannulation
|  |  |  | x | restraint (hand-cuff) attachment points
|  |  | e |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Stretchers should be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading, be equipped with a liquid-tight seal at the equipment entry point, and feature an integrated mains cable storage system; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Articulations
|  |  | a |  | Stretchers offered must be equipped with adjustable profile functions, including backrest, and adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents should offer stretchers with articulating knee Gatch (knee bend feature), and features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Side Rails
|  |  | a |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Collapsible side rails must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The following side rail options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | split side rails
|  |  |  | ii | flush-folding side rails
|  |  |  | iii | side rails without vertical bars
|  |  |  | iv | designed to safely and effectively support the patient during ingress and egress from the stretcher.
|  | 3.7 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Stretchers must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the stretcher. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• manufacturer's recommended mattress, make, model
|  |  |  | x | no. of mattress platform sections listing each (backrest section, etc.)
• includes X-ray translucent backrest with cassette carrier (Yes/No)
• includes patient anti-migration feature (Yes/No)
• includes articulating knee Gatch, knee bend feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for stretcher to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xi | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | list all accessories that are available offered stretcher (IV pole, cylinder holder, external charger for secondary battery, secondary battery)
|  |  |  | xvi | list all options available for offered stretcher
• integrated patient belongings storage (Yes/No)
• medical records and chart holder (Yes/No)
• equipment mounting solutions (Yes/No)
• radiolucent backrest with X-ray cassette holder (Yes/No)
• secured drainage collection device attachments (Yes/No)
• motorised fifth wheel (Yes/No)
• list all others
|  |  |  | xvii | state the size of the display, H x W (cm) where applicable
|  |  |  | xviii | stretcher connects to a 240V power outlet? (Yes/No)
|  |  |  | xix | removable secondary backup battery that staff can easily swap out (Yes/No)
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• how long does battery last on motorised drive (km)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | warranty period offered on stretcher structural frame (months)
|  |  |  | xxvii | warranty period offered on stretcher electronics (months)
|  |  |  | xxviii | warranty period offered on stretcher handset & cables (months)
|  |  |  | xxix | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.04 Hospital Stretchers, Radiographic
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Hospital Stretcher, Radiographic
|  |  |  | i | A wheeled device featuring a radiolucent platform to support and transport recumbent patients during x-ray imaging. Commonly used in emergency departments for initial diagnostic imaging of trauma patients, and may include accessories like cassette-support attachments. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | A wheeled device featuring a radiolucent platform designed to support and transport recumbent patients during mobile C-arm fluoroscopy procedures. These trolleys are critical for enabling precise patient positioning and unimpeded x-ray images in various clinical settings, including operating rooms, interventional radiology suites, and emergency departments. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | Product Requirements
|  |  | a |  | Radiographic Hospital Stretchers must meet all mandatory criteria within one of the following subcategories, unless clause 2.0(b) or 2.0(c) stipulates a specific exception: - Subcategory 7.01: Manual Hospital Stretchers - Subcategory 7.02: Powered Hospital Stretchers - Subcategory 7.03: Motorised-drive Hospital Stretchers Respondents are required to complete the relevant subcategory for the offered stretcher. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Radiographic Hospital Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | radiolucent mattress platform and mattress
|  |  |  | ii | mobile X-ray cassette holders able to securely accommodate both European and US cassette sizes, allowing for full length supine imaging
|  |  |  | iii | retractable side rails must be able to drop below the mattress level
|  |  | c |  | Fluoroscopy Procedural Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | radiolucent full length deck and mattress
|  |  |  | ii | unrestricted full-body access during imaging procedures
|  |  |  | iii | sufficient clearance underneath stretcher to accommodate C-arm
|  |  |  | iv | retractable side rails must be able to drop below the mattress level
|  |  |  | v | safe working load of at least 200 kg
|  |  | d |  | Offers for bariatric models are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• manufacturer's recommended mattress, make, model
|  |  |  | x | no. of mattress platform sections listing each (backrest section, etc.)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for stretcher to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xi | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | list all accessories that are available offered stretcher (IV pole, cylinder holder, external charger for secondary battery, secondary battery)
|  |  |  | xvi | list all options available for offered stretcher
• secured drainage collection device attachments (Yes/No)
• fifth wheel (Yes/No)
• motorised-drive (Yes/No)
• list all others
|  |  |  | xvii | Radiographic (Trauma) Stretcher (Yes/No)
|  |  |  | xviii | Fluoroscopy Stretcher (Yes/No)
• fluoroscopy window, Width x Length (cm)
• C-arm access width (cm)
|  |  |  | xix | stretcher connects to a 240V power outlet? (Yes/No)
|  |  |  | xx | removable secondary backup battery that staff can easily swap out (Yes/No)
|  |  |  | xxi | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxiii | state the 'Drop Test' height for handset
|  |  |  | xxiv | manufacturer of offered equipment
• country of origin
|  |  |  | xxv | Year 1st manufactured
|  |  |  | xxvi | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvii | warranty period offered on stretcher structural frame (months)
|  |  |  | xxviii | warranty period offered on stretcher electronics (months)
|  |  |  | xxix | warranty period offered on stretcher handset & cables (months)
|  |  |  | xxx | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxi | Estimated lifespan of device (years of service)
|  |  |  | xxxii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxiii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.05
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.05 Hospital Stretchers, MRI Conditional
|  | 1 |  |  | Intended Purpose
|  |  | a |  | MRI Conditional Hospital Stretcher
|  |  |  | i | A stretcher specifically engineered using materials and components proven to be safe and to pose no known hazards within a designated MRI environment. It’s designed to operate effectively without interfering with the MRI scanner’s function or jeopardising the safety of patients and staff, provided the specified magnetic field strengths and conditions are met. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretchers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Stretchers must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The stretcher frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the stretcher's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Stretchers designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services will prioritise stretchers that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | under specific conditions, safe for use in the MR environment
|  |  |  | ii | constructed from non-ferromagnetic materials, including fasteners
|  |  |  | iii | safe working load of at least 250 kg
|  |  |  | iv | adjustable backrest
|  |  |  | v | an integrated IV pole mounting point
|  |  | b |  | The mattress support platform must have the dimensions for a mattress at least 180 cm in length and 60 cm in width | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents should offer additional options to Stretchers, such as height adjustable mattress platform. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The stretcher's castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Of the four castors, at least two must incorporate a brake. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with stretchers must meet all mandatory criteria in Category 7.08. Additionally, Respondents must complete Category 7.08 for any offered stretchers that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 7.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | c |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | height adjustable mattress platform
|  |  |  | ii | gas-strut adjustable backrest
|  |  |  | iii | side rails
|  |  |  | iv | Trendelenburg positioning
|  |  |  | v | oxygen-cylinder holder
|  |  | e |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | ii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iii | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | iv | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | v | weight of the offered equipment (kg)
|  |  |  | vi | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | vii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | viii | mattress platform dimensions (without extensions), Width x Length (cm)
• manufacturer's recommended mattress, make, model
|  |  |  | ix | no. of mattress platform sections listing each (backrest section, etc.)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
|  |  |  | x | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
|  |  |  | xi | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiii | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xiv | list all accessories that are available offered stretcher (IV pole, cylinder holder, external charger for secondary battery, secondary battery)
|  |  |  | xv | list all options available for offered stretcher
• height adjustable mattress platform (Yes/No)
• backrest gas-spring adjusted (Yes/No)
• Trendelenburg positioning (Yes/No)
• oxygen-cylinder holder (Yes/No)
• list all others
|  |  |  | xvi | manufacturer of offered equipment
• country of origin
|  |  |  | xvii | Year 1st manufactured
|  |  |  | xviii | Estimated device's end-of-sales life (state the year)
|  |  |  | xix | warranty period offered on stretcher structural frame (months)
|  |  |  | xx | warranty period offered on stretcher electronics (months)
|  |  |  | xxi | warranty period offered on stretcher handset & cables (months)
|  |  |  | xxii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxiii | Estimated lifespan of device (years of service)
|  |  |  | xxiv | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxv | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.06
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.06 Hospital Stretcher Chairs
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Hospital Stretcher Chair
|  |  |  | i | A battery-powered hospital stretcher (Cat 8.02) seamlessly converted from a stretcher to a chair and back again eliminating the need to transfer patients. It’s design allows for optimal patient support and positioning, offering both fully reclined and upright seated postures. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• has a longer wheelbase than a Procedure Chair (Cat 6.03), of approximately 1.5m
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretchers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Stretchers must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The stretcher frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the stretcher's frame, must be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Stretchers must be equipped with ergonomic handgrips to enhance caregiver safety during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Stretchers designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | For improved usability, stretcher controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Participating Health Services will prioritise stretchers that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | battery-powered electric motors for adjusting the mattress support platform height
|  |  |  | ii | battery-powered electric motors for the manipulation of stretcher profile articulations
|  |  |  | iii | safe working load of at least 250 kg
|  |  |  | iv | knee break to offer Fowler positioning
|  |  |  | v | include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator, where the installed components must not impede the stretcher's function
|  |  |  | vi | an integrated IV pole mounting point
|  |  |  | vii | centrally controlled locking mechanism and directional steering
|  |  |  | viii | feature integrated, non-removable, anti-pinch side rails
|  |  | b |  | The mattress support platform must have an adjustable height, and the dimensions for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | patient support surface of at least 180 cm in length and 60 cm in width
|  |  | c |  | Respondents should offer additional options to stretchers, such as secure and convenient attachment points for drainage collection devices, fifth wheel, frame-integrated restraint points, patient belongings and medical record holders. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The stretcher's castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with stretchers must meet all mandatory criteria in Category 7.08. Additionally, Respondents must complete Category 7.08 for any offered stretchers that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 7.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | c |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | patient weigh features
|  |  |  | ii | integrated patient belongings storage
|  |  |  | iii | medical records and chart holder
|  |  |  | iv | radiolucent backrest with integrated upright chest X-ray cassette holder
|  |  |  | v | equipment mounting solutions
|  |  |  | vi | secured drainage collection device attachments
|  |  |  | vii | fifth wheel
|  |  |  | viii | motorised drive
|  |  |  | ix | patient control panel/handset
|  |  |  | x | arm boards for cannulation
|  |  | e |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Stretchers should be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading, be equipped with a liquid-tight seal at the equipment entry point, and feature an integrated mains cable storage system; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Articulations
|  |  | a |  | Stretchers offered must be equipped with adjustable profile functions, including backrest, and adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must offer stretchers with articulating knee Gatch (knee bend feature), and features to mitigate patient migration when the backrest is elevated. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Side Rails
|  |  | a |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Collapsible side rails must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The following side rail options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | split side rails
|  |  |  | ii | flush-folding side rails
|  |  |  | iii | side rails without vertical bars
|  |  |  | iv | designed to safely and effectively support the patient during ingress and egress from the stretcher.
|  | 3.7 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Stretchers must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the stretcher. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed end/s must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed end/s must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• manufacturer's recommended mattress, make, model
|  |  |  | x | no. of mattress platform sections listing each (backrest section, etc.)
• includes X-ray translucent backrest with cassette carrier (Yes/No)
• includes patient anti-migration feature (Yes/No)
• includes articulating knee Gatch, knee bend feature (Yes/No)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for stretcher to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xi | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xiv | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xv | list all accessories that are available offered stretcher (IV pole, cylinder holder, external charger for secondary battery, secondary battery)
|  |  |  | xvi | list all options available for offered stretcher
• integrated patient belongings storage (Yes/No)
• medical records and chart holder (Yes/No)
• equipment mounting solutions (Yes/No)
• radiolucent backrest with X-ray cassette holder (Yes/No)
• secured drainage collection device attachments (Yes/No)
• fifth wheel (Yes/No)
• list all others
|  |  |  | xvii | state the size of the display, H x W (cm) where applicable
|  |  |  | xviii | stretcher connects to a 240V power outlet? (Yes/No)
|  |  |  | xix | removable secondary backup battery that staff can easily swap out (Yes/No)
|  |  |  | xx | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxi | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xxii | state the 'Drop Test' height for handset
|  |  |  | xxiii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiv | Year 1st manufactured
|  |  |  | xxv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvi | warranty period offered on stretcher structural frame (months)
|  |  |  | xxvii | warranty period offered on stretcher electronics (months)
|  |  |  | xxviii | warranty period offered on stretcher handset & cables (months)
|  |  |  | xxix | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxx | Estimated lifespan of device (years of service)
|  |  |  | xxxi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.07
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.07 Hospital Stretchers, Ophthalmology / ENT
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Ophthalmology / ENT Hospital Stretcher
|  |  |  | i | A hospital stretcher, offered in battery-powered or hydraulic models, designed to meet the unique demands of ophthalmology and ENT departments, providing optimal access for delicate head and neck procedures. It provides a stable and adaptable platform crucial for examinations, minor procedures, and patient transport within these specialised medical fields. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretchers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Stretchers must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.4 |  |  | Occupational Health and Safety
|  |  | a |  | The stretcher frame must be free of protrusions that could damage the mattress or cause injury, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | designed to reduce the risk of tripping or falling.
|  |  | b |  | Impact-resistant corner protectors, protruding past the stretcher's frame, should be fitted on all four corners. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Stretchers must be equipped with ergonomic handgrips to enhance caregiver safety during patient transport. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | All metal surfaces, excluding stainless steel, must be powder coated. The supplier should provide a variety of colour options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Stretchers designs without scissor lift mechanisms or other pinch/crush hazards are preferred, especially where children may access the area under the stretcher. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | For improved usability, stretcher controls should feature intuitive layouts and clearly labelled functions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Participating Health Services will prioritise stretchers that facilitate easy movement and operation. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Stretcher must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | include a headrest, preferably adjustable
|  |  |  | iii | safe working load of at least 200 kg
|  |  |  | iv | include an integrated mounting point for a transportable medical gas cylinder (up to 185mm diameter) and its regulator, where the installed components must not impede the stretcher's function
|  |  |  | v | an integrated IV pole mounting point
|  |  |  | vi | centrally controlled locking mechanism and directional steering
|  |  |  | vii | feature integrated, non-removable, anti-pinch side rails
|  |  | b |  | The mattress support platform must have an adjustable height, and the dimensions for a | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | mattress at least 180 cm in length and 60 cm in width
|  |  | c |  | Respondents should offer additional options to Stretchers, such as secure and convenient attachment points for drainage collection devices, fifth wheel, patient belongings and medical record holders. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 250kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Platform
|  |  | a |  | The mattress support platform must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | use only solid panel mattress supports and not mesh,
|  |  |  | ii | be labelled with the recommended mattress dimensions,
|  |  |  | iii | ensure the mattress is held securely in place, without slippage or deformation, throughout all possible adjustable positions, and
|  |  |  | iv | be equipped with a horizontal level, or a self-leveling mechanism that restores the support platform to a horizontal plane after any adjustments,
|  | 3.2 |  |  | Castors and Braking Systems
|  |  | a |  | The stretcher's castors must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The braking system for beds must incorporate a centrally controlled locking mechanism and directional steering, both operable via a single foot pedal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with stretchers must meet all mandatory criteria in Category 7.08. Additionally, Respondents must complete Category 7.08 for any offered stretchers that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 7.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  | c |  | An optional IV pole with multiple prongs must be available. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | additional head supports
|  |  |  | ii | integrated patient belongings storage
|  |  |  | iii | medical records and chart holder
|  |  |  | iv | radiolucent backrest with integrated upright chest X-ray cassette holder
|  |  |  | v | equipment mounting solutions
|  |  |  | vi | secured drainage collection device attachments
|  |  | e |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | Stretchers should be powered by a mains connection. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | c |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading, be equipped with a liquid-tight seal at the equipment entry point, and feature an integrated mains cable storage system; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | All cable must be routed and secured to eliminate any risk of crushing by moving components. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | f |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Articulations
|  |  | a |  | Stretchers offered must be equipped with adjustable profile functions, including backrest, knee break, and adjustments for Trendelenburg and reverse Trendelenburg positioning. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.6 |  |  | Side Rails
|  |  | a |  | Side rails must comply with IEC 60601-2-52. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Collapsible side rails must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | maintain secure positioning and intended function during operation
|  |  |  | ii | be intuitive and safe to operate by a single person, requiring minimal force
|  |  |  | iii | incorporate a robust positive locking mechanism that requires a deliberate action to disengage
|  |  |  | iv | when collapsed, the uppermost surface positioned at or below the top of an uncompressed mattress
|  |  |  | v | when collapsed, designed to fall vertically or slide and remain fully contained beneath the bed frame
|  |  |  | vi | accommodate the full range of pressure-relieving mattresses without compromising the security, functionality, or intended gap dimensions of the side rails in their raised position
|  |  |  | vii | be designed to eliminate foreseeable risks of patient entrapment or injury in all operational and collapsed states
|  |  |  | viii | maintain intended height and functionality relative to the mattress surface throughout the full range of bed platform adjustments, including height, articulation (e.g., Trendelenburg), and any other positional changes
|  |  |  | ix | prevent unintentional removal by staff during routine use, cleaning, linen changes, or maintenance
|  |  | c |  | Side rail release mechanisms must be easily accessible and allow for one-handed operation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Side rail release mechanisms should be equipped with dampeners for controlled deceleration during lowering. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The following side rail options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | split side rails
|  |  |  | ii | flush-folding side rails
|  |  |  | iii | side rails without vertical bars
|  |  |  | iv | designed to safely and effectively support the patient during ingress and egress from the stretcher.
|  | 3.7 |  |  | CPR Release
|  |  | a |  | Upon activation of the CPR release mechanism, the backrest must rapidly and safely return to a horizontal position ensuring a controlled descent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | preferably within 10 seconds. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Stretchers must have a CPR release mechanism that is clearly identifiable and easily accessible from either side of the stretcher. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Access to the CPR release lever must not be impeded by the side rails. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Bed ends must be removable by one person within 15 seconds, without the use of tools. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Bed ends must be lightweight for easy removal and handling by a single person | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file size less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stretcher offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | the external dimensions of the offered stretcher, Width x Length (cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | SWL of offered stretcher (kg)
• therapeutic weight range (min/max kg)
|  |  |  | viii | adjustable height range from floor level of the mattress platform (cm)
|  |  |  | ix | mattress platform dimensions (without extensions), Width x Length (cm)
• manufacturer's recommended mattress, make, model
|  |  |  | x | headrest cushion dimensions, Width x Length (cm)
• describe the headrest cushion (adjustable, height, tilt)
|  |  |  | xi | no. of mattress platform sections listing each (backrest section, etc.)
• includes X-ray translucent backrest with cassette carrier (Yes/No)
• includes articulating knee break feature (Yes/No)
• maximum Trendelenburg/reverse Trendelenburg articulation (degrees)
• time for stretcher to return to horizontal plane from a 60° backrest incline upon activation of the CPR release mechanism (sec)
|  |  |  | xii | describe the side rails (full, 3/4 length, split)
• height of side rail from mattress platform (cm)
• maximum width including side rails (cm)
• side rail can be used as an assistive support for exiting (Yes/No)
- side rail assistive support SWL (kg)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• describe steering (2 wheels, all wheels)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | stretcher frame equipped with caregiver handgrips (Yes/No)
|  |  |  | xvi | list all accessories that are available offered stretcher (IV pole, cylinder holder, external charger for secondary battery, secondary battery)
|  |  |  | xvii | list all options available for offered stretcher
• integrated patient belongings storage (Yes/No)
• medical records and chart holder (Yes/No)
• equipment mounting solutions (Yes/No)
• radiolucent backrest with X-ray cassette holder (Yes/No)
• secured drainage collection device attachments (Yes/No)
• head holder (Yes/No)
• fifth wheel (Yes/No)
• list all others
|  |  |  | xviii | state the size of the display, H x W (cm) where applicable
|  |  |  | xix | stretcher connects to a 240V power outlet? (Yes/No)
|  |  |  | xx | removable secondary backup battery that staff can easily swap out (Yes/No)
|  |  |  | xxi | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xxii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• controller
• handset (if applicable)
|  |  |  | xiii | state the 'Drop Test' height for handset
|  |  |  | xxiv | manufacturer of offered equipment
• country of origin
|  |  |  | xxv | Year 1st manufactured
|  |  |  | xxvi | Estimated device's end-of-sales life (state the year)
|  |  |  | xxvii | warranty period offered on stretcher structural frame (months)
|  |  |  | xxviii | warranty period offered on stretcher electronics (months)
|  |  |  | xxix | warranty period offered on stretcher handset & cables (months)
|  |  |  | xxx | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxi | Estimated lifespan of device (years of service)
|  |  |  | xxxii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxiii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 7.08
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 7.0 |  |  |  | Patient Stretchers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 7.08 Stretcher Mattresses
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Mattress for Patient Stretchers
|  |  |  | i | A large, fabric-covered case field with a resilient material, such as foam rubber, designed to be used on a hospital stretcher to provide patients a comfortable surface and prevent pressure injuries during short-term stays, and transport within a healthcare facility. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Stretcher Mattresses offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered product must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Mattress Requirements
|  |  | a |  | The stretcher mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 200kg,
|  |  |  | ii | have a depth of at least 12cm
|  |  |  | iii | a hinged profile system that allows the mattress to adapt seamlessly to various articulated stretcher positions.
|  |  |  | iv | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  |  | v | sufficiently firm and supportive to enable the effective administration of CPR
|  |  |  | vi | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | vii | free of CFCs
|  |  |  | viii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions
|  |  | b |  | Mattresses prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric mattresses are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Participating Health Services may prefer radiolucent mattresses. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Labelling
|  |  | a |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | date of manufacture
|  |  |  | iii | model designation
|  |  |  | iv | batch or serial number
|  |  |  | v | Safe Working Load (SWL) expressed in kilograms
|  |  |  | vi | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | vii | correct orientation and positioning of the support surface
|  |  | b |  | Participating Health Services must be permitted to write the facility ID on mattresses without voiding the warranty. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The mattress cover should include a space to write the date of purchase. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress Cover
|  |  | a |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
|  | 3.3 |  |  | Pressure Distribution
|  |  | a |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 4 |  |  | Product Information
|  |  | a |  | For each stretcher mattress offered, the Respondent must submit:
|  |  |  | i | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each stretcher mattress offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | weight of the offered mattress (kg)
|  |  |  | vi | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | vii | the external dimensions of the offered mattress, Width x Length x Height (cm)
• list all materials used in the construction of the mattress core
• list all materials (foam types) used in the construction of the mattress core
• describe the antimicrobial treatment applied to the mattress core
• detail the microclimate management technologies integrated into the mattress
|  |  |  | viii | mattress width uniform from end to end (Yes/No)
• if No, describe width changes
|  |  |  | ix | mattress uniform in height and density across its surface (Yes/No)
• if No, what are the differences or inconsistencies?
|  |  |  | x | castellated cells (Yes/No)
• size of castellation, Width x Length (cm)
|  |  |  | xi | hinged for backrest (Yes/No)
|  |  |  | xii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
• describe the steps required to reveal the mattress core
|  |  |  | xiii | mattress cover includes handles (Yes/No)
• if Yes, describe the method used for joining and bonding handles to mattress cover
• if Yes, reinforced backing present to support handles (Yes/No)
|  |  |  | xiv | manufacturer of offered equipment
• country of origin
|  |  |  | xv | Year 1st manufactured
|  |  |  | xvi | Estimated device's end-of-sales life (state the year)
|  |  |  | xvii | warranty period of mattress (months)
|  |  |  | xviii | warranty period of mattress cover (months)
|  |  |  | xix | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xx | Estimated lifespan of device (years of service)
|  |  |  | xxi | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 8.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 8.0 |  |  |  | Operating Tables | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 8.01 Universal Operating Tables
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Operating Table
|  |  |  | i | A mobile, electrically powered table featuring a stable top surface supported by a fixed pedestal and a movable swivel base. Its surface comprises multiple articulated sections that can be independently elevated or lowered to contour the various anatomical positions, ensuring optimal support for the patient and enhanced exposure of the surgical field. Designed for compatibility with specialised accessories, such as positioners and traction devices, the table meets the diverse requirements of numerous clinical specialities. Additionally, designed to protect the patient from excessive manipulation, trauma, and abrasion. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Offered equipment must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The offered equipment must comply with the performance requirements specified in IEC 60601.2.46. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Operating Tables must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Operating Tables must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 350kg
|  |  |  | ii | operate on battery power
|  |  |  | iii | • run on mains power in the event of battery failure
|  |  |  | iv | • connect to mains power to charge batteries
|  |  |  | v | mobile base that locks into place for use with self-leveling brakes
|  |  |  | vi | height adjustable table top
|  |  |  | vii | include an accessory siderail
|  |  |  | viii | remote control
|  |  |  | ix | manual override controls
|  |  | b |  | The table top must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have at least 3 articulating sections
|  |  |  | ii | radiolucent
|  |  |  | iii | features radiology plate holders
|  |  |  | iv | both leg and head sections should be removable
|  |  |  | v | both Trendelenburg and reverse Trendelenburg positioning
|  |  |  | vi | lateral tilt
|  |  |  | vii | slide functionality
|  |  |  | viii | C-arm accessibility
|  |  | c |  | Participating Health Services may prefer operating tables with heated mattresses. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must be able to offer a comprehensive Service and Maintenance agreement for the tendered equipment to the Participating Health Service. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Respondents should offer additional options, such as motorised-drive | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Offers for bariatric operating tables are encouraged with a high safe working load significantly exceeding 350kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Offers for paediatric operating tables with a narrower table top are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with all operating tables must meet all mandatory criteria in Category 8.03. Additionally, Respondents must complete Category 8.03 for any offered operating table that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories, Options, and Consumables for the offered with operating tables in Categories 8.05 to 8.26 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  |  | iii | list all consumables and disposables that can be used with offered equipment in Category 8.22 to 8.26
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading, be equipped with a liquid-tight seal at the equipment entry point, and feature an integrated mains cable storage system; should be orange in colour and 4.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries should provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Articulations
|  |  | a |  | Operating tables offered must achieve the following articulations: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | at least 20° for both Trendelenburg and reverse Trendelenburg
|  |  |  | ii | at least 10° lateral tilt, left and right
|  |  |  | iii | back section capable of articulating at least 55° upward from the horizontal plane and 25° downward from the horizontal plane
|  |  |  | iv | leg section capable of articulating at least 20° upward from the horizontal plane and 90° downward from the horizontal plane
|  |  |  | v | head section capable of articulating at least 45° upward from the horizontal plane and 80° downward from the horizontal plane
|  | 4 |  |  | Product Information
|  |  | a |  | For each model operating table offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model equipment offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered operating table, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | articulation drive system (electrohydraulic/electromechanical type)
|  |  |  | ix | operating table integrates with surgical robotics console (Yes/No)
• list compatible robotic system/s
|  |  |  | x | SWL of offered operating table:
• flat table, no movement, centred on column (kg)
• flat table, lift, centred on column (kg)
• min. SWL, for all articulated positions including max. longitudinal slide (kg)
|  |  |  | xi | adjustable height range from floor level of the table(cm)
|  |  |  | xii | table slide range, to head (cm)/to foot (cm)
• without additional accessories, which anatomical regions of an 180cm tall patient may be challenging to visualise on the offered operating table using a standard C-arm
|  |  |  | xiii | table rotation, where applicable (degrees)
|  |  |  | xiv | no. of sections listing each (back, head, etc.)
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• lateral tilt, left and right (degrees)
• head section articulation, up/down (degrees)
• back section articulation, up/down (degrees)
• leg section articulation, up/down (degrees)
• flex/reflex articulation (degrees)
|  |  |  | xv | describe the accessory siderails (size, brand)
|  |  |  | xvi | mobile or fixed table
|  |  |  | xvii | castor size, diameter (cm)
• describe castor (material, single/dual,
• conductive casters (Yes/No)
|  |  |  | xviii | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xix | list the patient positionings table can articulate to
|  |  |  | xx | list all accessories available for offered operating table
|  |  |  | xxi | list all options available for offered operating table
• perineal cutout (Yes/No)
• height of kidney elevation, where applicable (cm)
• types of controls available (hand, foot, IR, etc.)
• motorised-drive (Yes/No)
• heated pads (Yes/No)
• list all others
|  |  |  | xxii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the battery capacity (Ah) and nominal voltage (V)
• state the battery capacity (no. hours of use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery charge
• state the lifespan of the rechargeable battery (in years) under normal use
|  |  |  | xxiii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• table electronics
• handset
|  |  |  | xxiv | state the 'Drop Test' height for handset
|  |  |  | xxv | manufacturer of offered equipment
• country of origin
|  |  |  | xxvi | Year 1st manufactured
|  |  |  | xxvii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxviii | warranty period offered on structural frame (months)
|  |  |  | xxix | warranty period offered on electronics (months)
|  |  |  | xxx | warranty period offered on handset & cables (months)
|  |  |  | xxxi | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxii | Estimated lifespan of device (years of service)
|  |  |  | xxxiii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxxiv | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 8.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 8.0 |  |  |  | Operating Tables | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 8.02 Spinal Surgery Operating Tables
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Spinal Surgery Operating Table
|  |  |  | i | An operating table specifically designed to facilitate optimal patient positioning and unobstructed surgical team access for spinal surgery procedures. These tables typically feature a radiolucent, modular, open-frame design that allows for free abdominal hang. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Offered equipment must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The offered equipment must comply with the performance requirements specified in IEC 60601.2.46. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Operating table must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Operating Tables must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 200 kg
|  |  |  | ii | operate on battery power
|  |  |  | iii | • run on mains power in the event of battery failure
|  |  |  | iv | • connect to mains power to charge batteries
|  |  |  | v | two-column design with interchangeable table tops
|  |  |  | vi | mobile table
|  |  |  | vii | full length C-arm accessibility
|  |  |  | viii | O-arm compatibility
|  |  |  | ix | remote control
|  |  |  | x | manual override controls
|  |  | b |  | The table top must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | radiolucent
|  |  |  | ii | both Trendelenburg and reverse Trendelenburg positioning
|  |  |  | iii | electronic lateral tilt
|  |  |  | iv | manual rotation
|  |  | c |  | The Respondent must be able to offer a comprehensive Service and Maintenance agreement for the tendered equipment to the Participating Health Service. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric operating tables are encouraged with a high safe working load significantly exceeding 350kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Offers for paediatric operating tables with a narrower table top are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | Mattresses included with all operating tables must meet all mandatory criteria in Category 8.03. Additionally, Respondents must complete Category 8.03 for any offered operating table that come with a mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The Respondent must list all purchasable Accessories, Options, and Consumables for the offered with operating tables in Categories 8.05 to 8.26 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  |  | iii | list all consumables and disposables that can be used with offered equipment in Category 8.22 to 8.26
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading, be equipped with a liquid-tight seal at the equipment entry point, and feature an integrated mains cable storage system; should be orange in colour and 4.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries should provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Articulations
|  |  | a |  | Operating tables offered must achieve the following articulations: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | at least 10° for both Trendelenburg and reverse Trendelenburg
|  |  |  | ii | at least 15° electronic rotation
|  |  |  | iii | 360° manual rotation
|  |  |  | iv | table top should hinge down (flexion)
|  |  |  | v | table top should hinge up (extension)
|  | 4 |  |  | Product Information
|  |  | a |  | For each model operating table offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model equipment offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | vi | the external dimensions of the offered operating table, Width x Length (cm)
|  |  |  | vii | weight of the offered equipment (kg)
|  |  |  | viii | articulation drive system (electrohydraulic/electromechanical type)
|  |  |  | ix | operating table integrates with surgical robotics console (Yes/No)
• list the compatible robotics system/s
|  |  |  | x | SWL of offered operating table (kg)
|  |  |  | xi | adjustable height range from floor level of the table(cm)
|  |  |  | xii | no. of various table top models
• state Trendelenburg/reverse Trendelenburg articulation (degrees)
• electronic rotation (degrees)
• manual rotation (degrees)
• table flexion, hinge down (degrees)
• table extension, hinge up (degrees)
|  |  |  | xiii | castor size, diameter (cm)
• describe castor (material, single/dual,
• conductive casters (Yes/No)
|  |  |  | xiv | describe locking mechanism (central, each caster, pedal, indicator)
|  |  |  | xv | list the patient positionings table can articulate to
|  |  |  | xvi | list all accessories available for offered operating table
|  |  |  | xvii | list all options available for offered operating table
|  |  |  | xviii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the battery capacity (Ah) and nominal voltage (V)
• state the battery capacity (no. hours of use)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery charge
• state the lifespan of the rechargeable battery (in years) under normal use
|  |  |  | xix | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• table electronics
• handset
|  |  |  | xx | state the 'Drop Test' height for handset
|  |  |  | xxi | manufacturer of offered equipment
• country of origin
|  |  |  | xxii | Year 1st manufactured
|  |  |  | xxiii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiv | warranty period offered on structural frame (months)
|  |  |  | xxv | warranty period offered on electronics (months)
|  |  |  | xxvi | warranty period offered on handset & cables (months)
|  |  |  | xxvii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxviii | Estimated lifespan of device (years of service)
|  |  |  | xxix | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxx | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 8.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 8.0 |  |  |  | Operating Tables | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 8.03 Operating Table Mattresses
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Operating Table Mattress
|  |  |  | i | A foam-field cushion designed as a component of an operating table, providing a supportive surface on which the patient is positioned during surgical procedures. It’s typically constructed from high-density, fire-retardant foam and is enclosed in a waterproof, antimicrobial cover that can withstand strong disinfectants. A reusable device. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Offered equipment should be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Mattress Requirements
|  |  | a |  | The operating table mattress is to conform to the various articulated sections of the operating table. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The operating table mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified table SWL
|  |  |  | ii | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | iii | free of CFCs
|  |  |  | iv | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions
|  |  |  | v | radiolucent
|  |  | c |  | Operating table mattresses prone to microbial growth must be treated with an antimicrobial agent during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Mattress Cover
|  |  | a |  | The mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | biocompatible
|  |  |  | vi | static dissipative
|  |  |  | vii | resistant to degradation by hospital-grade disinfectants
|  |  |  | viii | treated with an antimicrobial agent
|  |  |  | ix | latex-free
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | made from a robust, durable material
|  |  |  | ii | remain firmly and safely secured in position on the operating table
|  | 3.2 |  |  | Pressure Distribution
|  |  | a |  | The submission of a pressure-mapping report is highly encouraged, as Participating Health Services may prioritise products that include one. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The report should include both 2D and 3D heat maps or similar visual representation of the pressure-mapping results, and a detailed description of the testing methodology used.
|  |  |  | ii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), may be preferred.
|  |  |  | iii | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 4 |  |  | Product Information
|  |  | a |  | For each mattress offered, the Respondent must submit:
|  |  |  | i | PDF brochure featuring offered product | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each mattress offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number, where applicable | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s (for mattress)
|  |  |  | iii | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | iv | weight of the offered product (kg)
|  |  |  | v | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | vi | the external dimensions, Width x Length x Height (cm)
• list all materials used in the construction of the mattress/pad core
|  |  |  | vii | manufacturer of offered equipment
• country of origin
|  |  |  | viii | warranty period offered (months)
|  |  |  | ix | Year 1st manufactured
|  |  |  | x | Estimated device's end-of-sales life (state the year)
|  |  |  | xi | Estimated expected service life of device (years of service)

## 8.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 8.0 |  |  |  | Operating Tables | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 8.04 Patient Positioning Pads & Mattress Pads
|  | a |  |  | Intended Purpose
|  |  | a |  | Patient Positioning Pads & Mattress Pads
|  |  | e |  | Specialised pads used to supplement a mattress, providing enhanced support and pressure relief during extended periods of use or for specific activities. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Single-use pads are to be tendered in the operating tables consumables category.
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered reusable product casings (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered reusable products must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The pad must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | latex free
|  |  |  | ii | durable to withstand repeated use
|  |  |  | iii | non-slip
|  |  |  | iv | radiolucent and MRI-safe
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | PDF brochure featuring offered product | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | In the Tender Response Worksheet, the Respondent must:
|  |  |  | i | ARTG number, where applicable | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | the hospital grade disinfectants that are compatible to use on the offered product
• detergents that are compatible to use on the offered product
|  |  |  | iii | weight of the offered product (kg)
|  |  |  | iv | the external dimensions, Width x Length x Height (cm)
• list all materials used in the construction of the mattress/pad core
|  |  |  | v | describe the purpose of the pad (in 5 words or less)
|  |  |  | vi | manufacturer of offered product
• country of origin
|  |  |  | vii | warranty period offered (months)
|  |  |  | viii | Year 1st manufactured
|  |  |  | ix | Estimated product's end-of-sales life (state the year)
|  |  |  | x | Estimated lifespan of reusable product (months of service)

## 9.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 9.0 |  |  |  | Patient Movers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 9.01 Bed Movers
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Bed Mover
|  |  |  | i | A battery-powered device that hitches to hospital beds and stretches, to assist moving them within the hospital; significantly reducing the physical strain and risk of injury for caregivers from pushing and pulling beds manually, especially when navigating busy corridors, inclines, or lifts. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Bed Movers offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Occupational Health and Safety
|  |  | a |  | Eliminate the strenuous pushing and pulling of hospital beds and stretchers | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The bed mover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | load capacity of moving at least 400kg
|  |  |  | ii | include an emergency cut-out/stop button and a horn
|  |  |  | iii | compatible with multiple brands and models of beds and stretchers, featuring a simple hitching method
|  |  |  | iv | operated by a single person with variable speed control
|  |  |  | v | effortlessly move across various surfaces, including carpet
|  |  |  | vi | features a key start mechanism
|  |  | b |  | Bed movers should feature integrated battery chargers. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents are encouraged to offer models with a load capacity significantly exceeding 400kg. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Power Supply
|  |  | a |  | Bed movers should incorporate a mains power connection for battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment with rechargeable batteries must facilitate in-situ charging, eliminating the need for battery removal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments for operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must be durable enough to endure multiple 1.0m drops. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Bed mover offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | v | Compatibility chart - designed to be used with hospital beds stretches and trolleys | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model bed mover offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | external dimensions, Width x Length (cm)
• attached bed mover extends bed length (in cm)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | load capacity (kg)
|  |  |  | viii | turning circle (cm)
|  |  |  | ix | wheel size, diameter (cm)
• describe wheels/castors (material, etc.)
|  |  |  | x | list all accessories that are available offered product
|  |  |  | xi | list and describe the purpose of all controls found on the control panel
|  |  |  | xii | state the size of the display, H x W (cm) where applicable
|  |  |  | xiii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (distance travelled on single charge, km)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity
|  |  |  | xiv | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
|  |  |  | xv | state the 'Drop Test' height for handset
|  |  |  | xvi | manufacturer of offered equipment
• country of origin
|  |  |  | xvii | warranty period offered (months)
|  |  |  | xviii | Year 1st manufactured
|  |  |  | xix | Estimated device's end-of-sales life (state the year)
|  |  |  | xx | Estimated lifespan of device (years of service)
|  |  |  | xxi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 9.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 9.0 |  |  |  | Patient Movers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 9.02 Patient Lifters, Ceiling-mounted / Mobile-gantry
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Patient Lifters, Ceiling-mounted / Mobile-gantry
|  |  |  | i | An electrically powered system to safely and efficiently transfer individuals. It features an overhead track, are lifting cassette or motor, a patient suspension device (like a sling or harness), and a handheld control unit to control vertical movement. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | including lifters for:
• ceiling mounted rails
• mobile free standing rail systems
• swing lifts
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Patient lifters offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  |  |  | vi | emergency use instructions
|  |  |  | vii | pictogram instructions
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The patient lifter must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | fully suspend patient
|  |  |  | ii | safe working load of at least 200kg
|  |  |  | iii | include an emergency cut-out/stop button
|  |  |  | iv | include an emergency lowering feature
|  |  |  | v | battery operated
|  |  |  | vi | audible and visual alarms for low battery and overload conditions
|  |  |  | vii | quiet and smooth operation, to reduce patient anxiety
|  |  | b |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Mobile freestanding rail systems on castors, must be equipped with easily accessible, foot-operated brakes on all castors that require minimal force to engage/disengage. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Slings
|  |  | a |  | Respondents offering mobile patient lifters must also offer a range of patient support slings designed for use with the offered equipment. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Slings should be available in both reusable and single-patient-use options. Repositioning slings are encouraged. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Respondents offering reusable slings must include the laundering instructions with their tender submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Third-party providers are encouraged to offer patient support slings exclusively. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | third-party manufacturers of patient slings must submit compatibility matrices for patient hoists
|  |  | c |  | Manufacturers of slings must submit compatibility matrices detailing their slings' compatibility with the various devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 9.05 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  |  | iii | list all slings that can be used with offered equipment in either Category 9.10 or Category 9.11
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | weighing scales
|  |  |  | ii | interchangeable sling bars
|  |  |  | iii | an external charger, enabling the off-device charging of batteries
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment with rechargeable batteries must facilitate in-situ charging, eliminating the need for battery removal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments for operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must be durable enough to endure multiple 1.0m drops. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Controls intended for patient use must be easy to use, such as large tactile buttons or levers, with clear high-contrasting colours | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Patient Lifter offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
• including laundering instructions for the reusable slings
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | v | Compatibility chart outlining recommended sling and hoist pairings | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Patient Lifter offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | safe working load (kg)
|  |  |  | vi | type of mount (ceiling track/freestanding track/boom arm/etc.)
|  |  |  | vii | freestanding track system coverage range, Length x Width (m)
• track height range (m)
• mobile freestanding track system, on castors (Yes/No)
|  |  |  | viii | lifting unit (fixed/detachable) from track
|  |  |  | ix | the external dimensions of lifting unit, Width x Length x Height (cm)
|  |  |  | x | weight of the lifting unit (kg)
|  |  |  | xi | lifting stroke (cm)
|  |  |  | xii | lifting speed, maximum capacity load, SWL (mm/sec) 50mm/sec
|  |  |  | xiii | lifter compatible with repositioning slings (Yes/No)
• list sling bar/s suitable for repo slings
|  |  |  | xiv | list all accessories that are available offered product
|  |  |  | xv | list all options that are available offered product
• interchangeable sling bar (Yes/No)
- no. of sling suspension points available on sling bars
• weigh scales (Yes/No)
• pivot frame (Yes/No)
• Jordan frame (Yes/No)
• list all others
|  |  |  | xvi | list and describe the purpose of all controls found on the control panel
|  |  |  | xvii | state the size of the display, H x W (cm) where applicable
|  |  |  | xviii | battery charger (track mounted/portable)
|  |  |  | xix | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. of SWL 1m lifts)
• gauge indicating remaining battery capacity
|  |  |  | xx | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• lifting unit
• handset
• power supply
|  |  |  | xxi | state the 'Drop Test' height for handset
|  |  |  | xxii | manufacturer of offered equipment
• country of origin
|  |  |  | xxiii | Year 1st manufactured
|  |  |  | xxiv | Estimated device's end-of-sales life (state the year)
|  |  |  | xxv | warranty period offered on structural frame / track system (months)
|  |  |  | xxvi | warranty period offered on electronics / lifting unit (months)
|  |  |  | xxvii | warranty period offered on handset & cables (months)
|  |  |  | xxviii | Estimated lifespan of device (years of service)
|  |  |  | xxix | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxx | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 9.03
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 9.0 |  |  |  | Patient Movers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 9.03 Patient Lifters, Mobile
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Patient Lifters, Mobile
|  |  |  | i | A battery-powered device designed to safely and easily transfer incapacitated patients or individuals with disabilities over short distances within a treatment facility (e.g., from a bed to a bathroom). The lift features a V-shaped or U-shaped support base on casters, a motorised lifting mechanism, a vertical mast, a pivoting boom, a swivel bar, and a patient holding device such as a sling. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Patient lifters offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  |  |  | vi | emergency use instructions
|  |  |  | vii | pictogram instructions
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The patient lifter must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | fully suspend patient
|  |  |  | ii | safe working load of at least 200kg
|  |  |  | iii | include electric lifting mechanism/s
|  |  |  | iv | include an emergency cut-out/stop button
|  |  |  | v | include manual emergency lowering
|  |  |  | vi | battery operated
|  |  |  | vii | audible and visual alarms for low battery
|  |  | b |  | Patient lifter should include audible and visual alarms for overload conditions | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Participating Health Services may prefer patient lifters with a swappable battery system, where a spare battery can be kept in a dedicated charging station. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Slings
|  |  | a |  | Respondents offering mobile patient lifters must also offer a range of patient support slings designed for use with the offered equipment. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Slings should be available in both reusable and single-patient-use options. Repositioning slings are encouraged. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | Respondents offering reusable slings must include the laundering instructions with their tender submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Third-party providers are encouraged to offer patient support slings exclusively. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | third-party manufacturers of patient slings must submit compatibility matrices for patient hoists
|  |  | c |  | Manufacturers of slings must submit compatibility matrices detailing their slings' compatibility with the various devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Castors
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | At least two of the castors must incorporate a brake. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 9.05 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  |  | iii | list all slings that can be used with offered equipment in either Category 9.10 or Category 9.11
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | weighing scales
|  |  |  | ii | interchangeable sling bars
|  |  |  | iii | an external charger, enabling the off-device charging of batteries
|  |  | c |  | All optional features offered must comply with all applicable Occupational Health and Safety standards and meet the Human Factors Engineering requirements specified within this tender. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment with rechargeable batteries must facilitate in-situ charging, eliminating the need for battery removal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments for operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must be durable enough to endure multiple 1.0m drops. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Controls intended for patient use must be easy to use, such as large tactile buttons or levers, with clear high-contrasting colours | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model patient lifter offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
• including laundering instructions for the reusable slings
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | v | Compatibility chart outlining recommended sling and hoist pairings | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | safe working load (kg)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | the minimum storge dimensions of the offered equipment, Height x Width x Length (cm)
|  |  |  | viii | lifting stroke (cm)
• minimum lift height from floor (cm)
|  |  |  | ix | lifter compatible with repositioning slings (Yes/No)
• list sling bar/s suitable for repo slings
|  |  |  | x | turning circle (cm)
|  |  |  | xi | no. of electric motor/s
• list function of each
|  |  |  | xii | castor size/s, diameter (cm)
• no. of castors
• describe wheels/castors (dual, material, etc.)
• internal leg width range, open/closed (cm)
• external leg width range, open/closed (cm)
• leg clearance, floor to top of leg to slide under examination table / bed (cm)
|  |  |  | xiii | list all accessories that are available offered product
|  |  |  | xiv | list all options that are available offered product
• interchangeable sling bar (Yes/No)
- no. of sling suspension points available on sling bars
• weigh scales (Yes/No)
• pivot frame (Yes/No)
• Jordan frame (Yes/No)
• list all others
|  |  |  | xv | list and describe the purpose of all controls found on the handset
|  |  |  | xvi | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. of SWL lifts)
• charge time for a fully depleted battery (hours)
• gauge indicating remaining battery capacity (Yes/No)
|  |  |  | xvii | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• lifting unit
• handset
• power supply
|  |  |  | xviii | state the 'Drop Test' height for handset
|  |  |  | xix | manufacturer of offered equipment
• country of origin
|  |  |  | xx | Year 1st manufactured
|  |  |  | xxi | Estimated device's end-of-sales life (state the year)
|  |  |  | xxii | warranty period offered on structural frame (months)
|  |  |  | xxiii | warranty period offered on electronics (months)
|  |  |  | xxiv | warranty period offered on handset & cables (months)
|  |  |  | xxv | Estimated lifespan of device (years of service)
|  |  |  | xxvi | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 9.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 9.0 |  |  |  | Patient Movers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 9.04 Patient Lifters, Air-Assisted or other
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Patient Lifters, Air-Assisted
|  |  |  | i | An assembly of devices consisting of a mattress and a pump unit. The mattress is designed with features like patient straps, handles for transfer, air hose inlets. When activated, the pump unit inflates the mattress with air, causing it to rise several feet to elevate a patient to a desired height for easier transfers. A dedicated patient lifting system that uses air inflation to safely and efficiently raise individuals. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Other Patient Lifter Systems
|  |  |  | i | Any other powered device that facilitates the lifting of an individual, and which has not yet been specified.
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Patient lifters offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | The equipment (air supply unit) must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | model designation
|  |  |  | iii | serial number
|  |  |  | iv | power rating label
|  |  | b |  | The mattress must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number or lot number
|  |  |  | v | date of manufacture
|  |  |  | vi | mattress width
|  |  |  | vii | correct orientation and positioning of the support surface
|  |  |  | viii | pictogram instructions
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The patient lifter must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | air-supply unit connect to mains power
|  |  |  | ii | safe working load of at least 200kg
|  |  |  | iii | mattress placed under patient on floor and inflated to bed height
|  |  | b |  | The air-supply unit should be equipped with a battery backup system. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Transfer Mattresses
|  |  | a |  | Respondents offering air-assisted patient lifters must also offer transfer mattress/es designed for use with the offered system. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Transfer mattresses offered should be available in both reusable and single-patient-use options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Respondents offering reusable mattresses must include the laundering instructions with their tender submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Third-party providers are encouraged to offer transfer mattresses exclusively. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | third-party manufacturers of transfer mattresses must submit compatibility matrices for air-supply unit
|  |  | c |  | Manufacturers of mattresses must submit compatibility matrices detailing their mattress compatibility with the various devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 9.05 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  |  | iii | list all mattresses that can be used with offered equipment in either Category 9.14 or Category 9.15
|  |  | b |  | An optional ergonomically designed cart is encouraged for the easy, safe, and secure transport of the pump across the hospital; also providing a dedicated storage for the pump. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 5.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment with rechargeable batteries must facilitate in-situ charging, eliminating the need for battery removal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments for operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model air-assisted patient lift offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
• including laundering instructions for the reusable mattress
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | v | Compatibility chart outlining recommended mattress/cushion | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model air-assisted patient lift offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | safe working load (kg)
|  |  |  | vi | external dimension of mattress/es / cushion/s full inflated, Width x Length x Height (cm)
• mattress patient surface area, Width x Length (cm)
|  |  |  | vii | dimension of air-supply unit, Width x Length x Height (cm)
|  |  |  | viii | weight of the mattress/cushion (kg)
|  |  |  | ix | weight of the air-supply unit (kg)
|  |  |  | x | inflation time
|  |  |  | xi | no. of output settings on air-supply unit
• air flow range (L/min)
• maximum air pressure
|  |  |  | xii | list all accessories that are available offered product
|  |  |  | xiii | list all options that are available offered product
|  |  |  | xiv | list and describe the purpose of all controls found on the control panel
|  |  |  | xv | state the size of the display, H x W (cm) where applicable
|  |  |  | xvi | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. of SWL 1m lifts)
• gauge indicating remaining battery capacity
|  |  |  | xvii | state the Ingress Protection (IP) rating for both solids and liquids for the offered air-supply unit
|  |  |  | xviii | manufacturer of offered equipment
• country of origin
|  |  |  | xix | Year 1st manufactured
|  |  |  | xx | Estimated device's end-of-sales life (state the year)
|  |  |  | xxi | warranty period offered (months)
|  |  |  | xxii | Estimated lifespan of device (years of service)
|  |  |  | xxiii | Estimated lifespan of reusable mattress (years of service)
|  |  |  | xxiv | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 9.06
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 9.0 |  |  |  | Patient Movers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 9.06 Stand Transfer Aids, Mobile
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Stand Transfer Aids, Mobile
|  |  |  | i | Designed for healthcare professionals to safely transfer patients with disabilities over short distances in an upright (standing) position; featuring a rigid, low-centre-of-gravity metal frame with upper support bars for the user and bottom bars mounted on casters with brakes. Available in various configurations to suit specific needs, may include body-securing devices such as belts, footplates, and knee pads to help patients maintain a standing position. Other configurations may incorporate powered mechanisms to assist in moving the frame or to raise and lower the patient. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• can be a non-powered device, or
• can be an electrically-powered device
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  |  |  | vi | emergency use instructions
|  |  |  | vii | pictogram instructions
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Any stainless steel components must be marine grade 316 (preferred) or 304. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Stand Transfer Aid must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 175kg
|  |  |  | ii | hand grip for the patient to hold
|  |  |  | iii | knee support to provide leverage
|  |  |  | iv | footplate for a secure base
|  |  |  | v | for powered models, includes a secondary backup battery and corresponding charger.
|  |  |  | vi | quiet and smooth operation, to reduce patient anxiety
|  |  | b |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 175kg. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for models specifically designed and sized for paediatric use are encouraged. A reduced SWL will be viewed as appropriate. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Slings
|  |  | a |  | Respondents offering Standing Transfer Aids should also offer a range of patient support slings designed for use with the offered equipment. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Slings should be available in both reusable and single-patient-use options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Respondents offering reusable slings must include the laundering instructions with their tender submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Third-party providers are encouraged to offer patient support slings exclusively. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | third-party manufacturers of patient slings must submit compatibility matrices for patient hoists
|  |  | c |  | Manufacturers of slings must submit compatibility matrices detailing their slings' compatibility with the various devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Castors
|  |  | a |  | All castors must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | have fully enclosed housings to prevent debris ingress,
|  |  |  | ii | be anti-static and non-marking, and
|  |  |  | iii | provide smooth, low-rolling-resistance movement over typical hospital flooring, including tile, linoleum, and carpet, under maximum anticipated patient and equipment load.
|  |  | b |  | Participating Health Services may preference double-wheel castor systems. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | At least two of the castors must incorporate a brake. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 9.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  |  | iii | list all slings that can be used with offered equipment in either Category 9.12 or Category 9.13
|  |  | b |  | The following options should be offered: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | weighing scale for powered models
|  |  |  | ii | an external charger for powered models, enabling the off-device charging of batteries
|  | 3.4 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 2.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Participating Health Services may prefer in-situ charging of equipment with rechargeable batteries to eliminate the need for battery removal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.5 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments for operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Handsets must be durable enough to endure multiple 1.0m drops. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model stand transfer aid offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
• including laundering instructions for the reusable slings
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | v | Compatibility chart outlining recommended sling and device pairings | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model stand transfer aid offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | safe working load (kg)
|  |  |  | vi | weight of the offered equipment (kg)
|  |  |  | vii | the minimum storge dimensions of the offered equipment, Height x Width x Length (cm)
|  |  |  | viii | for powered devices, the lifting stroke (cm)
|  |  |  | ix | turning circle (cm)
|  |  |  | x | optimum patient size, height (cm)
|  |  |  | xi | outline the type of arm support e.g. crossbar handle for pt. to hold, arc armrest supports with hand grips, etc.
|  |  |  | xii | no. of sling suspension points
|  |  |  | xiii | no. of electric motor/s, if relevant
• list function of each
|  |  |  | xiv | castor size/s, diameter (cm)
• no. of castors
• describe wheels/castors (dual, material, etc.)
• internal leg width range, open/closed (cm)
• external leg width range, open/closed (cm)
• leg clearance, floor to top of leg to slide under examination table / bed (cm)
|  |  |  | xv | list all accessories that are available offered product
|  |  |  | xvi | list all options that are available offered product
|  |  |  | xvii | list and describe the purpose of all controls found on the handset
|  |  |  | xviii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. of SWL lifts)
• gauge indicating remaining battery capacity (Yes/No)
|  |  |  | xix | state the Ingress Protection (IP) rating for both solids and liquids for the offered:
• lifting unit
• handset
• power supply
|  |  |  | xx | state the 'Drop Test' height for handset
|  |  |  | xxi | manufacturer of offered equipment
• country of origin
|  |  |  | xxii | Year 1st manufactured
|  |  |  | xxiii | Estimated device's end-of-sales life (state the year)
|  |  |  | xxiv | warranty period offered on structural frame (months)
|  |  |  | xxv | warranty period offered on electronics (months)
|  |  |  | xxvi | warranty period offered on handset & cables (months)
|  |  |  | xxvii | Estimated lifespan of device (years of service)
|  |  |  | xxviii | availability of service training for health service biomedical engineers (Yes/No)
• can health service biomedical engineers service the equipment offered (Yes/No)
|  |  |  | xxix | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 9.07
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 9.0 |  |  |  | Patient Movers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 9.07 Full-body Transfer Boards
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Full-body Transfer Board
|  |  |  | i | A manual device used by healthcare providers to facilitate the physical movement and transfer of patients, allowing them to slide smoothly between surfaces such as a bed, stretcher, or imaging table. Typically constructed as a rigid, low-friction sliding board, these devices are commonly made of materials like wood or plastic. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Equipment offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | Equipment must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number
|  |  |  | v | date of manufacture
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | c |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | Full-body Transfer Boards must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 200kg
|  |  |  | ii | include a hanging hole
|  |  |  | iii | handles on each side
|  |  |  | iv | anti-slip / grip-pads on the underside
|  |  | b |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for built-in weighing scale models are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model full-body transfer board offered the Respondent must submit:
|  |  |  | i | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model full-body transfer board offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iii | the hospital grade disinfectants that are compatible to use on the offered equipment
|  |  |  | iv | safe working load (kg)
|  |  |  | v | weight of the offered equipment (kg)
|  |  |  | vi | the dimensions of the offered product, Length x Width x Thickness (cm)
|  |  |  | vii | no. of handles
|  |  |  | viii | Radiolucent (Yes/No)
|  |  |  | ix | bendable/foldable (Yes/No)
|  |  |  | x | built-in weighing scales (Yes/No)
|  |  |  | xi | manufacturer of offered equipment
• country of origin
|  |  |  | xii | warranty period offered (months)
|  |  |  | xiii | Year 1st manufactured
|  |  |  | xiv | Estimated device's end-of-sales life (state the year)
|  |  |  | xv | Estimated lifespan of device (years of service)
|  |  |  | xvi | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 9.08
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 9.0 |  |  |  | Patient Movers | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 9.08 Air-Assisted Transfer Devices
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Air-Assisted Transfer Device
|  |  |  | i | An assembly of devices consisting of the mattress and pump unit, designed to assist healthcare providers in sliding patients between surfaces like beds, stretches, or imaging tables. The mattress, equipped with patient straps, transfer handles, and air hose inlets, inflates to several inches high when activated by the pump unit. It is not intended for patient lifting. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Offered devices must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered equipment conforms to, within their Tender Response Worksheet submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Labelling
|  |  | a |  | The equipment (air supply unit) must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | model designation
|  |  |  | iii | serial number
|  |  |  | iv | power rating label
|  |  | b |  | The mattress must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | manufacturers name
|  |  |  | iii | model designation
|  |  |  | iv | serial number or lot number
|  |  |  | v | date of manufacture
|  |  |  | vi | mattress width
|  |  |  | vii | correct orientation and positioning of the support surface
|  |  |  | viii | pictogram instructions
|  | 2.3 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The patient lifter must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | air-supply unit connect to mains power
|  |  |  | ii | safe working load of at least 200kg
|  |  |  | iii | mattress placed under patient inflates to a height of several inches
|  |  | b |  | The air-supply unit should be equipped with a battery backup system. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged with a high safe working load significantly exceeding 200kg. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Transfer Mattresses
|  |  | a |  | Respondents offering air-assisted transfer devices must also offer transfer mattress/es designed for use with the offered system. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Transfer mattresses offered should be available in both reusable and single-patient-use options. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
[5 empty] | Respondents offering reusable mattresses must include the laundering instructions with their tender submission. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Third-party providers are encouraged to offer transfer mattresses exclusively. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | third-party manufacturers of transfer mattresses must submit compatibility matrices for air-supply unit
|  |  | c |  | Manufacturers of mattresses must submit compatibility matrices detailing their mattress compatibility with the various devices. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Options & Accessories
|  |  | a |  | The Respondent must list all purchasable Accessories and Options for the offered equipment in Category 9.09 of the Tender Response Worksheet. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Accessory is a non-essential, often peripheral, component that provides supplementary functionality, enhanced user convenience, or aesthetic customisation. It is not required for the core equipment to operate as intended and can typically be added or removed post-purchase. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Option is a customised variation or an additional component that enhances the equipment’s functionality or performance without changing its primary intended purpose. Options are typically chosen and added to the equipment at the time of purchase.
|  |  |  | iii | list all mattresses that can be used with offered equipment in either Category 9.14 or Category 9.15
|  |  | b |  | An optional ergonomically designed cart is encouraged for the easy, safe, and secure transport of the pump across the hospital; also providing a dedicated storage for the pump. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Power Supply
|  |  | a |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | b |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 5.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | d |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Equipment with rechargeable batteries must facilitate in-situ charging, eliminating the need for battery removal. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.4 |  |  | Controls
|  |  | a |  | Controls must be designed to enable safe and intuitive adjustments for operator, featuring clear, conveniently located, and easily accessible controls with unambiguous and durable labelling. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model air-assisted patient lift offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
• including laundering instructions for the reusable mattress
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | a comprehensive Total Cost of Ownership (TCO) analysis covering the first ten years of the product's anticipated lifecycle. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | v | Compatibility chart outlining recommended mattress/cushion | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model air-assisted patient lift offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | safe working load (kg)
|  |  |  | vi | dimension of mattress/es / cushion/s, Width x Length x Height (cm)
• mattress patient surface area, Width x Length (cm)
|  |  |  | vii | dimension of air-supply unit, Width x Length x Height (cm)
|  |  |  | viii | weight of the mattress/cushion (kg)
|  |  |  | ix | weight of the air-supply unit (kg)
|  |  |  | x | inflation time
|  |  |  | xi | no. of output settings on air-supply unit
• air flow range (L/min)
• maximum air pressure
|  |  |  | xii | list all accessories that are available offered product
|  |  |  | xiii | list all options that are available offered product
|  |  |  | xiv | list and describe the purpose of all controls found on the control panel
|  |  |  | xv | state the size of the display, H x W (cm) where applicable
|  |  |  | xvi | state the Ingress Protection (IP) rating for both solids and liquids for the offered air-supply unit
|  |  |  | xvii | manufacturer of offered equipment
• country of origin
|  |  |  | xviii | warranty period offered (months)
|  |  |  | xix | Year 1st manufactured
|  |  |  | xx | Estimated device's end-of-sales life (state the year)
|  |  |  | xxi | Estimated lifespan of device (years of service)
|  |  |  | xxii | Estimated lifespan of reusable mattress (years of service)
|  |  |  | xxiii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 10.01
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 10.0 |  |  |  | Body Support Surfaces | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 10.01 Hospital Bed Mattresses
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Hospital Bed Mattress
|  |  |  | i | A fabric-encased foam support, featuring internal segmentation or compartmentalisation, designed to be placed on a mattress support platform to provide a comfortable surface for patients to lie, rest, and sleep, while distributing their body weight over a broad area to relieve pressure points and prevent pressure injuries. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Mattresses offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | including flammability standards, and
|  |  |  | ii | compliance with the ISO 20342 series of standards is encouraged; additionally, manufacturers should account for the requirements of the parts of this standard that are currently under development.
|  |  | c |  | The mattress must comply with the durability and performance requirements specified in the ISO 20342 standard (all relevant parts) and any associated collateral and particular standards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Mattress Requirements
|  |  | a |  | The hospital bed mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | safe working load of at least 200kg,
|  |  |  | ii | have a depth of at least 15cm (20cm for Bariatric models)
|  |  |  | iii | a hinged profile system that allows the mattress to adapt seamlessly to various articulated bed positions
|  |  |  | iv | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  |  | v | sufficiently firm and supportive to enable the effective administration of CPR
|  |  |  | vi | free of CFCs
|  |  |  | vii | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | viii | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  | b |  | Mattresses prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric mattresses are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Labelling
|  |  | a |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | date of manufacture
|  |  |  | iii | model designation
|  |  |  | iv | batch or serial number
|  |  |  | v | Safe Working Load (SWL) expressed in kilograms
|  |  |  | vi | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | vii | correct orientation and positioning of the support surface
|  |  | b |  | Participating Health Services must be permitted to write the facility ID on mattresses without voiding the warranty. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The mattress cover should include a space to write the date of purchase. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress Cover
|  |  | a |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids (stitching is pervious to fluids)
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
|  |  | c |  | Mattress covers should be treated with an antimicrobial agent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.3 |  |  | Pressure Distribution
|  |  | a |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 4 |  |  | Product Information
|  |  | a |  | For each model mattress offered, the Respondent must submit:
|  |  |  | i | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model mattress offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | weight of the offered mattress (kg)
|  |  |  | vi | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | vii | the external dimensions of the offered mattress, Width x Length x Height (cm)
• list all materials used in the construction of the mattress core
• list all materials (foam types) used in the construction of the mattress core
• describe the antimicrobial treatment applied to the mattress core
• detail the microclimate management technologies integrated into the mattress
|  |  |  | viii | mattress width uniform from end to end (Yes/No)
• if No, describe width changes
|  |  |  | ix | mattress uniform in height and density across its surface (Yes/No)
• if No, what are the differences or inconsistencies?
|  |  |  | x | castellated cells (Yes/No)
• size of castellation, Width x Length (cm)
|  |  |  | xi | hinged for backrest (Yes/No)
|  |  |  | xii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
• describe the steps required to reveal the mattress core
|  |  |  | xiii | mattress cover includes handles (Yes/No)
• if Yes, describe the method used for joining and bonding handles to mattress cover
• if Yes, reinforced backing present to support handles (Yes/No)
|  |  |  | xiv | manufacturer of offered equipment
• country of origin
|  |  |  | xv | Year 1st manufactured
|  |  |  | xvi | warranty period of mattress (months)
|  |  |  | xvii | warranty period of mattress cover (months)
|  |  |  | xviii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xix | Estimated product's end-of-sales life (state the year)
|  |  |  | xx | Estimated expected service life of mattress (years of service)
|  |  |  | xxi | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 10.02
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 10.0 |  |  |  | Body Support Surfaces | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 10.02 Mattresses, Hybrid
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Hybrid Mattress
|  |  |  | i | A hospital bed mattress (Cat 11.01) which offers enhanced comfort and pressure redistribution by using other non-polyurethane foam components, e.g., gel toppers, static pre-filled air cells, etc., while still a reactive static surface. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Mattresses offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | including flammability standards, and
|  |  |  | ii | compliance with the ISO 20342 series of standards is encouraged; additionally, manufacturers should account for the requirements of the parts of this standard that are currently under development.
|  |  | c |  | The mattress must comply with the durability and performance requirements specified in the ISO 20342 standard (all relevant parts) and any associated collateral and particular standards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Mattress Requirements
|  |  | a |  | The mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | provide therapeutic pressure relief primarily through changes in pressure in the air cells
|  |  |  | ii | safe working load of at least 200kg,
|  |  |  | iii | have a depth of at least 15cm (20cm for Bariatric models)
a hybrid design, combining both dynamic and reactive therapies in a single mattress
|  |  |  | iv | non-powered
|  |  |  | v | a hinged profile system that allows the mattress to adapt seamlessly to various articulated bed positions
|  |  |  | vi | deliver effective pressure redistribution to minimise the risk of pressure injuries, and not "bottom-out" regardless of patient weight distribution, placement, or positioning within the specified SWL.
|  |  |  | vii | sufficiently firm and supportive to enable the effective administration of CPR
|  |  |  | viii | free of CFCs
|  |  |  | ix | constructed that prevents disintegration, fragmentation, or shedding of particles during normal use and cleaning of the mattress
|  |  |  | x | made from non-toxic, biocompatible materials that will not cause any deleterious effects on body tissue or elicit adverse skin reactions with prolonged contact
|  |  | b |  | Mattresses prone to microbial growth must have antimicrobial agents integrated into their inner core during manufacturing. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric mattresses are encouraged with a high safe working load significantly exceeding 200kg | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Labelling
|  |  | a |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | manufacturers name
|  |  |  | ii | date of manufacture
|  |  |  | iii | model designation
|  |  |  | iv | batch or serial number
|  |  |  | v | Safe Working Load (SWL) expressed in kilograms
|  |  |  | vi | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | vii | correct orientation and positioning of the support surface
|  |  | b |  | Participating Health Services must be permitted to write the facility ID on mattresses without voiding the warranty. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The mattress cover should include a space to write the date of purchase. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress Cover
|  |  | a |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids (stitching is pervious to fluids)
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
|  |  | c |  | Mattress covers should be treated with an antimicrobial agent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Pressure Distribution
|  |  | a |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  |  | b |  | Mattresses using viscoelastic materials must ensure that their physical properties (e.g., firmness, recovery time) remain consistent within the temperature range of 20 to 40°C. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Stretcher offered, the Respondent must submit:
|  |  |  | i | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model General Ward Bed offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the hospital grade disinfectants that are compatible to use on the offered equipment
• detergents that are compatible to use on the offered equipment
|  |  |  | v | weight of the offered mattress (kg)
|  |  |  | vi | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | vii | the external dimensions of the offered mattress, Width x Length x Height (cm)
• list all materials (foam types) used in the construction of the mattress core
• describe the mattress core (sidewalls, layers, compartments, etc.)
• describe the antimicrobial treatment applied to the mattress core
• detail the microclimate management technologies integrated into the mattress
• detail the hybrid support offered
|  |  |  | viii | mattress width uniform from end to end (Yes/No)
• if No, describe width changes
|  |  |  | ix | mattress uniform in height and density across its surface (Yes/No)
• if No, what are the differences or inconsistencies?
|  |  |  | x | castellated cells (Yes/No)
• size of castellation, Width x Length (cm)
|  |  |  | xi | hinged for backrest (Yes/No)
|  |  |  | xii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
• describe the steps required to reveal the mattress core
|  |  |  | xiii | mattress cover includes handles (Yes/No)
• if Yes, describe the method used for joining and bonding handles to mattress cover
• if Yes, reinforced backing present to support handles (Yes/No)
|  |  |  | xiv | manufacturer of offered equipment
• country of origin
|  |  |  | xv | Year 1st manufactured
|  |  |  | xvi | warranty period of mattress (months)
|  |  |  | xvii | warranty period of mattress cover (months)
|  |  |  | xviii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xix | Estimated device's end-of-sales life (state the year)
|  |  |  | xx | Estimated expected service life of mattress (years of service)
|  |  |  | xxi | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxii | Sustainability
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 10.04
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 10.0 |  |  |  | Body Support Surfaces | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 10.04 Mattress Systems, Alternating-Pressure
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Alternating-Pressure Mattress System
|  |  |  | i | An assembly of devices consisting of a mattress composed of multiple interconnected air cells (or air bladders) and a dedicated electric pump, which functions as a control unit. The control unit, connected to the mattress, systematically inflates and deflates these air cells in a cyclical pattern. This cyclical inflation and deflation actively shift the support points for the patient, thereby alternating the pressure distribution across their body. The primary purpose of these systems is to relieve pressure points for comfort and, to prevent and manage pressure injuries. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Systems offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | compliance with the ISO 20342 series of standards is encouraged; additionally, manufacturers should account for the requirements of the parts of this standard that are currently under development.
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Alternating-Pressure Mattress System must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a complete mattress replacement system, not an overlay
|  |  |  | ii | provide active therapeutic pressure relief primarily through cyclical changes in air pressure
|  |  |  | iii | include a CPR mode function enabling effective cardiopulmonary resuscitation when necessary
|  |  |  | iv | capable of sustaining therapeutic pressure in the event of a power failure of a few minutes by either having an internal battery or a readily accessible 'transport mode'
|  |  |  | v | continuously prevent bottoming out for a patient at its maximum rated weight capacity, throughout the entire alternating cycle, irrespective of the bed profile position.
|  |  | b |  | Offers for bariatric models are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Pump and Control Unit
|  |  | a |  | The Pump Control Unit must have: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | universal bed attachments
|  |  |  | ii | sensors located either in the pump unit or the mattress to monitor the air pressure in the mattress cells
|  |  |  | iii | a static pressure mode
|  |  |  | iv | automatic reversion to dynamic operation after a period of static operation
|  |  |  | v | audible and visual alarms for power failure and low pressure
|  |  |  | vi | control panel protected from accidental operation with a surface shaped to facilitate easy cleaning
|  |  | b |  | The pump unit noise emission should not exceed 40 dB at a distance of 1 meter. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connecting Tubes
|  |  | c |  | The connecting tubing between the pump and the mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable and from kink resistant materials
|  |  |  | ii | not poses any safety risk
|  |  |  | iii | adequative separation between connectors to allow easy disconnection of mattress
|  |  |  | iv | securely connect to pump unit to resist a minimum pull force of 25kg
|  |  |  | v | tubing attachments sufficiently robust and designed to accommodate the normal use of bed side rails
Power Supply
|  |  | d |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | e |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | g |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | l |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress
|  |  | a |  | The mattress core must be solely composed of multiple air cells that are interconnected and are continuously inflated or deflated with air via an electric pump. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The mattress must feature a power cord management system designed to keep power cables securely off the floor and away from under the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For safe and effective compatibility with most hospital beds, the mattress's length and width must closely match the dimensions of the bed's mattress support platform. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The mattress must feature a coverlet over the air cells. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The mattress should not exceed 10 kg in weight. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | Participating Health Services have expressed interest in the ability to customise mattresses with additional transverse cells. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | g |  | Mattresses must compensate for changes in bed position, especially when the backrest is elevated, to prevent bottoming out. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | The mattress must feature individual cells of adequate size and inflate to sufficient pressure, providing continuous air support for the patient's entire body throughout the mattress's inflation-deflation cycle.
|  |  | h |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.3 |  |  | Mattress Cover
|  |  | a |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids (stitching is pervious to fluids)
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
|  |  | c |  | Mattress covers should be treated with an antimicrobial agent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Labelling
|  |  | d |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | iii | correct orientation and positioning of the support surface
|  | 3.4 |  |  | CPR Release
|  |  | a |  | The system must include a clearly identifiable and easily accessible emergency (CPR) release mechanism for emergency procedures. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The emergency (CPR) mechanism must be designed and constructed to actively prevent unintentional or accidental deflation of the mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Quick guide instructions on activating the emergency mechanism must be prominently affixed to the connecting tubes and/or pump unit. These instructions must remain clearly visible and accessible when the bed is fully made up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Alternating-Pressure Mattress System offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Alternating-Pressure Mattress System offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered pump control unit
• detergents that are compatible to use on the offered pump control unit
|  |  |  | vi | the hospital grade disinfectants that are compatible to use on the offered mattress
• detergents that are compatible to use on the offered mattress
|  |  |  | vii | weight of the offered mattress (kg)
• external dimensions of the offered mattress fully deflated for storage/transport, Width x Length x Height (cm)
|  |  |  | viii | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | the external dimensions of the offered mattress fully inflated, Width x Length x Height (cm)
• list all materials used in the construction of the mattress core
• no. of transverse air cells
- can the mattress length be extended with additional transverse cells (Yes/No)
• describe the mattress core (sidewalls, pressure zones, compartments, etc.)
• does the mattress have side supports (Yes/No)
• detail the microclimate management technologies integrated into the mattress
|  |  |  | x | length of time mattress remains inflated with a 100kg subject in transport mode (minutes)
|  |  |  | xi | connecting tube length (cm)
|  |  |  | xii | emergency release mechanism (inflates/deflates) mattress
• time for mattress to deflate, ready for CPR (sec)
|  |  |  | xiii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
• describe the steps required to reveal the mattress core
|  |  |  | xiv | mattress cover includes handles (Yes/No)
• if Yes, describe the method used for joining and bonding handles to mattress cover
• if Yes, reinforced backing present to support handles (Yes/No)
|  |  |  | xv | the external dimensions of the offered pump, Width x Length x Height (cm)
|  |  |  | xvi | weight of the offered pump (kg)
|  |  |  | xvii | pump's maximum noise emission at 1 meter (dB)
|  |  |  | xviii | list all modes of operation (Transport, Static, Continuous Low Pressure, etc.)
|  |  |  | xix | automatically reverts to alternating-pressure from a static pressure mode (Yes/No)
• if Yes, time interval (range) on static pressure mode (minutes)
|  |  |  | xx | state the size of the display, H x W (cm) where applicable
• what information is shown on the display, list all
|  |  |  | xxi | list all alarms
|  |  |  | xxi | inflation level / comfort setting (Yes/No)
• comfort adjustment set (Automatic, Manual, Both)
• if Yes, the range (kg, mmHg, other)
|  |  |  | xxiii | pump output on inflation, into free air (L/min)
• inflation time from start-up to mattress maximum pressure (minutes)
• describe the alternating system (1:2, 1:3, Low Air Loss, Powered Static, etc.)
• minimum pressure level (range) during the deflation cycle (mmHg)
• in alternating mode, the cycle time/range (minutes)
|  |  |  | xxiv | unit has an air inlet filter (Yes/No)
|  |  |  | xxv | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity (Yes/No)
|  |  |  | xxvi | state the Ingress Protection (IP) rating for both solids and liquids for the pump
|  |  |  | xxvii | manufacturer of offered system
• country of origin
|  |  |  | xxviii | Year 1st manufactured
|  |  |  | xxix | warranty period of Pump Control Unit (months)
|  |  |  | xxx | warranty period of mattress (months)
|  |  |  | xxxiv | warranty period of mattress cover (months)
|  |  |  | xxxii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxiii | Estimated end-of-sales life of system (state the year)
|  |  |  | xxxiv | Estimated expected service life of Pump Control Unit (years of service)
|  |  |  | xxxv | Estimated expected service life of mattress (years of service)
|  |  |  | xxxvi | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxxvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 10.05
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 10.0 |  |  |  | Body Support Surfaces | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 10.05 Mattress Systems, Alternating-Pressure, Hybrid
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Alternating-Pressure Mattress System, Hybrid
|  |  |  | i | A mattress that offers the continuous support and comfort of a high-quality foam mattress, while also providing the active, dynamic pressure relief of an alternating air system, leading to a more comprehensive and often more comfortable solution for pressure injury management. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
| h | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Systems offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | compliance with the ISO 20342 series of standards is encouraged; additionally, manufacturers should account for the requirements of the parts of this standard that are currently under development.
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Mattress System must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a complete mattress replacement system, not an overlay
|  |  |  | ii | have a depth of at least 15cm (20cm for Bariatric models)
|  |  |  | iii | a hybrid design, combining both dynamic and reactive therapies in a single mattress
|  |  |  | iv | a powered system
|  |  |  | v | provide therapeutic pressure relief primarily through cyclical changes in air pressure in the air cells
|  |  |  | vi | sufficiently firm and supportive to enable the effective administration of CPR
|  |  |  | vii | capable of sustaining therapeutic pressure in the event of a power failure
|  |  |  | viii | continuously prevent bottoming out for a patient at its maximum rated weight capacity, throughout the entire alternating cycle, irrespective of the bed profile position.
|  |  | b |  | Offers for bariatric models are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Pump and Control Unit
|  |  | a |  | The Pump Control Unit must have: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | universal bed attachments
|  |  |  | ii | sensors located either in the pump unit or the mattress to monitor the air pressure in the mattress cells
|  |  |  | iii | a static (transport) mode
|  |  |  | iv | automatic reversion to dynamic operation after a period of static operation
|  |  |  | v | adjustable comfort/pressure levels
|  |  |  | vi | audible and visual alarms for power failure and low pressure
|  |  |  | vii | control panel protected from accidental operation with a surface shaped to facilitate easy cleaning
|  |  | b |  | The pump unit noise emission should not exceed 40 dB at a distance of 1 meter. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Connecting Tubes
|  |  | c |  | The connecting tubing between the pump and the mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable and from kink resistant materials
|  |  |  | ii | not poses any safety risk
|  |  |  | iii | adequative separation between connectors to allow easy disconnection of mattress
|  |  |  | iv | securely connect to pump unit to resist a minimum pull force of 25kg
|  |  |  | v | tubing attachments sufficiently robust and designed to accommodate the normal use of bed side rails
Power Supply
|  |  | d |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | e |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | g |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | l |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress
|  |  | a |  | The mattress must combine two or more technologies, typically air cells with a static medium. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The mattress must feature a power cord management system designed to keep power cables securely off the floor and away from under the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For safe and effective compatibility with most hospital beds, the mattress's length and width must closely match the dimensions of the bed's mattress support platform. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The mattress should not exceed 10 kg in weight. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | e |  | The mattress must compensate for changes in bed position, especially when the backrest is elevated, to prevent bottoming out. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | should feature patient movement sensors, such as those detecting Fowler position, to automatically adjust air pressure. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The mattress must feature individual cells of adequate size and inflate to sufficient pressure, providing continuous air support for the patient's entire body throughout the mattress's inflation-deflation cycle. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Multizone mattresses are encouraged | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.3 |  |  | Mattress Cover
|  |  | a |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids (stitching is pervious to fluids)
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
|  |  | c |  | Mattress covers should be treated with an antimicrobial agent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Labelling
|  |  | d |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | iii | correct orientation and positioning of the support surface
|  | 3.4 |  |  | CPR Release
|  |  | a |  | The system should feature a clearly identifiable and easily accessible emergency CPR release mechanism. Upon activation, this mechanism must rapidly deploy a firm support surface beneath the patient’s thoracic region to facilitate effective CPR. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The emergency (CPR) mechanism must be designed and constructed to actively prevent unintentional or accidental deflation of the mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Quick guide instructions on activating the emergency mechanism must be prominently affixed to the connecting tubes and/or pump unit. These instructions must remain clearly visible and accessible when the bed is fully made up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Alternating-Pressure Mattress System offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Alternating-Pressure Mattress System offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered pump control unit
• detergents that are compatible to use on the offered pump control unit
|  |  |  | vi | the hospital grade disinfectants that are compatible to use on the offered mattress
• detergents that are compatible to use on the offered mattress
|  |  |  | vii | weight of the offered mattress (kg)
• external dimensions of the offered mattress fully deflated for storage/transport, Width x Length x Height (cm)
|  |  |  | viii | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | the external dimensions of the offered mattress fully inflated, Width x Length x Height (cm)
• list all materials used in the construction of the mattress core
- describe the foam section
- no. of air cells
• air cell alignment (transverse only/longitudinal/both)
• list the various mattress zones (head, lower leg, etc.)
• mattress features cell-in-cell design (Yes/No)
• describe the mattress core (sidewalls, pressure zones, compartments, etc.)
• detail the microclimate management technologies integrated into the mattress
|  |  |  | x | connecting tube length (cm)
|  |  |  | xi | emergency release mechanism (inflates/deflates) mattress
• time for mattress to deflate, ready for CPR (sec)
|  |  |  | xii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
• describe the steps required to reveal the mattress core
• describe cable management features (sleeve for mains power cable)
|  |  |  | xiii | mattress cover includes handles (Yes/No)
• if Yes, describe the method used for joining and bonding handles to mattress cover
• if Yes, reinforced backing present to support handles (Yes/No)
|  |  |  | xiv | the external dimensions of the offered pump, Width x Length x Height (cm)
|  |  |  | xv | weight of the offered pump (kg)
|  |  |  | xvi | pump's maximum noise emission at 1 meter (dB)
|  |  |  | xvii | list all modes of operation (Transport, Static, CLP, LAL, max inflation, Deep Cell Therapy)
|  |  |  | xviii | time interval (range) on static (transport) mode (minutes)
|  |  |  | xix | state the size of the display, H x W (cm) where applicable
• what information is shown on the display, list all
|  |  |  | xx | list all alarms
|  |  |  | xxi | inflation level / comfort setting
• comfort adjustment set (Automatic, Manual, Both)
• mattress comfort setting range (kg, mmHg, other)
|  |  |  | xxii | location of pressure sensor (pump/cells/beneath mattress)
• function and location of other sensors (body weight distribution/fowler boost)
|  |  |  | xxiii | pump output on inflation, into free air (L/min)
• inflation time from start-up to mattress maximum pressure (minutes)
• describe the alternating system (1:2, 1:3, Low Air Loss, Powered Static, etc.)
• minimum pressure level (range) during the deflation cycle (mmHg)
• in alternating mode, the cycle time/range (minutes)
|  |  |  | xxiv | unit has an air inlet filter (Yes/No)
|  |  |  | xxv | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity (Yes/No)
|  |  |  | xxvi | state the Ingress Protection (IP) rating for both solids and liquids for the pump
|  |  |  | xxvii | manufacturer of offered system
• country of origin
|  |  |  | xxviii | Year 1st manufactured
|  |  |  | xxix | Estimated end-of-sales life of system (state the year)
|  |  |  | xxx | warranty period of Pump Control Unit (months)
|  |  |  | xxxi | warranty period of mattress (months)
|  |  |  | xxxii | warranty period of mattress cover (months)
|  |  |  | xxxiii | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxiv | Estimated expected service life of Pump Control Unit (years of service)
|  |  |  | xxxv | Estimated expected service life of mattress (years of service)
|  |  |  | xxxvi | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxxvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 10.06
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 10.0 |  |  |  | Body Support Surfaces | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 10.06 Mattress Systems, Low Air Loss
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Low Air Loss Mattress System
|  |  |  | i | An electrically powered device that provides dynamic pressure relief for patients by precisely controlling any flow through its dedicated mattress made of interconnected synthetic fabric air-field cushions. These systems are designed to lose air through the cushions at a controlled rate to regulate skin temperature and moisture of the patient’s skin and are essential for the care of patients requiring specialised pressure redistribution, such as individuals with, or susceptible to, pressure injury. The system comprises of the air-inflated pillow mattress, a compressor with a feed system, filters and a built-in control unit that manages the air cycle. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Systems offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | including flammability standards, and
|  |  |  | ii | compliance with the ISO 20342 series of standards is encouraged; additionally, manufacturers should account for the requirements of the parts of this standard that are currently under development.
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Low Air Loss Mattress System must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a complete mattress replacement system, not an overlay
|  |  |  | ii | constant airflow through the top of the mattress
|  |  |  | iii | include a CPR mode function enabling effective cardiopulmonary resuscitation when necessary
|  |  |  | iv | capable of sustaining therapeutic pressure in the event of a power failure of a few minutes by either having an internal battery or a readily accessible 'transport mode'
|  |  |  | v | continuously prevent bottoming out for a patient at its maximum rated weight capacity, throughout the entire alternating cycle, irrespective of the bed profile position.
|  |  | b |  | The mattress may provide active therapeutic pressure relief primarily through cyclical changes in air pressure | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Pump and Control Unit
|  |  | a |  | The Pump Control Unit must have: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | universal bed attachments
|  |  |  | ii | sensors located either in the pump unit or the mattress to monitor the air pressure in the mattress cells
|  |  |  | iii | a static (transport) mode
|  |  |  | iv | automatic comfort/pressure levels
|  |  |  | v | audible and visual alarms for power failure and low pressure
|  |  |  | vi | control panel protected from accidental operation with a surface shaped to facilitate easy cleaning
|  |  | b |  | The pump unit should: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | automatic reversion to dynamic operation after a period of static operation, where relevant
|  |  |  | ii | not exceed 40 dB noise emission at a distance of 1 meter.
Connecting Tubes
|  |  | c |  | The connecting tubing between the pump and the mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable and from kink resistant materials
|  |  |  | ii | not poses any safety risk
|  |  |  | iii | adequative separation between connectors to allow easy disconnection of mattress
|  |  |  | iv | securely connect to pump unit to resist a minimum pull force of 25kg
|  |  |  | v | tubing attachments sufficiently robust and designed to accommodate the normal use of bed side rails
Power Supply
|  |  | d |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | e |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | g |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | l |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress
|  |  | a |  | The mattress must feature a power cord management system designed to keep power cables securely off the floor and away from under the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For safe and effective compatibility with most hospital beds, the mattress's length and width must closely match the dimensions of the bed's mattress support platform. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | The mattress should feature a coverlet over the air cells. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The mattress should not exceed 10 kg in weight. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | e |  | The mattress must compensate for changes in bed position, especially when the backrest is elevated, to prevent bottoming out, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | should feature patient movement sensors, such as those detecting Fowler position, to automatically adjust air pressure.
|  |  | f |  | The mattress must feature individual cells of adequate size and inflate to sufficient pressure, providing continuous air support for the patient's entire body throughout the mattress's inflation-deflation cycle. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | g |  | Multizone mattresses are encouraged | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.3 |  |  | Mattress Cover
|  |  | a |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids (stitching is pervious to fluids)
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
|  |  | c |  | Mattress covers should be treated with an antimicrobial agent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Labelling
|  |  | d |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | iii | correct orientation and positioning of the support surface
|  | 3.4 |  |  | CPR Release
|  |  | a |  | The system must feature a clearly identifiable and easily accessible emergency CPR release mechanism. Upon activation, this mechanism must rapidly deploy a firm support surface beneath the patient’s thoracic region to facilitate effective CPR. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The emergency (CPR) mechanism must be designed and constructed to actively prevent unintentional or accidental deflation of the mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Quick guide instructions on activating the emergency mechanism must be prominently affixed to the connecting tubes and/or pump unit. These instructions must remain clearly visible and accessible when the bed is fully made up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Low-Air Loss Mattress System offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Low-Air Loss Mattress System offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered pump control unit
• detergents that are compatible to use on the offered pump control unit
|  |  |  | vi | the hospital grade disinfectants that are compatible to use on the offered mattress
• detergents that are compatible to use on the offered mattress
|  |  |  | vii | weight of the offered mattress (kg)
• external dimensions of the offered mattress fully deflated for storage/transport, Width x Length x Height (cm)
|  |  |  | viii | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | the external dimensions of the offered mattress fully inflated, Width x Length x Height (cm)
• list all materials used in the construction of the mattress core
- describe the foam section
• cell alignment (transverse only/longitudinal/both)
• list the various mattress zones (head 3 static cells; lower leg 5 cells)
• mattress features cell-in-cell design (Yes/No)
• describe the mattress core (sidewalls, pressure zones, compartments, etc.)
• detail the microclimate management technologies integrated into the mattress
|  |  |  | x | connecting tube length (cm)
|  |  |  | xi | emergency release mechanism (inflates/deflates) mattress
• time for mattress to deflate, ready for CPR (sec)
|  |  |  | xii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
• describe the steps required to reveal the mattress core
• describe cable management features (sleeve for mains power cable)
|  |  |  | xii | mattress cover includes handles (Yes/No)
• if Yes, describe the method used for joining and bonding handles to mattress cover
• if Yes, reinforced backing present to support handles (Yes/No)
|  |  |  | xiv | the hospital grade disinfectants that are compatible to use on the offered pump control unit
|  |  |  | xv | the external dimensions of the offered pump, Width x Length x Height (cm)
|  |  |  | xvi | weight of the offered pump (kg)
|  |  |  | xvii | pump's maximum noise emission at 1 meter (dB)
|  |  |  | xviii | list all modes of operation (Transport, Static, CLP, LAL, max inflation, Deep Cell Therapy)
|  |  |  | xix | time interval (range) on static (transport) mode (minutes)
|  |  |  | xx | state the size of the display, H x W (cm) where applicable
• what information is shown on the display, list all
|  |  |  | xxi | list all alarms
|  |  |  | xxii | automatic inflation level / comfort setting (Yes/No)
• mattress comfort setting range (kg, mmHg, other)
|  |  |  | xxiii | location of pressure sensor (pump/cells/beneath mattress)
• function and location of other sensors (body weight distribution/fowler boost)
|  |  |  | xxiv | pump output on inflation, into free air (L/min)
• inflation time from start-up to mattress maximum pressure (minutes)
• describe the alternating system (1:2, 1:3, Low Air Loss, Powered Static, etc.)
• minimum pressure level (range) during the deflation cycle (mmHg)
• in alternating mode, the cycle time/range (minutes)
|  |  |  | xxv | unit has an air inlet filter (Yes/No)
|  |  |  | xxvi | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity (Yes/No)
|  |  |  | xxvii | state the Ingress Protection (IP) rating for both solids and liquids for the pump
|  |  |  | xxviii | manufacturer of offered system
• country of origin
|  |  |  | xxix | Year 1st manufactured
|  |  |  | xxx | Estimated end-of-sales life of system (state the year)
|  |  |  | xxxi | warranty period of Pump Control Unit (months)
|  |  |  | xxxii | warranty period of mattress (months)
|  |  |  | xxxiii | warranty period of mattress cover (months)
|  |  |  | xxxiv | Estimated expected service life of Pump Control Unit (years of service)
|  |  |  | xxxv | Estimated expected service life of mattress (years of service)
|  |  |  | xxxvi | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxxvii | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 10.07
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here> |  | Model Name: <enter here> Model Number: <enter here> Ordering Part Number: <enter here>
| Cat 10.0 |  |  |  | Body Support Surfaces | Response Type | Compliance | Respondent's Response | Compliance | Respondent's Response | Compliance | Respondent's Response
Subcategory 10.07 Mattress Systems, Lateral Rotation/Percussion/Prone Position
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Mattress Systems, Lateral Rotation
|  |  |  | i | An electrically powered device that features a dedicated mattress composed of interconnected, air-filled synthetic fabric cushions. It provides dynamic pressure relief and facilitates lateral patient turning, either intermittently or continuously, through controlled airflow. These systems are designed to prevent and treat respiratory issues, skin complications, and related problems in critically ill, immobile patients, such as those who are ventilated or sedated. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | Mattress Systems, Percussion
|  |  |  | i | An electrically powered device that features a dedicated mattress composed of interconnected, air-filled synthetic fabric cushions. It’s designed to provide dynamic pressure relief and deliver percussion and/or vibration therapy. This therapy helps mobilise respiratory secretions in a patient’s lungs, making them easier to clear. Users can configure and deliver specific modes of therapy for preset durations and intensity levels.
|  |  | c |  | Mattress Systems, Prone Position
|  |  |  | i | A mattress system designed for the safe and effective positioning of patients in the prone position; simultaneously prevents and treats pressure injuries, ideal for managing patients with ARDS or other severe respiratory issues where prone therapy is used to improve lung function and oxygenation.
|  | 2 |  |  | General Requirements
|  | 2.1 |  |  | Standards and Compliance
|  |  | a |  | Systems offered must be included in the ARTG | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | All tendered medical electrical equipment must meet the durability and performance requirements specified in AS IEC 60601-1 and all relevant collateral and particular standards, as applicable to the equipment's intended use. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Respondents must list all the Standards that the offered product conforms to, within their Tender Response Worksheet submission,
|  |  |  | i | including flammability standards, and
|  |  |  | ii | compliance with the ISO 20342 series of standards is encouraged; additionally, manufacturers should account for the requirements of the parts of this standard that are currently under development. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 2.2 |  |  | Infection Control
|  |  | a |  | To facilitate cleaning and disinfection, the offered equipment casing (i.e., external coverings) must: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | be made from non-porous, smooth (preferred) durable material that is easily cleanable and resistant to common disinfectants,
|  |  |  | ii | be of colour that visibly contrasts with common contaminants, and
|  |  |  | iii | have minimal surface irregularities, such as coarse fabric, rough or pitted surfaces, holes, sharp edges, channels, crevices, or other potential dirt traps.
|  |  | b |  | Offered equipment must be capable of being cleaned without degradation with either:
|  |  |  | i | a TGA registered neutral detergent, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Isopropyl Alcohol 70% | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offered equipment (excluding the touchscreens) must be capable of being disinfected without degradation with at least one of the following disinfectant types:
|  |  |  | i | quaternary ammonium compounds, or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | Sodium Hypochlorite (bleach solutions), or | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | hydrogen peroxide solutions. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The Respondent must ensure that an Australian supplier continuously stocks the manufacturer’s recommended disinfectant. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3 |  |  | Product Requirements
|  |  | a |  | The Mattress System must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a complete mattress replacement system, not an overlay
|  |  |  | ii | provide active therapeutic pressure relief primarily through cyclical changes in air pressure
|  |  |  | iii | include a CPR mode function enabling effective cardiopulmonary resuscitation when necessary
|  |  |  | iv | capable of sustaining therapeutic pressure in the event of a power failure of a few minutes by either having an internal battery or a readily accessible 'transport mode'
|  |  |  | v | continuously prevent bottoming out for a patient at its maximum rated weight capacity, throughout the entire alternating cycle, irrespective of the bed profile position.
|  |  | b |  | The system must be capable of delivering:
|  |  |  | i | lateral rotation therapy, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | percussion/vibration therapy, | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | prone position | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | or a combined application of these modalities. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Offers for bariatric models are encouraged. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.1 |  |  | Pump and Control Unit
|  |  | a |  | The Pump Control Unit must have: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | universal bed attachments
|  |  |  | ii | sensors located either in the pump unit or the mattress to monitor the air pressure in the mattress cells
|  |  |  | iii | a static (transport) mode
|  |  |  | iv | adjustable comfort/pressure levels
|  |  |  | v | audible and visual alarms for power failure and low pressure
|  |  |  | vi | control panel protected from accidental operation with a surface shaped to facilitate easy cleaning
|  |  | b |  | The pump unit should: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | automatic reversion to dynamic operation after a period of static operation, where relevant
|  |  |  | ii | not exceed 40 dB noise emission at a distance of 1 meter.
Connecting Tubes
|  |  | c |  | The connecting tubing between the pump and the mattress must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | durable and from kink resistant materials
|  |  |  | ii | not poses any safety risk
|  |  |  | iii | adequative separation between connectors to allow easy disconnection of mattress
|  |  |  | iv | securely connect to pump unit to resist a minimum pull force of 25kg
|  |  |  | v | tubing attachments sufficiently robust and designed to accommodate the normal use of bed side rails
Power Supply
|  |  | d |  | All mains powered offered equipment must be supplied with an appropriate Australian plug (Type I) and be compatible with standard Australian 240V AC, 50Hz power outlets. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Mains Power Cable
|  |  | e |  | Power cables must exhibit high crush resistance to withstand repeated bed castor loading and be equipped with a liquid-tight seal at the equipment entry point; should be orange in colour and 3.0 m long. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | f |  | The equipment must incorporate a strain relief on the power cable, or a locking feature on the power inlet, to prevent unintended disconnection of the 240Vac power cable. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Batteries
|  |  | g |  | Disposable batteries must be a common household size, such as AA, AAA, C, D, or 9-volt. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Button batteries present a serious swallowing hazard, particularly to young children. Therefore, their use should be minimised whenever possible. When button batteries are essential, equipment must incorporate a secure, child-resistant battery compartment requiring a tool (not a coin) for access. Furthermore, the instruction manual must include detailed instructions and safety warnings regarding button battery handling and potential hazards. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Equipment designed to operate on battery power must include a clearly discernible low-battery indicator positioned ergonomically for ease of observation. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | j |  | Equipment using rechargeable batteries must charge the batteries in-situ, eliminating the need for specialised external battery charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | k |  | Rechargeable batteries must provide a minimum of two years of operational life, with minimal degradation in performance when subjected to partial charging. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | l |  | Rechargeable batteries should have long-lasting charge retention and short recharge times. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 3.2 |  |  | Mattress
|  |  | a |  | The mattress core must be solely composed of multiple air cells that are interconnected and are continuously inflated or deflated with air via an electric pump. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The mattress must feature a power cord management system designed to keep power cables securely off the floor and away from under the bed. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | For safe and effective compatibility with most hospital beds, the mattress's length and width must closely match the dimensions of the bed's mattress support platform. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | d |  | The mattress must feature a coverlet over the air cells. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | e |  | The mattress should not exceed 10 kg in weight. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Pressure Distribution
|  |  | f |  | The mattress must compensate for changes in bed position, especially when the backrest is elevated, to prevent bottoming out, and | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | should feature patient movement sensors, such as those detecting Fowler position, to automatically adjust air pressure.
|  |  | g |  | The mattress must feature individual cells of adequate size and inflate to sufficient pressure, providing continuous air support for the patient's entire body throughout the mattress's inflation-deflation cycle. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | h |  | Multizone mattresses are encouraged | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | i |  | Within 24 months of the contract start date, the supplier must provide a comprehensive pressure-mapping report for the product. This report must include a heatmap or a similar visual representation of the pressure-mapping results. | Info to Respondent | --Please Select-- |  | --Please Select-- |  | --Please Select--
• Acceptable Submissions: Existing pressure-mapping reports for tendered products will be deemed acceptable.
|  |  |  | i | The report should also include 3D pressure heatmaps.
|  |  |  | ii | Reports submitted after the tender closing date must include the following:
• a detailed description of the testing methodology used
• for human test subject: the gender, weight, height, and ethnicity (to account for anthropometric variability) of the participant/s
• for mannequin-based testing (if applicable): the make and model of the mannequin, along with its simulated anthropometric parameters (body dimensions, weight, and height)
• a heatmap or equivalent visual representation of the pressure-mapping results.
|  |  |  | iii | Test methodologies that adhere to recognised industry standards, such as RESNA SS-1 or ISO 20342 series (once published), are given strong preference.
|  |  |  | iv | All costs associated with this testing are the sole responsibility of the Respondent.
|  | 3.3 |  |  | Mattress Cover
|  |  | a |  | The offered mattress must be provided with a cover. The mattress cover must meet the following requirements: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | waterproof, resistant to penetration by biological fluids
|  |  |  | ii | vapour permeable
|  |  |  | iii | stain resistant
|  |  |  | iv | fire retardant
|  |  |  | v | designed with a low friction coefficient and low shear characteristics
|  |  |  | vi | biocompatible
|  |  |  | vii | exhibits multi-directional stretch and is designed to minimise wrinkling, remaining in close conformity with the mattress during all bed articulations
|  |  |  | viii | static dissipative
|  |  |  | ix | resistant to degradation by hospital-grade disinfectants
|  |  | b |  | The mattress cover must incorporate the following features: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | a robust, durable, non-slip base
|  |  |  | ii | all seams must be rendered impervious to fluids (stitching is pervious to fluids)
|  |  |  | iii | opening to allow for the easy removal and fitting mattress core
|  |  |  | iv | if a zipper is used (waterproof zipper preferable), it must be protected by a waterfall flap to prevent fluid ingress into the mattress core
|  |  |  | v | if mattress handles are fitted, they must be securely attached using a method that maintains fluid-impermeability, with the attachment points being reinforced
|  |  | c |  | Mattress covers should be treated with an antimicrobial agent. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
Labelling
|  |  | d |  | Mattresses must be clearly and permanently labelled with the following: | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | i | Safe Working Load (SWL) expressed in kilograms
|  |  |  | ii | mattress weight (for Manual Handling) expressed in kilograms
|  |  |  | iii | correct orientation and positioning of the support surface
|  | 3.4 |  |  | CPR Release
|  |  | a |  | The system must feature a clearly identifiable and easily accessible emergency CPR release mechanism. Upon activation, this mechanism must rapidly deploy a firm support surface beneath the patient’s thoracic region to facilitate effective CPR. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | The emergency (CPR) mechanism must be designed and constructed to actively prevent unintentional or accidental deflation of the mattress. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | c |  | Quick guide instructions on activating the emergency mechanism must be prominently affixed to the connecting tubes and/or pump unit. These instructions must remain clearly visible and accessible when the bed is fully made up. | Select Compliance | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  | 4 |  |  | Product Information
|  |  | a |  | For each model Mattress System offered, the Respondent must submit:
|  |  |  | i | complete operator user manuals in PDF format, and | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iii | an image of the offered equipment in JPEG format, with file less than 500KB | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | iv | pressure-mapping report | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  | b |  | For each model Mattress System offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | ARTG number | Data/Info Required | --Please Select-- |  | --Please Select-- |  | --Please Select--
|  |  |  | ii | UMDNS code and term
|  |  |  | iii | all the relevant product safety, performance, and essential performance standards that the offered equipment adheres to
• including Flammability Standard/s
|  |  |  | iv | the specific Electromagnetic Compatibility standard with which the offered equipment is compliant with
|  |  |  | v | the hospital grade disinfectants that are compatible to use on the offered pump control unit
• detergents that are compatible to use on the offered pump control unit
|  |  |  | vi | the hospital grade disinfectants that are compatible to use on the offered mattress
• detergents that are compatible to use on the offered mattress
|  |  |  | vii | weight of the offered mattress (kg)
• external dimensions of the offered mattress fully deflated for storage/transport, Width x Length x Height (cm)
|  |  |  | viii | SWL of offered mattress (kg)
• therapeutic weight range (min/max kg)
|  |  |  | ix | the external dimensions of the offered mattress fully inflated, Width x Length x Height (cm)
• list all materials used in the construction of the mattress core
• no. of transverse air cells
• list the various mattress zones (head 3 static cells; lower leg 5 cells)
• mattress features cell-in-cell design (Yes/No)
• describe the mattress core (sidewalls, pressure zones, compartments, etc.)
• detail the microclimate management technologies integrated into the mattress
length of time mattress remains inflated with a 100kg subject in transport mode (minutes)
|  |  |  | x | connecting tube length (cm)
|  |  |  | xi | emergency release mechanism (inflates/deflates) mattress
• time for mattress to deflate, ready for CPR (sec)
|  |  |  | xii | mattress cover:
• moisture vapour transmission rate (MVTR) (g/m2/day)
• describe the antimicrobial treatment applied to the mattress cover
• list and describe all methods used for joining and bonding the mattress cover fabric
• describe the steps required to reveal the mattress core
• describe cable management features (sleeve for mains power cable)
|  |  |  | xiii | mattress cover includes handles (Yes/No)
• if Yes, describe the method used for joining and bonding handles to mattress cover
• if Yes, reinforced backing present to support handles (Yes/No)
|  |  |  | xiv | the hospital grade disinfectants that are compatible to use on the offered pump control unit
|  |  |  | xv | the external dimensions of the offered pump, Width x Length x Height (cm)
|  |  |  | xvi | weight of the offered pump (kg)
|  |  |  | xvii | pump's maximum noise emission at 1 meter (dB)
|  |  |  | xviii | lateral therapy/percussion vibration therapy/combined therapy
|  |  |  | xix | list all modes of operation (Transport, Static, CLP, LAL, max inflation, Deep Cell Therapy)
|  |  |  | xx | time interval (range) on static (transport) mode (minutes)
|  |  |  | xxi | state the size of the display, H x W (cm) where applicable
• what information is shown on the display, list all
|  |  |  | xxii | list all alarms
|  |  |  | xxiii | inflation level / comfort setting (Yes/No)
• comfort adjustment set (Automatic, Manual, Both)
• if Yes, the range (kg, mmHg, other)
|  |  |  | xxiv | location of pressure sensor (pump/cells/beneath mattress)
• function and location of other sensors (body weight distribution/fowler boost)
|  |  |  | xxv | pump output on inflation, into free air (L/min)
• inflation time from start-up to mattress maximum pressure (minutes)
• describe the alternating system (1:2, 1:3, Low Air Loss, Powered Static, etc.)
• minimum pressure level (range) during the deflation cycle (mmHg)
• in alternating mode, the cycle time/range (minutes)
|  |  |  | xxvi | unit has an air inlet filter (Yes/No)
|  |  |  | xxvii | state the battery chemistry (SLA, Nicad, Li-ion, Alkaline, etc.)
• state the lifespan of the rechargeable battery (in years) under normal use
• state the battery capacity (no. hours of continuous use)
• gauge indicating remaining battery capacity (Yes/No)
|  |  |  | xxviii | state the Ingress Protection (IP) rating for both solids and liquids for the pump
|  |  |  | xxix | manufacturer of offered system
• country of origin
|  |  |  | xxx | Year 1st manufactured
|  |  |  | xxxi | Estimated end-of-sales life of system (state the year)
|  |  |  | xxxii | warranty period of Pump Control Unit (months)
|  |  |  | xxxiii | warranty period of mattress (months)
|  |  |  | xxxiv | warranty period of mattress cover (months)
|  |  |  | xxxv | list any extended warranty offerings
• list extended warranty exclusions
|  |  |  | xxxvi | Estimated expected service life of Pump Control Unit (years of service)
|  |  |  | xxxvii | Estimated expected service life of mattress (years of service)
|  |  |  | xxxviii | Estimated expected service life of mattress cover (years of service)
|  |  |  | xxxix | Sustainability:
• equipment packaging that can be locally recycled (%)
• recycled content in equipment packaging (%)
• equipment packaging removal completed by supplier (Yes/No)
• supplier offers a comprehensive equipment disposal program, ensuring responsible recycling and certified data destruction (Yes/No)

## 13.0
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  |  |  | No extra columns to be added by Respondent
| Cat 13.0 |  |  |  | Spare Parts | Response Type | Compliance | Respondent's Response
|  | 1 |  |  | Intended Purpose
|  |  | a |  | A spare part is any component or assembly necessary for performing schedule preventative or unscheduled corrective maintenance on the offered equipment to ensure its continued operation. | Info to Respondent | --Please Select--
|  | 2 |  |  | Spare Parts
|  |  | a |  | The Respondent must provide a complete list of all OEM spare parts necessary to ensure the offered equipment remains fully functional for its entire operational life cycle. This list must be provided in the Tender Response Worksheet. | Select Compliance | --Please Select--
|  |  | b |  | A maintenance kit contains the spare parts needed for routine periodic maintenance. | Select Compliance | --Please Select--
|  |  |  | i | The Respondent must list all maintenance kits required for the offered equipment's lifespan within the Tender Response Worksheet.
|  |  | c |  | The Respondent must list in the Tender Response Worksheet, all proprietary testers, calibration kits, specialised tools, cables and dongles that are exclusively obtainable from the OEM and are essential for the repair and ongoing maintenance of the offered equipment. | Select Compliance | --Please Select--
|  |  |  | i | Generic parts such as mains power leads, screws, washers, and fuses, should not be included in response. | Info to Respondent | --Please Select--
• Parts that are functionally equivalent to standard commercial off-the-shelf items, and are readily available from major distributors (e.g., RS Components) must not be included in response.
|  |  | d |  | For each spare part or assembly listed in the Tender Response Worksheet, the Respondent must specify the model equipment for which it is required. | Select Compliance | --Please Select--
|  |  |  | i | The Respondent can provide spare parts for equipment that fall within scope of this tender, even if that specific equipment is not explicitly detailed in Respondent's reply. | Info to Respondent | --Please Select--
|  |  | e |  | Upon request, the Supplier must furnish all required spare parts and maintenance kits to enable trained PHS personnel to maintain the offered equipment. | Select Compliance | --Please Select--
|  | 3 |  |  | Product Information
|  |  | a |  | The Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | The Participating Health Service can purchase all necessary spare parts for the offered equipment (Yes/No) | Data/Info Required | --Please Select--
|  |  |  | ii | All proprietary spare parts are listed in the Tender Response Worksheet (Yes/No)
• brand and model of equipment listed for each spare part
|  |  |  | iii | All preventative maintenance kits are listed in the Tender Response Worksheet (Yes/No)
• brand and model of equipment listed for each preventative maintenance kit

## 14.0
return to Index page
Product Requirements
| Clause Identifiers |  |  |  | Text |  |  |  | No extra columns to be added by Respondent
| Cat 14.0 |  |  |  | Service & Maintenance Arrangements | Response Type | Compliance | Respondent's Response
|  | 1 |  |  | Intended Purpose
|  |  | a |  | Corrective Maintenance | Info to Respondent | --Please Select--
A fee paying service for single items where respondent identifies, isolates, and rectifies faults so that failed equipment is restored to operational condition within the manufacturer’s established tolerances.
This also includes any pay-as-you-go routine preventative maintenance services offered by the Supplier
|  |  | b |  | Maintenance Agreements
|  |  |  | i | A contract between the Supplier and Participating Health Service for Respondent to maintain Participating Health Service’s equipment. Maintenance agreements may range anywhere from full comprehensive, to partnership maintenance arrangements with in-house Biomedical Engineering Department, to parts-only arrangements.
|  |  | c |  | Razor-Razorblade Model (Price-Per-Treatment/Amortisation) Agreements
|  |  |  | i | A contract where the Supplier provides equipment to the Participating Health Service under a free-on-loan model, where equipment costs are recovered through consumable sales over a defined contract duration.
|  |  |  | ii | Throughout the term of the agreement, the Supplier will maintain the equipment in full working order, except where the Participating Health Service has mutually agreed in writing for an alternative party to assume maintenance responsibilities.
• Respondents are encouraged to offer service training to Participating Health Service staff for equipment upkeep.
|  |  |  | iii | At the end of the agreement, the Participating Health Service must have the option to purchase the equipment for a significantly reduced price.
|  |  | d |  | Managed Equipment Services Agreement
|  |  |  | i | A long-term agreement between a Participating Health Service and a Supplier, where the Supplier takes on the responsibility for the entire life-cycle management of the offered equipment, which can include:
• installation and commissioning
• maintenance and repair, which can include calibration and testing
• upgrades and technology refresh
• decommissioning and disposal
• reporting and performance monitoring
|  |  | e |  | Software Maintenance Agreements
|  |  |  | i | An agreement for the ongoing upkeep and improvement of licensed software. These provisions generally include, but are not limited to: software upgrades (encompassing minor code updates, patches, and major version releases), code fix modifications, performance enhancements, and the addition of new or enhanced functional capabilities integrated into the existing software solution.
|  |  |  | ii | Additionally, for software associated with medical devices, the agreement may also include provisions for necessary IT hardware upgrades throughout the useful service life of the medical devices.
|  |  | f |  | Technical Service Training
|  |  |  | i | A short form of education and instruction session/s offered by the Supplier to equip Participating Health Service staff with the knowledge, skills, and abilities required to effectively install, maintain, troubleshoot, and repair offered products, equipment, or systems.
|  |  |  | ii | Any Technician/Biomedical Engineer service training package offered to the Participating Health Service by the Supplier, including the initial free-of-charge servicing training and of any on-going annual training.
|  |  | g |  | Extended Warranties
|  |  |  | i | The provisions for the repairs and parts beyond the original manufacturer's warranty.
|  | 2 |  |  | General Requirements
|  |  | a |  | To ensure full transparency for all Participating Health Services, Supplies must submit to HSV any agreements they enter into with a Participating Health Service that covers products under this agreement within 30 days of signing. | Select Compliance | --Please Select--
|  |  | b |  | For all corrective maintenance performed on offered equipment for Participating Health Services, the Supplier must submit invoices to HSV as part of their quarterly sales reports. | Select Compliance | --Please Select--
|  |  | c |  | The Supplier must include all charges, including any provided free-of-charge, for technical service training delivered to Participating Health Services in their quarterly sales reports to HSV. | Select Compliance | --Please Select--
|  |  | d |  | The Respondent must detail any extended warranty option available beyond them manufacturer’s warranty, including their annual cost. The Participating Health Service will decide whether to purchase any extended warranty. | Select Compliance | --Please Select--
|  | 3 |  |  | Product Information
|  |  | a |  | For all agreement types, training packages, and extended warranties offered, the Respondent must submit:
|  |  |  | i | comprehensive product brochures in PDF format. | Data/Info Required | --Please Select--
|  |  | b |  | For each agreement type, training package, and extended warranty offered, the Respondent must provide the following information in the Product Description on the Tender Response Worksheet:
|  |  |  | i | list the various agreements Respondent can offer the Participating Health Service for the ongoing maintenance and management of offered equipment | Data/Info Required | --Please Select--
• outline all exclusions, if any
|  |  |  | ii | exchange program offered for faulty unit replacement (Yes/No)
|  |  |  | iii | will Respondent provide loan unit should equipment require repair (Yes/No)
• delivery time for loan unit and limitations
|  |  |  | iv | provide the corrective maintenance rates for:
• metro travel charge
• <160km travel charge
• >160km travel charge
• Labour rate - bench repairs (per hour)
• Labour rate - on-site business hours (per hour)
• Labour rate - on-site after hours (per hour)
|  |  |  | v | warranty period on Supplier's corrective maintenance repairs
|  |  |  | vi | Cybersecurity (for products with network connectivity):
• endpoint protection solutions compatible with the offered product

## Appendix 1-Categories
return to Index page
Appendix 1 - Category and Subcategory List
Category |  | Subcategory
1 | Hospital Beds, Acute Care | 1.01 | General Ward Beds, Single
|  | 1.02 | General Ward Beds, King Single
|  | 1.03 | General Ward High-Low Beds, Single
|  | 1.04 | General Ward High-Low Beds, King Single
|  | 1.05 | Bariatric Beds
|  | 1.06 | Critical Care Beds
|  | 1.07 | Stand-Up Beds
|  | 1.08 | Homecare Beds
2 | Hospital Beds, Mental Health | 2.01 | Behavioural Healthcare Beds
|  | 2.02 | Seclusion Room Beds
|  | 2.03 | Behavioural Healthcare Mattresses
|  | 2.04 | Mental Health Care Bed Accessories & Options
3 | Neonate & Paediatric Beds | 3.01 | Bassinets
|  | 3.02 | Cots
|  | 3.03 | Youth Beds
|  | 3.04 | Infant Incubators
|  | 3.05 | Infant Incubators, Transport
|  | 3.06 | Infant Incubator/Radiant Warmers
|  | 3.07 | Infant Resuscitation Trolleys
|  | 3.08 | Bassinet Mattresses/Accessories/Options
|  | 3.09 | Infant Cot Mattresses/Accessories/Options
|  | 3.1 | Infant Incubator/Radiant Warmer/Resus Trolley Accessories/Mattresses
|  | 3.11 | Infant Incubator/Radiant Warmer/Resus Trolley Temperature Probes
|  | 3.12 | Infant Incubator/Radiant Warmer/Resus Trolley Consumables
4 | O&G Beds | 4.01 | Birthing Beds
|  | 4.02 | Gynaecology Chairs
|  | 4.03 | Gynae Exam Couches
|  | 4.04 | Hospital Stretchers, Obstetric
|  | 4.05 | Birthing Bed Mattresses/Accessories/Options
|  | 4.06 | Gynaecology Chair/Couch Accessories & Options
|  | 4.07 | Ultrasound Exam Tables, Obstetric
5 | Procedure Chairs | 5.01 | Haemodialysis / Chemotherapy Chairs
|  | 5.02 | Procedure Chairs
|  | 5.03 | Podiatry Chairs
|  | 5.04 | Ophthalmology / ENT Chairs
|  | 5.05 | Blood Collection Phlebotomy Chairs
|  | 5.06 | Procedure Chair Accessories & Options
6 | Examination Couches | 6.01 | Examination Plinths
|  | 6.02 | Examination Tables, 2-section
|  | 6.03 | Examination Tables, 3-section
|  | 6.04 | Podiatry Examination Couches
|  | 6.05 | Tilt Tables
|  | 6.06 | Traction Tables
|  | 6.07 | Ultrasound Exam Tables
|  | 6.08 | Ultrasound Exam Tables, Echocardiogram
|  | 6.09 | Ophthalmology / ENT Examination Couches
|  | 6.1 | Examination Table Accessories & Options
|  | 6.11 | Examination Table Consumables
7 | Patient Stretchers | 7.01 | Hospital Stretchers, Manual
|  | 7.02 | Hospital Stretchers, Powered
|  | 7.03 | Hospital Stretchers, Motorised-Drive
|  | 7.04 | Hospital Stretchers, Radiographic
|  | 7.05 | Hospital Stretchers, MRI Conditional
|  | 7.06 | Hospital Stretcher Chairs
|  | 7.07 | Hospital Stretchers, Ophthalmology / ENT
|  | 7.08 | Stretcher Mattresses
|  | 7.09 | Stretcher Accessories & Options
8 | Operating Tables | 8.01 | Universal Operating Tables
|  | 8.02 | Spinal Surgery Operating Tables
|  | 8.03 | Operating Table Mattresses
|  | 8.04 | Patient Positioning Pads & Mattress Pads
|  | 8.05 | Operating Table Accessories, Head Supports
|  | 8.06 | Operating Table Accessories, Arm Supports
|  | 8.07 | Operating Table Accessories, Leg Supports
|  | 8.08 | Operating Table Accessories, Body Supports
|  | 8.09 | Operating Table Accessories, Clamps
|  | 8.1 | Operating Table Accessories, Carbon Components
|  | 8.11 | Operating Table Accessories, Controls & Cords
|  | 8.12 | Operating Table Accessories, Shuttles & Carts
|  | 8.13 | Operating Table Accessories, Neuro
|  | 8.14 | Operating Table Accessories, Spinal
|  | 8.15 | Operating Table Accessories, Orthopaedic
|  | 8.16 | Operating Table Accessories, Urology & Gynaecology
|  | 8.17 | Operating Table Accessories, Bariatric
|  | 8.18 | Operating Table Accessories, Paediatric
|  | 8.19 | Operating Table Accessories, Anaesthesia
|  | 8.2 | Operating Table Accessories, Radiology
|  | 8.21 | Operating Table Options
|  | 8.22 | Operating Table Consumables
|  | 8.23 | Operating Table Consumables, Other
|  | 8.24 | Operating Table SPU Consumables, Spinal & Prone
|  | 8.25 | Operating Table SPU Consumables, Positioning
|  | 8.26 | Operating Table SPU Consumables, Orthopaedic
9 | Patient Movers | 9.01 | Bed movers
|  | 9.02 | Patient Lifters, Ceiling-mounted / Mobile-gantry
|  | 9.03 | Patient Lifters, Mobile
|  | 9.04 | Patient Lifters, Air-Assisted or other
|  | 9.05 | Patient Lifters, Accessories & Options
|  | 9.06 | Stand Transfer Aids, Mobile
|  | 9.07 | Full-body Transfer Board
|  | 9.08 | Air-Assisted Transfer Devices
|  | 9.09 | Transfer Aids, Accessories & Options
|  | 9.1 | Patient Lifter Slings, SPU
|  | 9.11 | Patient Lifter Slings, reusable
|  | 9.12 | Stand Transfer Aid Slings, SPU
|  | 9.13 | Stand Transfer Aid Slings, reusable
|  | 9.14 | Patient Lifter/Transfer Mattresses, SPU
|  | 9.15 | Patient Lifter/Transfer Mattresses, reusable
10 | Body Support Surfaces | 10.01 | Hospital Bed Mattresses
|  | 10.02 | Mattresses, Hybrid
|  | 10.03 | Mattress Covers
|  | 10.04 | Mattress Systems, Alternating-Pressure
|  | 10.05 | Mattress Systems, Alternating-Pressure, Hybrid
|  | 10.06 | Mattress Systems, Low Air Loss
|  | 10.07 | Mattress Systems, Lateral Rotation/Percussion/Prone Position
|  | 10.08 | Body Support Surface Accessories
11 | Rental (for in-hospital use) | 11.01 | Bariatric Bed
|  | 11.02 | Intensive Care Unit Bed
|  | 11.03 | General Ward Floor-level Bed
|  | 11.04 | General Ward Bed
|  | 11.05 | Mattress Systems, Alternating-Pressure
|  | 11.06 | Patient Lifters/Transfer Aids
|  | 11.07 | Patient Lifter/Transfer Slings, reusable
|  | 11.08 | Hospital Bed Mattresses
|  | 11.09 | Homecare Beds
|  | 11.1 | Birthing Bed
12 | Accessories | 12.01 | Bed Accessories & Options (hospital beds & youth beds)
13 | Spare Parts | 13.01 | Bed Spare Parts
|  | 13.02 | Bassinet Spare Parts
|  | 13.03 | Infant Cot Spare Parts
|  | 13.04 | Incubator Spare Parts
|  | 13.05 | Radiant Warmer/Resuscitation Trolley Spare Parts
|  | 13.06 | Birthing Bed Spare Parts
|  | 13.07 | Procedural Chair Spare Parts
|  | 13.08 | Examination Table Spare Parts
|  | 13.09 | Hospital Stretcher Spare Parts
|  | 13.1 | Operating Table Spare Parts
|  | 13.11 | Bed Mover Spare Parts
|  | 13.12 | Patient Lifter Spare Parts
|  | 13.13 | Patient Transfer Aid Spare Parts
|  | 13.14 | Patient Lifter/Transfer Mattress System Spare Parts
|  | 13.15 | Alternating-Pressure Pump Spare Parts
|  | 13.16 | Alternating-Pressure Mattress/Cushion Spare Parts
14 | Service & Maintenance | 14.01 | Extended Warranties
|  | 14.02 | BME Training
|  | 14.03 | Preventative Maintenance
|  | 14.04 | Corrective Maintenance
|  | 14.05 | Service Maintenance Agreements
|  | 14.06 | Test/Service Equipment
Products Out-of-Scope
Bariatric Chairs | Bariatic Chairs | out-of-scope (Clinical Furniture in the future)
Bedside Lockers | Bedside Lockers | out-of-scope (Clinical Furniture in the future)
Chiropractic Tables |  | out-of-scope
Concave Mattresses |  | out-of-scope (PRG considered restraining)
Dental Chairs | Dental Chairs | out-of-scope
Equipment hire for patient's discharge |  | out-of-scope
Massage Table |  | out-of-scope
Mattress Overlays |  | out-of-scope
Mortuary Lifter |  | HPVC2026-115
Mortuary Stretcher |  | HPVC2026-115
Over Bed Tables | Over Bed Tables | out-of-scope (Clinical Furniture in the future)
Patient Day/ Bedside Chairs | Patient Day/ Bedside Chairs | HPVC2025-163
Pillows | Pillows | out-of-scope
Princess Chairs |  | out-of-scope (Clinical Furniture in the future)
Recliners | Patient Recliners | out-of-scope (Clinical Furniture in the future)
Repositioning and Off-loading Devices |  | out-of-scope
Roho Cushions | Roho Cushions | HPVC2025-163
Shower and Commode Chairs | Shower and Commode Chairs | HPVC2025-163
Shower Trolleys |  | HPVC2025-163
Static Seat Cushions | Seat Cushions | out-of-scope
Stools for Clinicians |  | out-of-scope (Clinical Furniture in the future)
Upright Rehabilitation Walkers |  | out-of-scope
Remanufacturing of SPU products |  | out-of-scope

## Appendix 2-Document Reference
return to Index page
Appendix 2 - Document References
Standards | Standard Number | Standard Name
References to standards include any amendments, revisions or consolidations to those standards | AS/NZS IEC 60601.1 | Medical Electrical Equipment – Part 1: General requirements for basic safety and essential performance
| AS IEC 60601.1.2 | Medical electrical equipment – General requirements for basic safety and essential performance – Collateral Standard: Electromagnetic disturbances – Requirements and tests
| AS IEC 60601.1.6 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral standard: Usability
| AS IEC 60601.1.8 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems
| AS IEC 60601.1.9 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: Requirements for environmentally conscious design
| AS IEC 60601.1.11 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment
| AS IEC 60601.1.12 | Medical electrical equipment - General requirements for basic safety and essential performance - Collateral Standard: Requirements for medical electrical equipment and medical electrical systems intended for use in the emergency medical services environment
| IEC 60601-2-52 | Medical electrical equipment - Part 2-52: Particular requirements for the basic safety and essential performance of medical beds
| ISO 20342 series (all of) | Assistive products for tissue integrity when lying down
| ISO 22882 | Castors and wheels - Requirements for castors for hospital beds
| AS/NZS 4187 | Reprocessing of reusable medical devices in health service organizations
| AS/NZS 8811.1 | Methods of testing infant products, Method 1: Sleep surfaces - Test for firmness
| ANSI/RESNA SS-1 | American National Standard for Support Surfaces – Volume 1: Requirements and Test Methods for Full Body Support Surfaces.
Standards for Associated Services | Standard Number | Standard Name
| AS/NZS 3551 | Management programs for medical equipment
| AS/NZS 2500 | Guide to the safe use of electricity in patient care
Legislation | Name
References to legislation includes any amendments, revisions or consolidations to those references | Therapeutic Goods Act 1989
Electricity Safety Act 1998
Occupational Health and Safety Act 2004
Building Act 1993
Guidelines and Other References | Issuing Body | Title
References to guidelines include any amendments, revisions or consolidations to those guidelines | NPUAP (2025) | Prevention and Treatment of Pressure Ulcers/Injuries: Clinical Practice Guideline
| AWMA (2012) | Pan Pacific Clinical Practice Guideline for the Prevention and Management of Pressure Injury
| NHMRC (2010) | Australian Guidelines for the Prevention and Control of Infection in Healthcare, Commonwealth of Australia
| Therapeutic Goods Administration (1991) | Therapeutic Goods Order No. 37: General Requirements for Labels for Therapeutic Goods
| Therapeutic Goods Administration (2004) | Uniform Recall Procedure for Therapeutic Goods, Commonwealth of Australia
| WorkSafe Victoria (2000) | Code of Practice for Manual Handling
| WorkSafe Victoria (2009) | Transferring People Safely
| WorkSafe Victoria (2007) | A Guide to Designing Workplaces for Safer Handling of People

## Appendix 3-Tables
return to Index page
Appendix 3 - Tables
Delivery
Order type | Maximum delivery time | Time prior to installation
Equipment | 90 days | 7 days
Accessories | 90 days
Consumables / semi-consumables | 5 days
Batteries | 5 days
Spare Parts (common & essential) | 5 days
Preventative Maintenance Kits | 5 days
Urgent Deliveries | < 24 business hours (metro) < 48 business hours (regional & rural)
Service & Maintenance
Device Type | 1st Level Support | 2nd Level Support [1] | 3rd Level Support [2] | Spare Parts Delivery 'common' and 'essential'
Portable Equipment | Phone call returned within 4 business hours | Courier to service base and Returned to PHS within 15 business days |  | < 5 days
Operating Tables | phone call returned within 2 hours | On-site within 12 business hours (metro) 3 business days (regional & rural) | Loan Unit within 20 business hours (metro) 5 business days (regional) | < 5 days
all other Equipment | Phone call returned within 4 business hours | Repair at Respondent's service centre and returned repaired equipment within 15 business days |  | < 5 days
ICT Systems, and Central Stations et al. | phone call returned within 2 hours | Remote Access within 4 hours | On-site within 12 business hours (metro) 24 business hours (regional & rural) | 8 business hours
|  | [1] Unless resolved by phone support | [2] Unless resolved by phone support/ remote access